## Alanine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Alanine pos | 9.49 | Target Compound | KD082321_082521_B1 | 3553115114 |  | Cal Std | C3H7NO2 | 8.016 | 500 | 443277707
Alanine pos | 9.49 | Target Compound | KD082321_082521_B2 | 2154673749 |  | Cal Std | C3H7NO2 | 3.773 | 250 | 571146589
Alanine pos | 9.49 | Target Compound | KD082321_082521_B3 | 1368656300 |  | Cal Std | C3H7NO2 | 2.016 | 125 | 678905607
Alanine pos | 9.49 | Target Compound | KD082321_082521_B4 | 799766933 |  | Cal Std | C3H7NO2 | 1.034 | 62.5 | 773548415
Alanine pos | 9.49 | Target Compound | KD082321_082521_B5 | 477020240 |  | Cal Std | C3H7NO2 | 0.521 | 31.25 | 914968685
Alanine pos | 9.49 | Target Compound | KD082321_082521_B6 | 245638662 |  | Cal Std | C3H7NO2 | 0.264 | 15.625 | 930240994
Alanine pos | 9.49 | Target Compound | KD082321_082521_B7 | 135662864 |  | Cal Std | C3H7NO2 | 0.142 | 7.813 | 957287224
Alanine pos | 9.49 | Target Compound | KD082321_082521_B8 | 73656146 |  | Cal Std | C3H7NO2 | 0.077 | 3.906 | 958825247
Alanine pos | 9.49 | Target Compound | KD082321_082521_B9 | 38789563 |  | Cal Std | C3H7NO2 | 0.039 | 1.953 | 984026848 | Calibration Amount Diff -37.332% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_B10 | 20835175 |  | Cal Std | C3H7NO2 | 0.021 | 0.977 | 1011743942 | Calibration Amount Diff -100.814% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_B11 | 10555842 |  | Cal Std | C3H7NO2 | 0.011 | 0.488 | 996418691 | Calibration Amount Diff -235.712% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_B12 | 6602347 |  | Cal Std | C3H7NO2 | 0.007 | 0.244 | 976604541 | Calibration Amount Diff -474.223% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_C1 | 3311057883 |  | Cal Std | C3H7NO2 | 7.644 | 500 | 433129710
Alanine pos | 9.49 | Target Compound | KD082321_082521_C2 | 2176779576 |  | Cal Std | C3H7NO2 | 3.688 | 250 | 590231095
Alanine pos | 9.49 | Target Compound | KD082321_082521_C3 | 1426983061 |  | Cal Std | C3H7NO2 | 1.979 | 125 | 721158441
Alanine pos | 9.49 | Target Compound | KD082321_082521_C4 | 888410945 |  | Cal Std | C3H7NO2 | 1.042 | 62.5 | 852408426
Alanine pos | 9.49 | Target Compound | KD082321_082521_C5 | 490171347 |  | Cal Std | C3H7NO2 | 0.532 | 31.25 | 920817398
Alanine pos | 9.49 | Target Compound | KD082321_082521_C6 | 251015812 |  | Cal Std | C3H7NO2 | 0.266 | 15.625 | 942713557
Alanine pos | 9.49 | Target Compound | KD082321_082521_C7 | 122500366 |  | Cal Std | C3H7NO2 | 0.126 | 7.813 | 973357354
Alanine pos | 9.49 | Target Compound | KD082321_082521_C8 | 81270069 |  | Cal Std | C3H7NO2 | 0.082 | 3.906 | 997096995
Alanine pos | 9.49 | Target Compound | KD082321_082521_C9 | 38601862 |  | Cal Std | C3H7NO2 | 0.038 | 1.953 | 1009296296 | Calibration Amount Diff -41.262% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_C10 | 20864986 |  | Cal Std | C3H7NO2 | 0.021 | 0.977 | 1011746699 | Calibration Amount Diff -100.618% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_C11 | 11575454 |  | Cal Std | C3H7NO2 | 0.011 | 0.488 | 1022928364 | Calibration Amount Diff -226.028% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_C12 | 5473871 |  | Cal Std | C3H7NO2 | 0.006 | 0.244 | 975898977 | Calibration Amount Diff -505.102% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_A1 | 3335833270 |  | Cal Std | C3H7NO2 | 7.471 | 500 | 446478618
Alanine pos | 9.49 | Target Compound | KD082321_082521_A2 | 2095804522 |  | Cal Std | C3H7NO2 | 3.592 | 250 | 583413312
Alanine pos | 9.49 | Target Compound | KD082321_082521_A3 | 1426414009 |  | Cal Std | C3H7NO2 | 2.05 | 125 | 695643130
Alanine pos | 9.49 | Target Compound | KD082321_082521_A4 | 828417567 |  | Cal Std | C3H7NO2 | 1.023 | 62.5 | 809711464
Alanine pos | 9.49 | Target Compound | KD082321_082521_A5 | 441656999 |  | Cal Std | C3H7NO2 | 0.526 | 31.25 | 839660970
Alanine pos | 9.49 | Target Compound | KD082321_082521_A6 | 248259676 |  | Cal Std | C3H7NO2 | 0.271 | 15.625 | 917360560
Alanine pos | 9.49 | Target Compound | KD082321_082521_A7 | 131098016 |  | Cal Std | C3H7NO2 | 0.136 | 7.813 | 966805896
Alanine pos | 9.49 | Target Compound | KD082321_082521_A8 | 83824180 |  | Cal Std | C3H7NO2 | 0.08 | 3.906 | 1046611702
Alanine pos | 9.49 | Target Compound | KD082321_082521_A9 | 39873159 |  | Cal Std | C3H7NO2 | 0.04 | 1.953 | 998446188 | Calibration Amount Diff -35.603% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_A10 | 20194685 |  | Cal Std | C3H7NO2 | 0.021 | 0.977 | 952367480 | Calibration Amount Diff -96.72% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_A11 | 9720453 |  | Cal Std | C3H7NO2 | 0.01 | 0.488 | 970820655 | Calibration Amount Diff -243.505% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD082321_082521_A12 | 6473090 |  | Cal Std | C3H7NO2 | 0.007 | 0.244 | 980687717 | Calibration Amount Diff -478.513% (absolute) > max calibration amount diff 20%

## Alanine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B1 | 443277707 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B2 | 571146589 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B3 | 678905607 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B4 | 773548415 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B5 | 914968685 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B6 | 930240994 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B7 | 957287224 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B8 | 958825247 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B9 | 984026848 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B10 | 1011743942 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B11 | 996418691 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_B12 | 976604541 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C1 | 433129710 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C2 | 590231095 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C3 | 721158441 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C4 | 852408426 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C5 | 920817398 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C6 | 942713557 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C7 | 973357354 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C8 | 997096995 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C9 | 1009296296 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C10 | 1011746699 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C11 | 1022928364 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_C12 | 975898977 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A1 | 446478618 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A2 | 583413312 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A3 | 695643130 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A4 | 809711464 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A5 | 839660970 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A6 | 917360560 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A7 | 966805896 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A8 | 1046611702 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A9 | 998446188 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A10 | 952367480 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A11 | 970820655 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD082321_082521_A12 | 980687717 |  | Cal Std | [13]C3H7[15]NO2 | 0 | 1 | 0

## Arginine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Arginine pos | 14.24 | Target Compound | KD082321_082521_B1 | 36997535161 |  | Cal Std | C6H14N4O2 | 8.809 | 500 | 4200042144
Arginine pos | 14.24 | Target Compound | KD082321_082521_B2 | 22267813939 |  | Cal Std | C6H14N4O2 | 4.207 | 250 | 5292543623
Arginine pos | 14.24 | Target Compound | KD082321_082521_B3 | 14328477410 |  | Cal Std | C6H14N4O2 | 2.293 | 125 | 6247797296
Arginine pos | 14.24 | Target Compound | KD082321_082521_B4 | 9022613627 |  | Cal Std | C6H14N4O2 | 1.204 | 62.5 | 7491385886
Arginine pos | 14.24 | Target Compound | KD082321_082521_B5 | 5672527361 |  | Cal Std | C6H14N4O2 | 0.643 | 31.25 | 8821371231
Arginine pos | 14.24 | Target Compound | KD082321_082521_B6 | 2953800014 |  | Cal Std | C6H14N4O2 | 0.325 | 15.625 | 9102549878
Arginine pos | 14.24 | Target Compound | KD082321_082521_B7 | 1651444439 |  | Cal Std | C6H14N4O2 | 0.175 | 7.813 | 9455289810
Arginine pos | 14.24 | Target Compound | KD082321_082521_B8 | 865241559 |  | Cal Std | C6H14N4O2 | 0.094 | 3.906 | 9181846257 | Calibration Amount Diff -36.323% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_B9 | 455038452 |  | Cal Std | C6H14N4O2 | 0.049 | 1.953 | 9254571717 | Calibration Amount Diff -110.927% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_B10 | 226932621 |  | Cal Std | C6H14N4O2 | 0.026 | 0.977 | 8893553232 | Calibration Amount Diff -266.923% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_B11 | 124359797 |  | Cal Std | C6H14N4O2 | 0.013 | 0.488 | 9415777889 | Calibration Amount Diff -585.345% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_B12 | 62779923 |  | Cal Std | C6H14N4O2 | 0.007 | 0.244 | 9360253937 | Calibration Amount Diff -1230.346% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_C1 | 35372753296 |  | Cal Std | C6H14N4O2 | 8.421 | 500 | 4200704963
Arginine pos | 14.24 | Target Compound | KD082321_082521_C2 | 22810485396 |  | Cal Std | C6H14N4O2 | 4.146 | 250 | 5501876382
Arginine pos | 14.24 | Target Compound | KD082321_082521_C3 | 15590016150 |  | Cal Std | C6H14N4O2 | 2.247 | 125 | 6938969354
Arginine pos | 14.24 | Target Compound | KD082321_082521_C4 | 9314306535 |  | Cal Std | C6H14N4O2 | 1.206 | 62.5 | 7721012493
Arginine pos | 14.24 | Target Compound | KD082321_082521_C5 | 5401597812 |  | Cal Std | C6H14N4O2 | 0.639 | 31.25 | 8459149025
Arginine pos | 14.24 | Target Compound | KD082321_082521_C6 | 3063905504 |  | Cal Std | C6H14N4O2 | 0.333 | 15.625 | 9213936023
Arginine pos | 14.24 | Target Compound | KD082321_082521_C7 | 1628624498 |  | Cal Std | C6H14N4O2 | 0.164 | 7.813 | 9930093750
Arginine pos | 14.24 | Target Compound | KD082321_082521_C8 | 934656239 |  | Cal Std | C6H14N4O2 | 0.099 | 3.906 | 9441076398 | Calibration Amount Diff -29.012% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_C9 | 478751789 |  | Cal Std | C6H14N4O2 | 0.048 | 1.953 | 9886688255 | Calibration Amount Diff -113.213% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_C10 | 253677062 |  | Cal Std | C6H14N4O2 | 0.026 | 0.977 | 9864371849 | Calibration Amount Diff -265.697% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_C11 | 132792960 |  | Cal Std | C6H14N4O2 | 0.013 | 0.488 | 9935114757 | Calibration Amount Diff -583.399% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_C12 | 64712707 |  | Cal Std | C6H14N4O2 | 0.007 | 0.244 | 9580929958 | Calibration Amount Diff -1229.185% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_A1 | 33082379854 |  | Cal Std | C6H14N4O2 | 7.973 | 500 | 4149115205
Arginine pos | 14.24 | Target Compound | KD082321_082521_A2 | 20961971858 |  | Cal Std | C6H14N4O2 | 3.985 | 250 | 5259889539
Arginine pos | 14.24 | Target Compound | KD082321_082521_A3 | 14387227188 |  | Cal Std | C6H14N4O2 | 2.285 | 125 | 6296164561
Arginine pos | 14.24 | Target Compound | KD082321_082521_A4 | 8647903601 |  | Cal Std | C6H14N4O2 | 1.193 | 62.5 | 7247108417
Arginine pos | 14.24 | Target Compound | KD082321_082521_A5 | 5085391811 |  | Cal Std | C6H14N4O2 | 0.626 | 31.25 | 8127579933
Arginine pos | 14.24 | Target Compound | KD082321_082521_A6 | 2938577235 |  | Cal Std | C6H14N4O2 | 0.331 | 15.625 | 8891085351
Arginine pos | 14.24 | Target Compound | KD082321_082521_A7 | 1512634709 |  | Cal Std | C6H14N4O2 | 0.17 | 7.813 | 8895375765
Arginine pos | 14.24 | Target Compound | KD082321_082521_A8 | 902464216 |  | Cal Std | C6H14N4O2 | 0.096 | 3.906 | 9447484561 | Calibration Amount Diff -34.343% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_A9 | 438018544 |  | Cal Std | C6H14N4O2 | 0.049 | 1.953 | 8941972466 | Calibration Amount Diff -111.493% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_A10 | 236732868 |  | Cal Std | C6H14N4O2 | 0.026 | 0.977 | 9200345908 | Calibration Amount Diff -265.608% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_A11 | 119560199 |  | Cal Std | C6H14N4O2 | 0.013 | 0.488 | 9136271954 | Calibration Amount Diff -586.834% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD082321_082521_A12 | 64159128 |  | Cal Std | C6H14N4O2 | 0.007 | 0.244 | 9185172125 | Calibration Amount Diff -1223.518% (absolute) > max calibration amount diff 20%

## Arginine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B1 | 4200042144 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B2 | 5292543623 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B3 | 6247797296 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B4 | 7491385886 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B5 | 8821371231 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B6 | 9102549878 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B7 | 9455289810 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B8 | 9181846257 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B9 | 9254571717 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B10 | 8893553232 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B11 | 9415777889 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_B12 | 9360253937 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C1 | 4200704963 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C2 | 5501876382 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C3 | 6938969354 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C4 | 7721012493 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C5 | 8459149025 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C6 | 9213936023 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C7 | 9930093750 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C8 | 9441076398 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C9 | 9886688255 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C10 | 9864371849 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C11 | 9935114757 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_C12 | 9580929958 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A1 | 4149115205 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A2 | 5259889539 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A3 | 6296164561 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A4 | 7247108417 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A5 | 8127579933 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A6 | 8891085351 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A7 | 8895375765 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A8 | 9447484561 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A9 | 8941972466 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A10 | 9200345908 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A11 | 9136271954 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD082321_082521_A12 | 9185172125 |  | Cal Std | [13]C6H14[15]N4O2 | 0 | 1 | 0

## Asparagine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B1 | 1429363359 |  | Cal Std | C4H8N2O3 | 7.3021 | 434 | 195747333 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -23.517% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B2 | 835002755 |  | Cal Std | C4H8N2O3 | 3.2067 | 217 | 260391318 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -29.2% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B3 | 604740254 |  | Cal Std | C4H8N2O3 | 1.9206 | 108.5 | 314871915 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -24.975% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B4 | 337034396 |  | Cal Std | C4H8N2O3 | 0.9225 | 54.25 | 365348115 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -24.52% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B5 | 222429912 |  | Cal Std | C4H8N2O3 | 0.5937 | 27.125 | 374653013 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -24.322% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B6 | 112469403 |  | Cal Std | C4H8N2O3 | 0.26177 | 13.5625 | 429654333 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -29.144% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B7 | 47256801 |  | Cal Std | C4H8N2O3 | 0.10568 | 6.7812 | 447151307 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -31.868% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B8 | 23919897 |  | Cal Std | C4H8N2O3 | 0.047182 | 3.3906 | 506970262 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -38.002% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B9 | 10680419 |  | Cal Std | C4H8N2O3 | 0.021117 | 1.6953 | 505782347 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -71.717% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B10 | 3999784 |  | Cal Std | C4H8N2O3 | 0.0080163 | 0.84766 | 498955765 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -120.32% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B11 | 2331634 |  | Cal Std | C4H8N2O3 | 0.0045366 | 0.42383 | 513957916 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -229.95% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_B12 | 1019228 |  | Cal Std | C4H8N2O3 | 0.0019981 | 0.21191 | 510102844 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -444.497% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C1 | 2345766763 |  | Cal Std | C4H8N2O3 | 12.9966 | 868 | 180490158 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -49.084% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C2 | 1435411380 |  | Cal Std | C4H8N2O3 | 5.8001 | 434 | 247479066 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -51.88% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C3 | 916743482 |  | Cal Std | C4H8N2O3 | 3.0792 | 217 | 297719271 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -48.858% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C4 | 548818799 |  | Cal Std | C4H8N2O3 | 1.5928 | 108.5 | 344562776 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -48.483% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C5 | 282199648 |  | Cal Std | C4H8N2O3 | 0.71574 | 54.25 | 394274894 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -50.191% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C6 | 168887303 |  | Cal Std | C4H8N2O3 | 0.3867 | 27.125 | 436743619 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -52.427% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C7 | 81719055 |  | Cal Std | C4H8N2O3 | 0.18315 | 13.5625 | 446174432 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -61.143% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C8 | 42365214 |  | Cal Std | C4H8N2O3 | 0.089623 | 6.7812 | 472705621 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -64.516% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C9 | 19646312 |  | Cal Std | C4H8N2O3 | 0.040481 | 3.3906 | 485321766 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -99.287% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C10 | 8249078 |  | Cal Std | C4H8N2O3 | 0.015735 | 1.6953 | 524255717 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -151.302% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C11 | 3213203 |  | Cal Std | C4H8N2O3 | 0.0064504 | 0.84766 | 498137318 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -250.423% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_C12 | 1992942 |  | Cal Std | C4H8N2O3 | 0.0040121 | 0.42383 | 496732109 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -476.209% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A1 | 1991883536 |  | Cal Std | C4H8N2O3 | 8.9406 | 434 | 222791565 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 75.04% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A2 | 1243466581 |  | Cal Std | C4H8N2O3 | 4.8292 | 217 | 257489316 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 66.795% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A3 | 755651314 |  | Cal Std | C4H8N2O3 | 2.445 | 108.5 | 309056150 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 86.341% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A4 | 427212676 |  | Cal Std | C4H8N2O3 | 1.1697 | 54.25 | 365235104 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 91.042% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A5 | 248387133 |  | Cal Std | C4H8N2O3 | 0.61513 | 27.125 | 403796403 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 89.844% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A6 | 117817274 |  | Cal Std | C4H8N2O3 | 0.27328 | 13.5625 | 431128005 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 94.927% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A7 | 65909054 |  | Cal Std | C4H8N2O3 | 0.14327 | 6.7812 | 460017897 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 88.946% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A8 | 28836397 |  | Cal Std | C4H8N2O3 | 0.060715 | 3.3906 | 474949928 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 104.268% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A9 | 13705003 |  | Cal Std | C4H8N2O3 | 0.025851 | 1.6953 | 530150111 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 76.328% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A10 | 4384031 |  | Cal Std | C4H8N2O3 | 0.008775 | 0.84766 | 499604308 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff 30.71% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A11 | 2507860 |  | Cal Std | C4H8N2O3 | 0.0046273 | 0.42383 | 541971636 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -100.489% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD082321_082521_A12 | 1431542 |  | Cal Std | C4H8N2O3 | 0.0028441 | 0.21191 | 503346271 | Rsquared 0.723 < R2Threshold 0.99. Calibration Amount Diff -300.693% (absolute) > max calibration amount diff 20%

## Asparagine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B1 | 645805578 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B2 | 797282437 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B3 | 899627897 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B4 | 973215977 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B5 | 1114934878 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B6 | 1139123322 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B7 | 1155892756 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B8 | 1120568928 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B9 | 1138448468 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B10 | 1171846610 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B11 | 1172666108 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_B12 | 1186582167 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C1 | 693412869 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C2 | 836000502 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C3 | 942350073 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C4 | 1017253801 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C5 | 1040713458 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C6 | 1098530700 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C7 | 1146092729 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C8 | 1174119647 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C9 | 1188432540 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C10 | 1166787917 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C11 | 1134646493 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_C12 | 1189424648 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A1 | 530353897 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A2 | 647773728 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A3 | 767291599 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A4 | 882658291 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A5 | 963386347 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A6 | 1032150859 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A7 | 1110494801 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A8 | 1219435350 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A9 | 1147403318 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A10 | 1162470381 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A11 | 1127607843 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD082321_082521_A12 | 1180753428 |  | Cal Std | [13]C4H8[15]N2O3 | 0 | 1 | 0

## Aspartate neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B1 | 6471418039 |  | Cal Std | C4H7NO4 | 8.539 | 8.5389 | 500 | 757876361
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B2 | 3606881552 |  | Cal Std | C4H7NO4 | 4.042 | 4.0417 | 250 | 892414404
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B3 | 2337288208 |  | Cal Std | C4H7NO4 | 2.169 | 2.1691 | 125 | 1077560170
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B4 | 1362070692 |  | Cal Std | C4H7NO4 | 1.127 | 1.1267 | 62.5 | 1208872686
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B5 | 908017521 |  | Cal Std | C4H7NO4 | 0.595 | 0.59495 | 31.25 | 1526196280
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B6 | 468822068 |  | Cal Std | C4H7NO4 | 0.301 | 0.30093 | 15.625 | 1557918233
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B7 | 252299285 |  | Cal Std | C4H7NO4 | 0.162 | 0.16166 | 7.813 | 1560678737
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B8 | 137251138 |  | Cal Std | C4H7NO4 | 0.087 | 0.087136 | 3.906 | 1575131251
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B9 | 72743063 |  | Cal Std | C4H7NO4 | 0.045 | 0.044523 | 1.953 | 1633838373 | Calibration Amount Diff -60.568% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B10 | 39220787 |  | Cal Std | C4H7NO4 | 0.023 | 0.02282 | 0.977 | 1718714597 | Calibration Amount Diff -157.049% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B11 | 18001051 |  | Cal Std | C4H7NO4 | 0.011 | 0.011218 | 0.488 | 1604695678 | Calibration Amount Diff -359.633% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_B12 | 9567673 |  | Cal Std | C4H7NO4 | 0.006 | 0.0057615 | 0.244 | 1660631059 | Calibration Amount Diff -756.041% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C1 | 6345840290 |  | Cal Std | C4H7NO4 | 8.317 | 8.317 | 500 | 762994750
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C2 | 3686909317 |  | Cal Std | C4H7NO4 | 4.006 | 4.0064 | 250 | 920264112
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C3 | 2419559580 |  | Cal Std | C4H7NO4 | 2.143 | 2.1426 | 125 | 1129274079
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C4 | 1532414828 |  | Cal Std | C4H7NO4 | 1.129 | 1.1287 | 62.5 | 1357684424
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C5 | 890916627 |  | Cal Std | C4H7NO4 | 0.592 | 0.59221 | 31.25 | 1504394457
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C6 | 479472020 |  | Cal Std | C4H7NO4 | 0.304 | 0.30398 | 15.625 | 1577290529
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C7 | 245500445 |  | Cal Std | C4H7NO4 | 0.149 | 0.14906 | 7.813 | 1646997600
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C8 | 150246812 |  | Cal Std | C4H7NO4 | 0.09 | 0.089925 | 3.906 | 1670805206
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C9 | 75511142 |  | Cal Std | C4H7NO4 | 0.044 | 0.044175 | 1.953 | 1709363054 | Calibration Amount Diff -61.658% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C10 | 38472402 |  | Cal Std | C4H7NO4 | 0.023 | 0.023383 | 0.977 | 1645348912 | Calibration Amount Diff -153.526% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C11 | 20617262 |  | Cal Std | C4H7NO4 | 0.012 | 0.011587 | 0.488 | 1779328762 | Calibration Amount Diff -355.003% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_C12 | 10200737 |  | Cal Std | C4H7NO4 | 0.006 | 0.0060298 | 0.244 | 1691728439 | Calibration Amount Diff -749.316% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A1 | 5789822896 |  | Cal Std | C4H7NO4 | 7.893 | 7.8926 | 500 | 733571775
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A2 | 3320633377 |  | Cal Std | C4H7NO4 | 3.844 | 3.8444 | 250 | 863748951
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A3 | 2351191630 |  | Cal Std | C4H7NO4 | 2.181 | 2.181 | 125 | 1078034957
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A4 | 1400400151 |  | Cal Std | C4H7NO4 | 1.129 | 1.129 | 62.5 | 1240377354
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A5 | 842794682 |  | Cal Std | C4H7NO4 | 0.581 | 0.5812 | 31.25 | 1450105307
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A6 | 455309278 |  | Cal Std | C4H7NO4 | 0.305 | 0.305 | 15.625 | 1492840184
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A7 | 224274528 |  | Cal Std | C4H7NO4 | 0.156 | 0.15582 | 7.813 | 1439325615
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A8 | 145162817 |  | Cal Std | C4H7NO4 | 0.088 | 0.088381 | 3.906 | 1642466717
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A9 | 71809580 |  | Cal Std | C4H7NO4 | 0.045 | 0.044972 | 1.953 | 1596757631 | Calibration Amount Diff -59.161% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A10 | 36463960 |  | Cal Std | C4H7NO4 | 0.024 | 0.024172 | 0.977 | 1508527791 | Calibration Amount Diff -148.584% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A11 | 19553599 |  | Cal Std | C4H7NO4 | 0.012 | 0.012166 | 0.488 | 1607202545 | Calibration Amount Diff -347.745% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD082321_082521_A12 | 9345286 |  | Cal Std | C4H7NO4 | 0.006 | 0.006154 | 0.244 | 1518575387 | Calibration Amount Diff -746.202% (absolute) > max calibration amount diff 20%

## Aspartate U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B1 | 757876361 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B2 | 892414404 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B3 | 1077560170 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B4 | 1208872686 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B5 | 1526196280 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B6 | 1557918233 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B7 | 1560678737 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B8 | 1575131251 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B9 | 1633838373 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B10 | 1718714597 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B11 | 1604695678 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_B12 | 1660631059 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C1 | 762994750 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C2 | 920264112 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C3 | 1129274079 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C4 | 1357684424 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C5 | 1504394457 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C6 | 1577290529 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C7 | 1646997600 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C8 | 1670805206 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C9 | 1709363054 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C10 | 1645348912 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C11 | 1779328762 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_C12 | 1691728439 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A1 | 733571775 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A2 | 863748951 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A3 | 1078034957 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A4 | 1240377354 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A5 | 1450105307 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A6 | 1492840184 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A7 | 1439325615 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A8 | 1642466717 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A9 | 1596757631 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A10 | 1508527791 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A11 | 1607202545 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD082321_082521_A12 | 1518575387 |  | Cal Std | [13]C4H7[15]NO4 | 0 | 1 | 0

## Cystine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Cystine pos | 10.6 | Target Compound | KD082321_082521_B1 | 2510417837 |  | Cal Std | C6H12N2O4S2 | 19.939 | 500 | 125902853
Cystine pos | 10.6 | Target Compound | KD082321_082521_B2 | 4425251409 |  | Cal Std | C6H12N2O4S2 | 8.455 | 250 | 523391356
Cystine pos | 10.6 | Target Compound | KD082321_082521_B3 | 2804943194 |  | Cal Std | C6H12N2O4S2 | 4.574 | 125 | 613250555
Cystine pos | 10.6 | Target Compound | KD082321_082521_B4 | 1558718744 |  | Cal Std | C6H12N2O4S2 | 2.354 | 62.5 | 662131795
Cystine pos | 10.6 | Target Compound | KD082321_082521_B5 | 899623692 |  | Cal Std | C6H12N2O4S2 | 1.234 | 31.25 | 728917490
Cystine pos | 10.6 | Target Compound | KD082321_082521_B6 | 424200672 |  | Cal Std | C6H12N2O4S2 | 0.624 | 15.625 | 679454187
Cystine pos | 10.6 | Target Compound | KD082321_082521_B7 | 206099728 |  | Cal Std | C6H12N2O4S2 | 0.329 | 7.813 | 625586261 | Calibration Amount Diff 22.517% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_B8 | 107697425 |  | Cal Std | C6H12N2O4S2 | 0.175 | 3.906 | 615197460 | Calibration Amount Diff 38.727% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_B9 | 54150872 |  | Cal Std | C6H12N2O4S2 | 0.088 | 1.953 | 617582179 | Calibration Amount Diff 57.083% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_B10 | 25988299 |  | Cal Std | C6H12N2O4S2 | 0.044 | 0.977 | 597037126 | Calibration Amount Diff 92.421% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_B11 | 12813293 |  | Cal Std | C6H12N2O4S2 | 0.021 | 0.488 | 606104418 | Calibration Amount Diff 161.809% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_B12 | 5598014 |  | Cal Std | C6H12N2O4S2 | 0.009 | 0.244 | 602296068 | Calibration Amount Diff 293.004% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_C1 | 2185095786 |  | Cal Std | C6H12N2O4S2 | 19.176 | 500 | 113950114
Cystine pos | 10.6 | Target Compound | KD082321_082521_C2 | 2970888602 |  | Cal Std | C6H12N2O4S2 | 8.546 | 250 | 347630682
Cystine pos | 10.6 | Target Compound | KD082321_082521_C3 | 2881513678 |  | Cal Std | C6H12N2O4S2 | 4.602 | 125 | 626102752
Cystine pos | 10.6 | Target Compound | KD082321_082521_C4 | 1617269018 |  | Cal Std | C6H12N2O4S2 | 2.362 | 62.5 | 684773401
Cystine pos | 10.6 | Target Compound | KD082321_082521_C5 | 770732652 |  | Cal Std | C6H12N2O4S2 | 1.238 | 31.25 | 622446505
Cystine pos | 10.6 | Target Compound | KD082321_082521_C6 | 397863889 |  | Cal Std | C6H12N2O4S2 | 0.637 | 15.625 | 624190470
Cystine pos | 10.6 | Target Compound | KD082321_082521_C7 | 196136681 |  | Cal Std | C6H12N2O4S2 | 0.304 | 7.813 | 645377696
Cystine pos | 10.6 | Target Compound | KD082321_082521_C8 | 115745272 |  | Cal Std | C6H12N2O4S2 | 0.18 | 3.906 | 644333600 | Calibration Amount Diff 41.877% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_C9 | 54920635 |  | Cal Std | C6H12N2O4S2 | 0.089 | 1.953 | 616958381 | Calibration Amount Diff 58.924% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_C10 | 27014837 |  | Cal Std | C6H12N2O4S2 | 0.044 | 0.977 | 608599987 | Calibration Amount Diff 94.789% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_C11 | 13160015 |  | Cal Std | C6H12N2O4S2 | 0.021 | 0.488 | 618427811 | Calibration Amount Diff 162.578% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_C12 | 5402472 |  | Cal Std | C6H12N2O4S2 | 0.009 | 0.244 | 600840938 | Calibration Amount Diff 289.664% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_A1 | 3647210807 |  | Cal Std | C6H12N2O4S2 | 17.847 | 500 | 204354842
Cystine pos | 10.6 | Target Compound | KD082321_082521_A2 | 4053943190 |  | Cal Std | C6H12N2O4S2 | 8.169 | 250 | 496240383
Cystine pos | 10.6 | Target Compound | KD082321_082521_A3 | 2826733005 |  | Cal Std | C6H12N2O4S2 | 4.568 | 125 | 618785591
Cystine pos | 10.6 | Target Compound | KD082321_082521_A4 | 1429137935 |  | Cal Std | C6H12N2O4S2 | 2.372 | 62.5 | 602622741
Cystine pos | 10.6 | Target Compound | KD082321_082521_A5 | 768072220 |  | Cal Std | C6H12N2O4S2 | 1.235 | 31.25 | 621827281
Cystine pos | 10.6 | Target Compound | KD082321_082521_A6 | 383653447 |  | Cal Std | C6H12N2O4S2 | 0.632 | 15.625 | 606618870
Cystine pos | 10.6 | Target Compound | KD082321_082521_A7 | 192684977 |  | Cal Std | C6H12N2O4S2 | 0.316 | 7.813 | 610515876
Cystine pos | 10.6 | Target Compound | KD082321_082521_A8 | 118773355 |  | Cal Std | C6H12N2O4S2 | 0.177 | 3.906 | 669373478 | Calibration Amount Diff 40.365% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_A9 | 54355731 |  | Cal Std | C6H12N2O4S2 | 0.089 | 1.953 | 613725636 | Calibration Amount Diff 58.302% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_A10 | 25832860 |  | Cal Std | C6H12N2O4S2 | 0.043 | 0.977 | 597394795 | Calibration Amount Diff 91.633% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_A11 | 11557345 |  | Cal Std | C6H12N2O4S2 | 0.02 | 0.488 | 580324076 | Calibration Amount Diff 155.055% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD082321_082521_A12 | 4898907 |  | Cal Std | C6H12N2O4S2 | 0.008 | 0.244 | 591240083 | Calibration Amount Diff 281.883% (absolute) > max calibration amount diff 20%

## Cystine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B1 | 125902853 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0 | 10.752:Peak area 125902853.263 is out of bounds (ISTD Minimum Recovery 285604274.104 and ISTD Max Recovery 856812822.311)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B2 | 523391356 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B3 | 613250555 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B4 | 662131795 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B5 | 728917490 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B6 | 679454187 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B7 | 625586261 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B8 | 615197460 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B9 | 617582179 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B10 | 597037126 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B11 | 606104418 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_B12 | 602296068 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C1 | 113950114 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0 | 10.755:Peak area 113950113.538 is out of bounds (ISTD Minimum Recovery 285604274.104 and ISTD Max Recovery 856812822.311)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C2 | 347630682 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C3 | 626102752 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C4 | 684773401 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C5 | 622446505 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C6 | 624190470 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C7 | 645377696 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C8 | 644333600 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C9 | 616958381 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C10 | 608599987 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C11 | 618427811 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_C12 | 600840938 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A1 | 204354842 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0 | 10.779:Peak area 204354841.619 is out of bounds (ISTD Minimum Recovery 285604274.104 and ISTD Max Recovery 856812822.311)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A2 | 496240383 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A3 | 618785591 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A4 | 602622741 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A5 | 621827281 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A6 | 606618870 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A7 | 610515876 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A8 | 669373478 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A9 | 613725636 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A10 | 597394795 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A11 | 580324076 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD082321_082521_A12 | 591240083 |  | Cal Std | [13]C6H12[15]N2O4S2 | 0 | 1 | 0

## Glutamate neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B1 | 6511317564 |  | Cal Std | C5H9NO4 | 8.536 | 500 | 762800115
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B2 | 3631035569 |  | Cal Std | C5H9NO4 | 3.995 | 250 | 908880601
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B3 | 2219332622 |  | Cal Std | C5H9NO4 | 2.151 | 125 | 1031963498
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B4 | 1193178305 |  | Cal Std | C5H9NO4 | 1.12 | 62.5 | 1065678259
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B5 | 707667278 |  | Cal Std | C5H9NO4 | 0.587 | 31.25 | 1206430544
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B6 | 357675517 |  | Cal Std | C5H9NO4 | 0.298 | 15.625 | 1198665688
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B7 | 186447450 |  | Cal Std | C5H9NO4 | 0.157 | 7.813 | 1184020945
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B8 | 104998543 |  | Cal Std | C5H9NO4 | 0.086 | 3.906 | 1225937130
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B9 | 55585608 |  | Cal Std | C5H9NO4 | 0.044 | 1.953 | 1268925336 | Calibration Amount Diff -71.331% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B10 | 27771204 |  | Cal Std | C5H9NO4 | 0.022 | 0.977 | 1251855507 | Calibration Amount Diff -179.919% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B11 | 13760081 |  | Cal Std | C5H9NO4 | 0.011 | 0.488 | 1236375053 | Calibration Amount Diff -400.47% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_B12 | 6059905 |  | Cal Std | C5H9NO4 | 0.005 | 0.244 | 1234069196 | Calibration Amount Diff -858.982% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C1 | 6398720446 |  | Cal Std | C5H9NO4 | 8.101 | 500 | 789910828
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C2 | 3750738200 |  | Cal Std | C5H9NO4 | 3.96 | 250 | 947221881
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C3 | 2297010702 |  | Cal Std | C5H9NO4 | 2.158 | 125 | 1064516365
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C4 | 1281861000 |  | Cal Std | C5H9NO4 | 1.122 | 62.5 | 1142184918
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C5 | 697570506 |  | Cal Std | C5H9NO4 | 0.588 | 31.25 | 1185654374
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C6 | 370791888 |  | Cal Std | C5H9NO4 | 0.304 | 15.625 | 1217926175
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C7 | 178050391 |  | Cal Std | C5H9NO4 | 0.145 | 7.813 | 1231200179
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C8 | 118341497 |  | Cal Std | C5H9NO4 | 0.09 | 3.906 | 1319677348
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C9 | 54844322 |  | Cal Std | C5H9NO4 | 0.042 | 1.953 | 1296754932 | Calibration Amount Diff -76.131% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C10 | 28315135 |  | Cal Std | C5H9NO4 | 0.022 | 0.977 | 1268275972 | Calibration Amount Diff -179.02% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C11 | 13487848 |  | Cal Std | C5H9NO4 | 0.01 | 0.488 | 1301469496 | Calibration Amount Diff -410.201% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_C12 | 5504193 |  | Cal Std | C5H9NO4 | 0.005 | 0.244 | 1214475409 | Calibration Amount Diff -868.597% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A1 | 6241256176 |  | Cal Std | C5H9NO4 | 7.744 | 500 | 805925064
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A2 | 3404841218 |  | Cal Std | C5H9NO4 | 3.844 | 250 | 885682753
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A3 | 2164885521 |  | Cal Std | C5H9NO4 | 2.146 | 125 | 1008991057
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A4 | 1215477818 |  | Cal Std | C5H9NO4 | 1.124 | 62.5 | 1081452485
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A5 | 655967806 |  | Cal Std | C5H9NO4 | 0.576 | 31.25 | 1138356519
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A6 | 367605750 |  | Cal Std | C5H9NO4 | 0.302 | 15.625 | 1216916122
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A7 | 182871565 |  | Cal Std | C5H9NO4 | 0.154 | 7.813 | 1189776288
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A8 | 109987438 |  | Cal Std | C5H9NO4 | 0.087 | 3.906 | 1267052602
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A9 | 55821932 |  | Cal Std | C5H9NO4 | 0.046 | 1.953 | 1224188523 | Calibration Amount Diff -65.636% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A10 | 27371744 |  | Cal Std | C5H9NO4 | 0.023 | 0.977 | 1185439056 | Calibration Amount Diff -174.169% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A11 | 13900023 |  | Cal Std | C5H9NO4 | 0.011 | 0.488 | 1248334384 | Calibration Amount Diff -400.4% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD082321_082521_A12 | 6542068 |  | Cal Std | C5H9NO4 | 0.005 | 0.244 | 1216153509 | Calibration Amount Diff -847.068% (absolute) > max calibration amount diff 20%

## Glutamate U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B1 | 762800115 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B2 | 908880601 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B3 | 1031963498 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B4 | 1065678259 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B5 | 1206430544 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B6 | 1198665688 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B7 | 1184020945 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B8 | 1225937130 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B9 | 1268925336 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B10 | 1251855507 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B11 | 1236375053 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_B12 | 1234069196 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C1 | 789910828 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C2 | 947221881 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C3 | 1064516365 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C4 | 1142184918 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C5 | 1185654374 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C6 | 1217926175 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C7 | 1231200179 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C8 | 1319677348 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C9 | 1296754932 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C10 | 1268275972 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C11 | 1301469496 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_C12 | 1214475409 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A1 | 805925064 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A2 | 885682753 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A3 | 1008991057 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A4 | 1081452485 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A5 | 1138356519 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A6 | 1216916122 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A7 | 1189776288 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A8 | 1267052602 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A9 | 1224188523 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A10 | 1185439056 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A11 | 1248334384 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD082321_082521_A12 | 1216153509 |  | Cal Std | [13]C5H9[15]NO4 | 0 | 1 | 0

## Glutamine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B1 | 9759949987 |  | Cal Std | C5H10N2O3 | 6.3363 | 434 | 1540320659 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -76.921% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B2 | 6503979990 |  | Cal Std | C5H10N2O3 | 3.0192 | 217 | 2154226812 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -79.268% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B3 | 4756187457 |  | Cal Std | C5H10N2O3 | 1.8637 | 108.5 | 2552027562 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -79.017% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B4 | 2815090907 |  | Cal Std | C5H10N2O3 | 0.94137 | 54.25 | 2990411855 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -80.56% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B5 | 1861707230 |  | Cal Std | C5H10N2O3 | 0.62142 | 27.125 | 2995894522 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -84.435% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B6 | 966156966 |  | Cal Std | C5H10N2O3 | 0.28676 | 13.5625 | 3369178088 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -94.63% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B7 | 444498438 |  | Cal Std | C5H10N2O3 | 0.12952 | 6.7812 | 3432009308 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -113.005% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B8 | 234824534 |  | Cal Std | C5H10N2O3 | 0.064094 | 3.3906 | 3663757007 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -152.614% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B9 | 131396337 |  | Cal Std | C5H10N2O3 | 0.035436 | 1.6953 | 3707968263 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -232.496% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B10 | 56792269 |  | Cal Std | C5H10N2O3 | 0.015872 | 0.84766 | 3578094435 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -395.543% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B11 | 31921390 |  | Cal Std | C5H10N2O3 | 0.0087648 | 0.42383 | 3642009208 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -714.896% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_B12 | 13851501 |  | Cal Std | C5H10N2O3 | 0.0036597 | 0.21191 | 3784889414 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -1358.986% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C1 | 16666991916 |  | Cal Std | C5H10N2O3 | 12.7879 | 868 | 1303340185 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -95.806% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C2 | 11748563972 |  | Cal Std | C5H10N2O3 | 5.8452 | 434 | 2009943876 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -96.68% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C3 | 8045181041 |  | Cal Std | C5H10N2O3 | 3.3107 | 217 | 2430064596 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -97.542% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C4 | 4992039526 |  | Cal Std | C5H10N2O3 | 1.8016 | 108.5 | 2770891557 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -99.797% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C5 | 2617488049 |  | Cal Std | C5H10N2O3 | 0.84983 | 54.25 | 3079997681 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -104.007% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C6 | 1603258248 |  | Cal Std | C5H10N2O3 | 0.47104 | 27.125 | 3403627578 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -114.892% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C7 | 838356142 |  | Cal Std | C5H10N2O3 | 0.24082 | 13.5625 | 3481223643 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -135.043% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C8 | 457587333 |  | Cal Std | C5H10N2O3 | 0.12947 | 6.7812 | 3534442295 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -175.037% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C9 | 239950472 |  | Cal Std | C5H10N2O3 | 0.06745 | 3.3906 | 3557439817 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -254.517% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C10 | 112450276 |  | Cal Std | C5H10N2O3 | 0.030868 | 1.6953 | 3642888299 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -415.074% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C11 | 58526679 |  | Cal Std | C5H10N2O3 | 0.016886 | 0.84766 | 3466052470 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -733.635% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_C12 | 24847765 |  | Cal Std | C5H10N2O3 | 0.0073213 | 0.42383 | 3393888373 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -1364.106% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A1 | 9577020619 |  | Cal Std | C5H10N2O3 | 5.863 | 434 | 1633466978 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 172.321% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A2 | 6730875502 |  | Cal Std | C5H10N2O3 | 3.2799 | 217 | 2052188458 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 169.381% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A3 | 4270419392 |  | Cal Std | C5H10N2O3 | 1.7175 | 108.5 | 2486439676 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 200.087% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A4 | 2430994815 |  | Cal Std | C5H10N2O3 | 0.87859 | 54.25 | 2766917688 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 211.807% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A5 | 1415283401 |  | Cal Std | C5H10N2O3 | 0.46727 | 27.125 | 3028841195 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 215.521% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A6 | 695794575 |  | Cal Std | C5H10N2O3 | 0.21837 | 13.5625 | 3186345452 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 218.884% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A7 | 415738035 |  | Cal Std | C5H10N2O3 | 0.12081 | 6.7812 | 3441277804 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 207.352% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A8 | 195242014 |  | Cal Std | C5H10N2O3 | 0.057044 | 3.3906 | 3422630562 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 207.024% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A9 | 108318924 |  | Cal Std | C5H10N2O3 | 0.029466 | 1.6953 | 3676046453 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff 132.625% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A10 | 44407295 |  | Cal Std | C5H10N2O3 | 0.01264 | 0.84766 | 3513285586 | Rsquared 0.331 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A11 | 25451955 |  | Cal Std | C5H10N2O3 | 0.0067604 | 0.42383 | 3764856051 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -341.771% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD082321_082521_A12 | 9438004 |  | Cal Std | C5H10N2O3 | 0.0028037 | 0.21191 | 3366291039 | Rsquared 0.331 < R2Threshold 0.99. Calibration Amount Diff -952.136% (absolute) > max calibration amount diff 20%

## Glutamine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B1 | 5072096375 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B2 | 6458141831 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B3 | 7049859501 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B4 | 7484856176 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B5 | 8044823744 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B6 | 8275049764 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B7 | 8180178056 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B8 | 7908693578 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B9 | 7889544264 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B10 | 8090264111 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B11 | 7834938983 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_B12 | 8384536196 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C1 | 5495172703 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C2 | 6641633381 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C3 | 7836825800 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C4 | 7969427232 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C5 | 8001599961 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C6 | 8187684152 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C7 | 8202372814 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C8 | 8673141027 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C9 | 8349826555 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C10 | 8362680866 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C11 | 8683401146 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_C12 | 8248925690 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A1 | 3310798983 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0 | 9.757:Peak area 3310798983.009 is out of bounds (ISTD Minimum Recovery 3723531500.431 and ISTD Max Recovery 11170594501.293)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A2 | 4351150710 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A3 | 5451036985 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A4 | 6426567185 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A5 | 6855858885 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A6 | 7726730615 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A7 | 7878975506 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A8 | 8620052812 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A9 | 8168950533 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A10 | 7930277170 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A11 | 8073217440 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD082321_082521_A12 | 7974977301 |  | Cal Std | [13]C5H10[15]N2O3 | 0 | 1 | 0

## Glycine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glycine neg | 10.11 | Target Compound | KD082321_082521_B1 | 320139455 |  | Cal Std | C2H5NO2 | 10.516 | 500 | 30444242 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_B2 | 147894426 |  | Cal Std | C2H5NO2 | 4.159 | 250 | 35557888 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_B3 | 79818752 |  | Cal Std | C2H5NO2 | 2.195 | 125 | 36363048 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_B4 | 38492450 |  | Cal Std | C2H5NO2 | 1.098 | 62.5 | 35068932 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_B5 | 18619611 |  | Cal Std | C2H5NO2 | 0.504 | 31.25 | 36907603 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_B6 | 9584714 |  | Cal Std | C2H5NO2 | 0.26 | 15.625 | 36873035 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_B7 | 3795686 |  | Cal Std | C2H5NO2 | 0.11 | 7.813 | 34553509 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 21.819% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_B8 | 1792914 |  | Cal Std | C2H5NO2 | 0.054 | 3.906 | 33250432 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 67.933% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_B9 | 537334 |  | Cal Std | C2H5NO2 | 0.015 | 1.953 | 35560011 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 130.754% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_B10 | 541605 |  | Cal Std | C2H5NO2 | 0.016 | 0.977 | 34630442 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 364.136% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_B11 | 793650 |  | Cal Std | C2H5NO2 | 0.024 | 0.488 | 33155206 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 919.163% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_B12 | 1723063 |  | Cal Std | C2H5NO2 | 0.048 | 0.244 | 35854369 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 2461.192% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_C1 | 312497195 |  | Cal Std | C2H5NO2 | 9.882 | 500 | 31622994 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_C2 | 158804798 |  | Cal Std | C2H5NO2 | 4.007 | 250 | 39636260 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_C3 | 92185165 |  | Cal Std | C2H5NO2 | 2.199 | 125 | 41928822 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_C4 | 47562066 |  | Cal Std | C2H5NO2 | 1.117 | 62.5 | 42580093 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_C5 | 23395137 |  | Cal Std | C2H5NO2 | 0.553 | 31.25 | 42336681 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_C6 | 11056338 |  | Cal Std | C2H5NO2 | 0.276 | 15.625 | 40001066 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_C7 | 6979544 |  | Cal Std | C2H5NO2 | 0.189 | 7.813 | 36987134 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 75.202% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_C8 | 3752265 |  | Cal Std | C2H5NO2 | 0.093 | 3.906 | 40559229 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 120.193% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_C9 | 1796087 |  | Cal Std | C2H5NO2 | 0.048 | 1.953 | 37499534 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 219.549% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_C10 | 1674473 |  | Cal Std | C2H5NO2 | 0.043 | 0.977 | 38517441 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 514.824% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_C11 | 766270 |  | Cal Std | C2H5NO2 | 0.02 | 0.488 | 38781738 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 873.868% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_C12 | 0 |  | Cal Std | C2H5NO2 | 0 | 0.244 | 36550620 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_A1 | 280854345 |  | Cal Std | C2H5NO2 | 8.776 | 500 | 32002067 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_A2 | 131664810 |  | Cal Std | C2H5NO2 | 3.902 | 250 | 33739465 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_A3 | 72847318 |  | Cal Std | C2H5NO2 | 2.19 | 125 | 33258602 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_A4 | 37572722 |  | Cal Std | C2H5NO2 | 1.054 | 62.5 | 35654121 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_A5 | 17370978 |  | Cal Std | C2H5NO2 | 0.492 | 31.25 | 35284842 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_A6 | 7689085 |  | Cal Std | C2H5NO2 | 0.211 | 15.625 | 36505496 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_A7 | 3767773 |  | Cal Std | C2H5NO2 | 0.109 | 7.813 | 34649893 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 21.066% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_A8 | 3281769 |  | Cal Std | C2H5NO2 | 0.082 | 3.906 | 40145842 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 105.612% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_A9 | 867217 |  | Cal Std | C2H5NO2 | 0.025 | 1.953 | 34878626 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 157.169% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_A10 | 905760 |  | Cal Std | C2H5NO2 | 0.026 | 0.977 | 34626731 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 421.081% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD082321_082521_A11 | 0 |  | Cal Std | C2H5NO2 | 0 | 0.488 | 35652430 | Rsquared 0.988 < R2Threshold 0.99
Glycine neg | 10.11 | Target Compound | KD082321_082521_A12 | 536479 |  | Cal Std | C2H5NO2 | 0.016 | 0.244 | 34412987 | Rsquared 0.988 < R2Threshold 0.99. Calibration Amount Diff 1757.361% (absolute) > max calibration amount diff 20%

## Glycine U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B1 | 30444242 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B2 | 35557888 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B3 | 36363048 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B4 | 35068932 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B5 | 36907603 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B6 | 36873035 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B7 | 34553509 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B8 | 33250432 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B9 | 35560011 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B10 | 34630442 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B11 | 33155206 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_B12 | 35854369 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C1 | 31622994 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C2 | 39636260 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C3 | 41928822 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C4 | 42580093 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C5 | 42336681 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C6 | 40001066 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C7 | 36987134 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C8 | 40559229 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C9 | 37499534 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C10 | 38517441 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C11 | 38781738 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_C12 | 36550620 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A1 | 32002067 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A2 | 33739465 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A3 | 33258602 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A4 | 35654121 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A5 | 35284842 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A6 | 36505496 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A7 | 34649893 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A8 | 40145842 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A9 | 34878626 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A10 | 34626731 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A11 | 35652430 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD082321_082521_A12 | 34412987 |  | Cal Std | [13]C2H5[15]NO2 | 0 | 1 | 0

## Isoleucine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B1 | 38743518466 |  | Cal Std | C6H13NO2 | 8.112 | 500 | 4775834017
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B2 | 24624888187 |  | Cal Std | C6H13NO2 | 3.888 | 250 | 6333966369
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B3 | 15116128076 |  | Cal Std | C6H13NO2 | 2.089 | 125 | 7237088876
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B4 | 8628275553 |  | Cal Std | C6H13NO2 | 1.093 | 62.5 | 7891520819
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B5 | 4665064002 |  | Cal Std | C6H13NO2 | 0.556 | 31.25 | 8384278182
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B6 | 2494089189 |  | Cal Std | C6H13NO2 | 0.28 | 15.625 | 8907463964
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B7 | 1333694432 |  | Cal Std | C6H13NO2 | 0.152 | 7.813 | 8799278780
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B8 | 672244046 |  | Cal Std | C6H13NO2 | 0.081 | 3.906 | 8268872721 | Calibration Amount Diff -23.296% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B9 | 355544107 |  | Cal Std | C6H13NO2 | 0.041 | 1.953 | 8696481935 | Calibration Amount Diff -80.746% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B10 | 188234955 |  | Cal Std | C6H13NO2 | 0.021 | 0.977 | 8840685952 | Calibration Amount Diff -191.514% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B11 | 93389198 |  | Cal Std | C6H13NO2 | 0.011 | 0.488 | 8632972253 | Calibration Amount Diff -422.36% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_B12 | 50595634 |  | Cal Std | C6H13NO2 | 0.006 | 0.244 | 8964604479 | Calibration Amount Diff -882.183% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C1 | 40239790608 |  | Cal Std | C6H13NO2 | 7.782 | 500 | 5171139025
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C2 | 24509796119 |  | Cal Std | C6H13NO2 | 3.876 | 250 | 6323557872
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C3 | 15190464666 |  | Cal Std | C6H13NO2 | 2.061 | 125 | 7370489703
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C4 | 8691857293 |  | Cal Std | C6H13NO2 | 1.06 | 62.5 | 8198607199
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C5 | 4515421498 |  | Cal Std | C6H13NO2 | 0.556 | 31.25 | 8117653316
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C6 | 2495522008 |  | Cal Std | C6H13NO2 | 0.294 | 15.625 | 8477835775
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C7 | 1204465639 |  | Cal Std | C6H13NO2 | 0.139 | 7.813 | 8664260118
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C8 | 752919303 |  | Cal Std | C6H13NO2 | 0.085 | 3.906 | 8834200148
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C9 | 359300131 |  | Cal Std | C6H13NO2 | 0.041 | 1.953 | 8663885861 | Calibration Amount Diff -78.797% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C10 | 181451731 |  | Cal Std | C6H13NO2 | 0.021 | 0.977 | 8707297502 | Calibration Amount Diff -194.518% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C11 | 92056119 |  | Cal Std | C6H13NO2 | 0.011 | 0.488 | 8554831072 | Calibration Amount Diff -423.117% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_C12 | 45933944 |  | Cal Std | C6H13NO2 | 0.005 | 0.244 | 8695505339 | Calibration Amount Diff -891.786% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A1 | 37902469851 |  | Cal Std | C6H13NO2 | 7.379 | 500 | 5136530466
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A2 | 23173653050 |  | Cal Std | C6H13NO2 | 3.628 | 250 | 6387573701
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A3 | 14838576509 |  | Cal Std | C6H13NO2 | 2.05 | 125 | 7239491161
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A4 | 8450536258 |  | Cal Std | C6H13NO2 | 1.061 | 62.5 | 7965266769
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A5 | 4473057809 |  | Cal Std | C6H13NO2 | 0.557 | 31.25 | 8029948529
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A6 | 2365247617 |  | Cal Std | C6H13NO2 | 0.284 | 15.625 | 8330371457
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A7 | 1243176340 |  | Cal Std | C6H13NO2 | 0.144 | 7.813 | 8619919646
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A8 | 757621571 |  | Cal Std | C6H13NO2 | 0.082 | 3.906 | 9238332021 | Calibration Amount Diff -22.117% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A9 | 357542264 |  | Cal Std | C6H13NO2 | 0.042 | 1.953 | 8413916886 | Calibration Amount Diff -75.4% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A10 | 188822361 |  | Cal Std | C6H13NO2 | 0.022 | 0.977 | 8575019675 | Calibration Amount Diff -186.682% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A11 | 91822202 |  | Cal Std | C6H13NO2 | 0.011 | 0.488 | 8605078666 | Calibration Amount Diff -424.313% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD082321_082521_A12 | 47795284 |  | Cal Std | C6H13NO2 | 0.006 | 0.244 | 8451433849 | Calibration Amount Diff -881.881% (absolute) > max calibration amount diff 20%

## Isoleucine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B1 | 4775834017 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B2 | 6333966369 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B3 | 7237088876 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B4 | 7891520819 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B5 | 8384278182 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B6 | 8907463964 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B7 | 8799278780 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B8 | 8268872721 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B9 | 8696481935 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B10 | 8840685952 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B11 | 8632972253 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_B12 | 8964604479 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C1 | 5171139025 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C2 | 6323557872 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C3 | 7370489703 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C4 | 8198607199 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C5 | 8117653316 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C6 | 8477835775 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C7 | 8664260118 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C8 | 8834200148 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C9 | 8663885861 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C10 | 8707297502 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C11 | 8554831072 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_C12 | 8695505339 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A1 | 5136530466 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A2 | 6387573701 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A3 | 7239491161 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A4 | 7965266769 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A5 | 8029948529 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A6 | 8330371457 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A7 | 8619919646 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A8 | 9238332021 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A9 | 8413916886 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A10 | 8575019675 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A11 | 8605078666 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD082321_082521_A12 | 8451433849 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0

## Leucine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Leucine pos | 6.74 | Target Compound | KD082321_082521_B1 | 40186901185 |  | Cal Std | C6H13NO2 | 8.148 | 500 | 4932075460
Leucine pos | 6.74 | Target Compound | KD082321_082521_B2 | 22846904095 |  | Cal Std | C6H13NO2 | 3.716 | 250 | 6148300072
Leucine pos | 6.74 | Target Compound | KD082321_082521_B3 | 13924273044 |  | Cal Std | C6H13NO2 | 2.022 | 125 | 6887120810
Leucine pos | 6.74 | Target Compound | KD082321_082521_B4 | 8018617155 |  | Cal Std | C6H13NO2 | 1.057 | 62.5 | 7587486680
Leucine pos | 6.74 | Target Compound | KD082321_082521_B5 | 4672784526 |  | Cal Std | C6H13NO2 | 0.571 | 31.25 | 8189781937
Leucine pos | 6.74 | Target Compound | KD082321_082521_B6 | 2364971526 |  | Cal Std | C6H13NO2 | 0.288 | 15.625 | 8218338186
Leucine pos | 6.74 | Target Compound | KD082321_082521_B7 | 1263800133 |  | Cal Std | C6H13NO2 | 0.151 | 7.813 | 8360610017
Leucine pos | 6.74 | Target Compound | KD082321_082521_B8 | 652139708 |  | Cal Std | C6H13NO2 | 0.084 | 3.906 | 7784977258
Leucine pos | 6.74 | Target Compound | KD082321_082521_B9 | 328019435 |  | Cal Std | C6H13NO2 | 0.043 | 1.953 | 7643901352 | Calibration Amount Diff -56.364% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_B10 | 176632443 |  | Cal Std | C6H13NO2 | 0.022 | 0.977 | 7983573199 | Calibration Amount Diff -151.452% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_B11 | 88051200 |  | Cal Std | C6H13NO2 | 0.011 | 0.488 | 7848961792 | Calibration Amount Diff -348.674% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_B12 | 42751754 |  | Cal Std | C6H13NO2 | 0.005 | 0.244 | 7879981961 | Calibration Amount Diff -752.085% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_C1 | 39181503208 |  | Cal Std | C6H13NO2 | 7.785 | 500 | 5033144801
Leucine pos | 6.74 | Target Compound | KD082321_082521_C2 | 22224536530 |  | Cal Std | C6H13NO2 | 3.656 | 250 | 6078091877
Leucine pos | 6.74 | Target Compound | KD082321_082521_C3 | 14346851271 |  | Cal Std | C6H13NO2 | 1.985 | 125 | 7227768693
Leucine pos | 6.74 | Target Compound | KD082321_082521_C4 | 8137525326 |  | Cal Std | C6H13NO2 | 1.078 | 62.5 | 7551473925
Leucine pos | 6.74 | Target Compound | KD082321_082521_C5 | 4180259713 |  | Cal Std | C6H13NO2 | 0.566 | 31.25 | 7383279849
Leucine pos | 6.74 | Target Compound | KD082321_082521_C6 | 2325278179 |  | Cal Std | C6H13NO2 | 0.288 | 15.625 | 8085575469
Leucine pos | 6.74 | Target Compound | KD082321_082521_C7 | 1140624468 |  | Cal Std | C6H13NO2 | 0.144 | 7.813 | 7908715032
Leucine pos | 6.74 | Target Compound | KD082321_082521_C8 | 687122663 |  | Cal Std | C6H13NO2 | 0.084 | 3.906 | 8164051320
Leucine pos | 6.74 | Target Compound | KD082321_082521_C9 | 347923708 |  | Cal Std | C6H13NO2 | 0.042 | 1.953 | 8373782834 | Calibration Amount Diff -60.914% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_C10 | 177314612 |  | Cal Std | C6H13NO2 | 0.022 | 0.977 | 7889317015 | Calibration Amount Diff -149.112% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_C11 | 88430003 |  | Cal Std | C6H13NO2 | 0.011 | 0.488 | 7981051578 | Calibration Amount Diff -350.52% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_C12 | 46990907 |  | Cal Std | C6H13NO2 | 0.006 | 0.244 | 8033306583 | Calibration Amount Diff -740.756% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_A1 | 38901710266 |  | Cal Std | C6H13NO2 | 7.361 | 500 | 5284839236
Leucine pos | 6.74 | Target Compound | KD082321_082521_A2 | 21803648179 |  | Cal Std | C6H13NO2 | 3.647 | 250 | 5978154825
Leucine pos | 6.74 | Target Compound | KD082321_082521_A3 | 14335429928 |  | Cal Std | C6H13NO2 | 2.062 | 125 | 6951829734
Leucine pos | 6.74 | Target Compound | KD082321_082521_A4 | 8239683020 |  | Cal Std | C6H13NO2 | 1.081 | 62.5 | 7619069881
Leucine pos | 6.74 | Target Compound | KD082321_082521_A5 | 4249132703 |  | Cal Std | C6H13NO2 | 0.546 | 31.25 | 7778097288
Leucine pos | 6.74 | Target Compound | KD082321_082521_A6 | 2243997278 |  | Cal Std | C6H13NO2 | 0.294 | 15.625 | 7643291354
Leucine pos | 6.74 | Target Compound | KD082321_082521_A7 | 1192152234 |  | Cal Std | C6H13NO2 | 0.149 | 7.813 | 7975184024
Leucine pos | 6.74 | Target Compound | KD082321_082521_A8 | 728004332 |  | Cal Std | C6H13NO2 | 0.085 | 3.906 | 8588774858
Leucine pos | 6.74 | Target Compound | KD082321_082521_A9 | 334157076 |  | Cal Std | C6H13NO2 | 0.042 | 1.953 | 7962511511 | Calibration Amount Diff -59.522% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_A10 | 172244538 |  | Cal Std | C6H13NO2 | 0.022 | 0.977 | 7874625148 | Calibration Amount Diff -153.128% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_A11 | 88637061 |  | Cal Std | C6H13NO2 | 0.011 | 0.488 | 7765843701 | Calibration Amount Diff -346.063% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD082321_082521_A12 | 45752197 |  | Cal Std | C6H13NO2 | 0.006 | 0.244 | 7707102460 | Calibration Amount Diff -738.435% (absolute) > max calibration amount diff 20%

## Leucine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B1 | 4932075460 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B2 | 6148300072 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B3 | 6887120810 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B4 | 7587486680 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B5 | 8189781937 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B6 | 8218338186 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B7 | 8360610017 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B8 | 7784977258 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B9 | 7643901352 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B10 | 7983573199 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B11 | 7848961792 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_B12 | 7879981961 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C1 | 5033144801 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C2 | 6078091877 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C3 | 7227768693 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C4 | 7551473925 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C5 | 7383279849 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C6 | 8085575469 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C7 | 7908715032 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C8 | 8164051320 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C9 | 8373782834 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C10 | 7889317015 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C11 | 7981051578 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_C12 | 8033306583 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A1 | 5284839236 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A2 | 5978154825 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A3 | 6951829734 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A4 | 7619069881 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A5 | 7778097288 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A6 | 7643291354 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A7 | 7975184024 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A8 | 8588774858 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A9 | 7962511511 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A10 | 7874625148 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A11 | 7765843701 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD082321_082521_A12 | 7707102460 |  | Cal Std | [13]C6H13[15]NO2 | 0 | 1 | 0

## Lysine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Lysine pos | 13.78 | Target Compound | KD082321_082521_B1 | 9479966795 |  | Cal Std | C6H14N2O2 | 8.98 | 500 | 1055640032
Lysine pos | 13.78 | Target Compound | KD082321_082521_B2 | 6193568824 |  | Cal Std | C6H14N2O2 | 4.169 | 250 | 1485795297
Lysine pos | 13.78 | Target Compound | KD082321_082521_B3 | 4051273665 |  | Cal Std | C6H14N2O2 | 2.247 | 125 | 1802997644
Lysine pos | 13.78 | Target Compound | KD082321_082521_B4 | 2432307773 |  | Cal Std | C6H14N2O2 | 1.169 | 62.5 | 2080529522
Lysine pos | 13.78 | Target Compound | KD082321_082521_B5 | 1433603254 |  | Cal Std | C6H14N2O2 | 0.613 | 31.25 | 2340498018
Lysine pos | 13.78 | Target Compound | KD082321_082521_B6 | 733968167 |  | Cal Std | C6H14N2O2 | 0.307 | 15.625 | 2388531258
Lysine pos | 13.78 | Target Compound | KD082321_082521_B7 | 377918080 |  | Cal Std | C6H14N2O2 | 0.163 | 7.813 | 2322196616
Lysine pos | 13.78 | Target Compound | KD082321_082521_B8 | 187246278 |  | Cal Std | C6H14N2O2 | 0.088 | 3.906 | 2121528629
Lysine pos | 13.78 | Target Compound | KD082321_082521_B9 | 95454991 |  | Cal Std | C6H14N2O2 | 0.045 | 1.953 | 2104171647 | Calibration Amount Diff -57.688% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_B10 | 51686954 |  | Cal Std | C6H14N2O2 | 0.024 | 0.977 | 2175526735 | Calibration Amount Diff -145.777% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_B11 | 26136549 |  | Cal Std | C6H14N2O2 | 0.012 | 0.488 | 2139277754 | Calibration Amount Diff -331.05% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_B12 | 13498180 |  | Cal Std | C6H14N2O2 | 0.006 | 0.244 | 2161514306 | Calibration Amount Diff -706.388% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_C1 | 9457571361 |  | Cal Std | C6H14N2O2 | 8.535 | 500 | 1108044877
Lysine pos | 13.78 | Target Compound | KD082321_082521_C2 | 6234012246 |  | Cal Std | C6H14N2O2 | 4.119 | 250 | 1513427426
Lysine pos | 13.78 | Target Compound | KD082321_082521_C3 | 4288819874 |  | Cal Std | C6H14N2O2 | 2.252 | 125 | 1904287017
Lysine pos | 13.78 | Target Compound | KD082321_082521_C4 | 2465798023 |  | Cal Std | C6H14N2O2 | 1.176 | 62.5 | 2096757872
Lysine pos | 13.78 | Target Compound | KD082321_082521_C5 | 1347060607 |  | Cal Std | C6H14N2O2 | 0.61 | 31.25 | 2207002654
Lysine pos | 13.78 | Target Compound | KD082321_082521_C6 | 731906792 |  | Cal Std | C6H14N2O2 | 0.315 | 15.625 | 2321081082
Lysine pos | 13.78 | Target Compound | KD082321_082521_C7 | 357155114 |  | Cal Std | C6H14N2O2 | 0.154 | 7.813 | 2326302394
Lysine pos | 13.78 | Target Compound | KD082321_082521_C8 | 214917385 |  | Cal Std | C6H14N2O2 | 0.093 | 3.906 | 2318715929
Lysine pos | 13.78 | Target Compound | KD082321_082521_C9 | 98274533 |  | Cal Std | C6H14N2O2 | 0.044 | 1.953 | 2228390642 | Calibration Amount Diff -61.502% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_C10 | 51954154 |  | Cal Std | C6H14N2O2 | 0.024 | 0.977 | 2202468858 | Calibration Amount Diff -146.799% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_C11 | 27327405 |  | Cal Std | C6H14N2O2 | 0.012 | 0.488 | 2205354722 | Calibration Amount Diff -328.95% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_C12 | 12517074 |  | Cal Std | C6H14N2O2 | 0.006 | 0.244 | 2021913926 | Calibration Amount Diff -707.695% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_A1 | 7971968918 |  | Cal Std | C6H14N2O2 | 8.188 | 500 | 973617297
Lysine pos | 13.78 | Target Compound | KD082321_082521_A2 | 5543942924 |  | Cal Std | C6H14N2O2 | 3.977 | 250 | 1394022820
Lysine pos | 13.78 | Target Compound | KD082321_082521_A3 | 3747708464 |  | Cal Std | C6H14N2O2 | 2.257 | 125 | 1660725754
Lysine pos | 13.78 | Target Compound | KD082321_082521_A4 | 2295545016 |  | Cal Std | C6H14N2O2 | 1.168 | 62.5 | 1965883490
Lysine pos | 13.78 | Target Compound | KD082321_082521_A5 | 1282775824 |  | Cal Std | C6H14N2O2 | 0.607 | 31.25 | 2111891446
Lysine pos | 13.78 | Target Compound | KD082321_082521_A6 | 706458408 |  | Cal Std | C6H14N2O2 | 0.314 | 15.625 | 2253192345
Lysine pos | 13.78 | Target Compound | KD082321_082521_A7 | 374996346 |  | Cal Std | C6H14N2O2 | 0.159 | 7.813 | 2357744186
Lysine pos | 13.78 | Target Compound | KD082321_082521_A8 | 208727480 |  | Cal Std | C6H14N2O2 | 0.089 | 3.906 | 2348077162
Lysine pos | 13.78 | Target Compound | KD082321_082521_A9 | 95899275 |  | Cal Std | C6H14N2O2 | 0.046 | 1.953 | 2095594175 | Calibration Amount Diff -56.488% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_A10 | 49541543 |  | Cal Std | C6H14N2O2 | 0.024 | 0.977 | 2073960610 | Calibration Amount Diff -144.999% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_A11 | 25721119 |  | Cal Std | C6H14N2O2 | 0.013 | 0.488 | 2049583156 | Calibration Amount Diff -327.04% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD082321_082521_A12 | 12655223 |  | Cal Std | C6H14N2O2 | 0.006 | 0.244 | 1996844216 | Calibration Amount Diff -704.146% (absolute) > max calibration amount diff 20%

## Lysine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B1 | 1055640032 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B2 | 1485795297 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B3 | 1802997644 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B4 | 2080529522 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B5 | 2340498018 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B6 | 2388531258 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B7 | 2322196616 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B8 | 2121528629 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B9 | 2104171647 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B10 | 2175526735 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B11 | 2139277754 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_B12 | 2161514306 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C1 | 1108044877 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C2 | 1513427426 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C3 | 1904287017 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C4 | 2096757872 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C5 | 2207002654 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C6 | 2321081082 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C7 | 2326302394 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C8 | 2318715929 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C9 | 2228390642 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C10 | 2202468858 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C11 | 2205354722 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_C12 | 2021913926 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A1 | 973617297 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0 | 13.810:Peak area 973617296.571 is out of bounds (ISTD Minimum Recovery 998792937.711 and ISTD Max Recovery 2996378813.132)
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A2 | 1394022820 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A3 | 1660725754 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A4 | 1965883490 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A5 | 2111891446 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A6 | 2253192345 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A7 | 2357744186 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A8 | 2348077162 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A9 | 2095594175 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A10 | 2073960610 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A11 | 2049583156 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD082321_082521_A12 | 1996844216 |  | Cal Std | [13]C6H14[15]N2O2 | 0 | 1 | 0

## Methionine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Methionine pos | 7.14 | Target Compound | KD082321_082521_B1 | 43260805218 |  | Cal Std | C5H11NO2S | 8.852 | 500 | 4886950727
Methionine pos | 7.14 | Target Compound | KD082321_082521_B2 | 26765322889 |  | Cal Std | C5H11NO2S | 4.114 | 250 | 6505535041
Methionine pos | 7.14 | Target Compound | KD082321_082521_B3 | 17307973570 |  | Cal Std | C5H11NO2S | 2.213 | 125 | 7821793566
Methionine pos | 7.14 | Target Compound | KD082321_082521_B4 | 10063602222 |  | Cal Std | C5H11NO2S | 1.163 | 62.5 | 8655056393
Methionine pos | 7.14 | Target Compound | KD082321_082521_B5 | 5874455802 |  | Cal Std | C5H11NO2S | 0.601 | 31.25 | 9779549180
Methionine pos | 7.14 | Target Compound | KD082321_082521_B6 | 2980474121 |  | Cal Std | C5H11NO2S | 0.306 | 15.625 | 9749086373
Methionine pos | 7.14 | Target Compound | KD082321_082521_B7 | 1645040233 |  | Cal Std | C5H11NO2S | 0.164 | 7.813 | 10052568537
Methionine pos | 7.14 | Target Compound | KD082321_082521_B8 | 840374024 |  | Cal Std | C5H11NO2S | 0.088 | 3.906 | 9541969313
Methionine pos | 7.14 | Target Compound | KD082321_082521_B9 | 429398500 |  | Cal Std | C5H11NO2S | 0.045 | 1.953 | 9626335847 | Calibration Amount Diff -53.641% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_B10 | 230737596 |  | Cal Std | C5H11NO2S | 0.023 | 0.977 | 10179814489 | Calibration Amount Diff -141.533% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_B11 | 115638413 |  | Cal Std | C5H11NO2S | 0.011 | 0.488 | 10111893995 | Calibration Amount Diff -320.677% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_B12 | 56665739 |  | Cal Std | C5H11NO2S | 0.006 | 0.244 | 10068407521 | Calibration Amount Diff -683.598% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_C1 | 42440162092 |  | Cal Std | C5H11NO2S | 8.523 | 500 | 4979440779
Methionine pos | 7.14 | Target Compound | KD082321_082521_C2 | 26260754413 |  | Cal Std | C5H11NO2S | 4.05 | 250 | 6484772859
Methionine pos | 7.14 | Target Compound | KD082321_082521_C3 | 17744325877 |  | Cal Std | C5H11NO2S | 2.204 | 125 | 8049914431
Methionine pos | 7.14 | Target Compound | KD082321_082521_C4 | 10481816029 |  | Cal Std | C5H11NO2S | 1.156 | 62.5 | 9068252114
Methionine pos | 7.14 | Target Compound | KD082321_082521_C5 | 5573857089 |  | Cal Std | C5H11NO2S | 0.599 | 31.25 | 9306529054
Methionine pos | 7.14 | Target Compound | KD082321_082521_C6 | 3037578158 |  | Cal Std | C5H11NO2S | 0.312 | 15.625 | 9741549225
Methionine pos | 7.14 | Target Compound | KD082321_082521_C7 | 1532269938 |  | Cal Std | C5H11NO2S | 0.153 | 7.813 | 10040355346
Methionine pos | 7.14 | Target Compound | KD082321_082521_C8 | 916367589 |  | Cal Std | C5H11NO2S | 0.091 | 3.906 | 10114143864
Methionine pos | 7.14 | Target Compound | KD082321_082521_C9 | 448645377 |  | Cal Std | C5H11NO2S | 0.045 | 1.953 | 10071685961 | Calibration Amount Diff -53.829% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_C10 | 232818011 |  | Cal Std | C5H11NO2S | 0.023 | 0.977 | 10239560367 | Calibration Amount Diff -141.099% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_C11 | 110609548 |  | Cal Std | C5H11NO2S | 0.011 | 0.488 | 9905999836 | Calibration Amount Diff -323.983% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_C12 | 55644113 |  | Cal Std | C5H11NO2S | 0.005 | 0.244 | 10165339419 | Calibration Amount Diff -687.374% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_A1 | 42487891521 |  | Cal Std | C5H11NO2S | 7.989 | 500 | 5318489606
Methionine pos | 7.14 | Target Compound | KD082321_082521_A2 | 25817075015 |  | Cal Std | C5H11NO2S | 3.912 | 250 | 6600078936
Methionine pos | 7.14 | Target Compound | KD082321_082521_A3 | 17294089459 |  | Cal Std | C5H11NO2S | 2.213 | 125 | 7813885785
Methionine pos | 7.14 | Target Compound | KD082321_082521_A4 | 10093031264 |  | Cal Std | C5H11NO2S | 1.147 | 62.5 | 8803005648
Methionine pos | 7.14 | Target Compound | KD082321_082521_A5 | 5536790555 |  | Cal Std | C5H11NO2S | 0.594 | 31.25 | 9328436111
Methionine pos | 7.14 | Target Compound | KD082321_082521_A6 | 3037243127 |  | Cal Std | C5H11NO2S | 0.308 | 15.625 | 9862539399
Methionine pos | 7.14 | Target Compound | KD082321_082521_A7 | 1571535017 |  | Cal Std | C5H11NO2S | 0.157 | 7.813 | 9982542749
Methionine pos | 7.14 | Target Compound | KD082321_082521_A8 | 954893254 |  | Cal Std | C5H11NO2S | 0.089 | 3.906 | 10753610151
Methionine pos | 7.14 | Target Compound | KD082321_082521_A9 | 452403925 |  | Cal Std | C5H11NO2S | 0.045 | 1.953 | 10004880999 | Calibration Amount Diff -51.769% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_A10 | 232909470 |  | Cal Std | C5H11NO2S | 0.023 | 0.977 | 9966167573 | Calibration Amount Diff -137.228% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_A11 | 115558444 |  | Cal Std | C5H11NO2S | 0.012 | 0.488 | 10030382695 | Calibration Amount Diff -319.636% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD082321_082521_A12 | 53994135 |  | Cal Std | C5H11NO2S | 0.006 | 0.244 | 9754056184 | Calibration Amount Diff -685.864% (absolute) > max calibration amount diff 20%

## Methionine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B1 | 4886950727 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B2 | 6505535041 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B3 | 7821793566 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B4 | 8655056393 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B5 | 9779549180 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B6 | 9749086373 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B7 | 10052568537 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B8 | 9541969313 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B9 | 9626335847 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B10 | 10179814489 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B11 | 10111893995 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_B12 | 10068407521 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C1 | 4979440779 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C2 | 6484772859 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C3 | 8049914431 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C4 | 9068252114 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C5 | 9306529054 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C6 | 9741549225 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C7 | 10040355346 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C8 | 10114143864 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C9 | 10071685961 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C10 | 10239560367 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C11 | 9905999836 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_C12 | 10165339419 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A1 | 5318489606 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A2 | 6600078936 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A3 | 7813885785 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A4 | 8803005648 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A5 | 9328436111 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A6 | 9862539399 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A7 | 9982542749 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A8 | 10753610151 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A9 | 10004880999 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A10 | 9966167573 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A11 | 10030382695 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD082321_082521_A12 | 9754056184 |  | Cal Std | [13]C5H11[15]NO2S | 0 | 1 | 0

## Phenylalanine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B1 | 69809652484 |  | Cal Std | C9H11NO2 | 8.579 | 500 | 8137041799
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B2 | 39889271119 |  | Cal Std | C9H11NO2 | 4.004 | 250 | 9962834656
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B3 | 24667178818 |  | Cal Std | C9H11NO2 | 2.156 | 125 | 11442751104
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B4 | 14050307616 |  | Cal Std | C9H11NO2 | 1.136 | 62.5 | 12373566203
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B5 | 8134723100 |  | Cal Std | C9H11NO2 | 0.596 | 31.25 | 13651622830
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B6 | 4011552664 |  | Cal Std | C9H11NO2 | 0.302 | 15.625 | 13303218122
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B7 | 2128012864 |  | Cal Std | C9H11NO2 | 0.162 | 7.813 | 13151492336
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B8 | 1131375386 |  | Cal Std | C9H11NO2 | 0.087 | 3.906 | 12986233321
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B9 | 584394778 |  | Cal Std | C9H11NO2 | 0.045 | 1.953 | 12923508685 | Calibration Amount Diff -58.219% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B10 | 294592229 |  | Cal Std | C9H11NO2 | 0.023 | 0.977 | 12847524565 | Calibration Amount Diff -156.583% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B11 | 145949882 |  | Cal Std | C9H11NO2 | 0.012 | 0.488 | 12460124246 | Calibration Amount Diff -354.429% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_B12 | 72958398 |  | Cal Std | C9H11NO2 | 0.006 | 0.244 | 12874799396 | Calibration Amount Diff -761.037% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C1 | 66722961659 |  | Cal Std | C9H11NO2 | 8.217 | 500 | 8120490936
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C2 | 39184566756 |  | Cal Std | C9H11NO2 | 3.92 | 250 | 9995356191
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C3 | 24689908708 |  | Cal Std | C9H11NO2 | 2.142 | 125 | 11524429438
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C4 | 14472620277 |  | Cal Std | C9H11NO2 | 1.128 | 62.5 | 12833257387
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C5 | 7331328153 |  | Cal Std | C9H11NO2 | 0.585 | 31.25 | 12529069005
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C6 | 4039426458 |  | Cal Std | C9H11NO2 | 0.308 | 15.625 | 13120176869
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C7 | 1887717819 |  | Cal Std | C9H11NO2 | 0.149 | 7.813 | 12634246555
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C8 | 1161756885 |  | Cal Std | C9H11NO2 | 0.09 | 3.906 | 12886349726
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C9 | 585482230 |  | Cal Std | C9H11NO2 | 0.044 | 1.953 | 13262707949 | Calibration Amount Diff -61.598% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C10 | 298410189 |  | Cal Std | C9H11NO2 | 0.023 | 0.977 | 13067279318 | Calibration Amount Diff -157.17% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C11 | 155487556 |  | Cal Std | C9H11NO2 | 0.012 | 0.488 | 13367909962 | Calibration Amount Diff -355.46% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_C12 | 70450057 |  | Cal Std | C9H11NO2 | 0.006 | 0.244 | 12502518861 | Calibration Amount Diff -761.84% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A1 | 64730738226 |  | Cal Std | C9H11NO2 | 7.891 | 500 | 8202949901
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A2 | 38522002982 |  | Cal Std | C9H11NO2 | 3.839 | 250 | 10035509667
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A3 | 24509971041 |  | Cal Std | C9H11NO2 | 2.174 | 125 | 11275161735
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A4 | 13972350160 |  | Cal Std | C9H11NO2 | 1.131 | 62.5 | 12353752371
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A5 | 7675955795 |  | Cal Std | C9H11NO2 | 0.59 | 31.25 | 13012892629
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A6 | 3943136936 |  | Cal Std | C9H11NO2 | 0.305 | 15.625 | 12941385659
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A7 | 2037363687 |  | Cal Std | C9H11NO2 | 0.157 | 7.813 | 12977587804
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A8 | 1214648986 |  | Cal Std | C9H11NO2 | 0.087 | 3.906 | 13902144959
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A9 | 573634056 |  | Cal Std | C9H11NO2 | 0.045 | 1.953 | 12620865440 | Calibration Amount Diff -57.49% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A10 | 303609811 |  | Cal Std | C9H11NO2 | 0.024 | 0.977 | 12791730371 | Calibration Amount Diff -151.523% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A11 | 149354080 |  | Cal Std | C9H11NO2 | 0.012 | 0.488 | 12972490148 | Calibration Amount Diff -356.948% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD082321_082521_A12 | 78379512 |  | Cal Std | C9H11NO2 | 0.006 | 0.244 | 13006445714 | Calibration Amount Diff -751.991% (absolute) > max calibration amount diff 20%

## Phenylalanine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B1 | 8137041799 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B2 | 9962834656 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B3 | 11442751104 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B4 | 12373566203 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B5 | 13651622830 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B6 | 13303218122 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B7 | 13151492336 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B8 | 12986233321 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B9 | 12923508685 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B10 | 12847524565 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B11 | 12460124246 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_B12 | 12874799396 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C1 | 8120490936 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C2 | 9995356191 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C3 | 11524429438 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C4 | 12833257387 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C5 | 12529069005 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C6 | 13120176869 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C7 | 12634246555 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C8 | 12886349726 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C9 | 13262707949 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C10 | 13067279318 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C11 | 13367909962 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_C12 | 12502518861 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A1 | 8202949901 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A2 | 10035509667 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A3 | 11275161735 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A4 | 12353752371 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A5 | 13012892629 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A6 | 12941385659 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A7 | 12977587804 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A8 | 13902144959 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A9 | 12620865440 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A10 | 12791730371 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A11 | 12972490148 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD082321_082521_A12 | 13006445714 |  | Cal Std | [13]C9H11[15]NO2 | 0 | 1 | 0

## Proline pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Proline pos | 8.15 | Target Compound | KD082321_082521_B1 | 74683196752 |  | Cal Std | C5H9NO2 | 8 | 500 | 9335264391
Proline pos | 8.15 | Target Compound | KD082321_082521_B2 | 48204861323 |  | Cal Std | C5H9NO2 | 3.793 | 250 | 12709846020
Proline pos | 8.15 | Target Compound | KD082321_082521_B3 | 33019165223 |  | Cal Std | C5H9NO2 | 2.083 | 125 | 15853229804
Proline pos | 8.15 | Target Compound | KD082321_082521_B4 | 20735600321 |  | Cal Std | C5H9NO2 | 1.099 | 62.5 | 18865278735
Proline pos | 8.15 | Target Compound | KD082321_082521_B5 | 12492003902 |  | Cal Std | C5H9NO2 | 0.577 | 31.25 | 21648782186
Proline pos | 8.15 | Target Compound | KD082321_082521_B6 | 6941892013 |  | Cal Std | C5H9NO2 | 0.298 | 15.625 | 23315044292
Proline pos | 8.15 | Target Compound | KD082321_082521_B7 | 3932553852 |  | Cal Std | C5H9NO2 | 0.159 | 7.813 | 24671570209
Proline pos | 8.15 | Target Compound | KD082321_082521_B8 | 2099291404 |  | Cal Std | C5H9NO2 | 0.087 | 3.906 | 24193049899 | Calibration Amount Diff -33.443% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_B9 | 1109951721 |  | Cal Std | C5H9NO2 | 0.045 | 1.953 | 24729270944 | Calibration Amount Diff -108.117% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_B10 | 568345471 |  | Cal Std | C5H9NO2 | 0.023 | 0.977 | 24587337645 | Calibration Amount Diff -262.943% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_B11 | 294120985 |  | Cal Std | C5H9NO2 | 0.012 | 0.488 | 24757474355 | Calibration Amount Diff -577.823% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_B12 | 147020481 |  | Cal Std | C5H9NO2 | 0.006 | 0.244 | 24593278365 | Calibration Amount Diff -1214.921% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_C1 | 74967406419 |  | Cal Std | C5H9NO2 | 7.668 | 500 | 9776428873
Proline pos | 8.15 | Target Compound | KD082321_082521_C2 | 48427035116 |  | Cal Std | C5H9NO2 | 3.744 | 250 | 12936079309
Proline pos | 8.15 | Target Compound | KD082321_082521_C3 | 33245412667 |  | Cal Std | C5H9NO2 | 2.051 | 125 | 16212811484
Proline pos | 8.15 | Target Compound | KD082321_082521_C4 | 20966764593 |  | Cal Std | C5H9NO2 | 1.088 | 62.5 | 19267215088
Proline pos | 8.15 | Target Compound | KD082321_082521_C5 | 12087550158 |  | Cal Std | C5H9NO2 | 0.577 | 31.25 | 20951068276
Proline pos | 8.15 | Target Compound | KD082321_082521_C6 | 6761412879 |  | Cal Std | C5H9NO2 | 0.301 | 15.625 | 22457604655
Proline pos | 8.15 | Target Compound | KD082321_082521_C7 | 3691040077 |  | Cal Std | C5H9NO2 | 0.148 | 7.813 | 24903309134
Proline pos | 8.15 | Target Compound | KD082321_082521_C8 | 2163326725 |  | Cal Std | C5H9NO2 | 0.09 | 3.906 | 24147991358 | Calibration Amount Diff -28.7% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_C9 | 1100506257 |  | Cal Std | C5H9NO2 | 0.044 | 1.953 | 24760295333 | Calibration Amount Diff -109.593% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_C10 | 583583275 |  | Cal Std | C5H9NO2 | 0.023 | 0.977 | 24940176602 | Calibration Amount Diff -261.03% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_C11 | 308312702 |  | Cal Std | C5H9NO2 | 0.012 | 0.488 | 25738935464 | Calibration Amount Diff -576.495% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_C12 | 151637072 |  | Cal Std | C5H9NO2 | 0.006 | 0.244 | 24953425334 | Calibration Amount Diff -1212.257% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_A1 | 72041446760 |  | Cal Std | C5H9NO2 | 7.294 | 500 | 9876773465
Proline pos | 8.15 | Target Compound | KD082321_082521_A2 | 47514175151 |  | Cal Std | C5H9NO2 | 3.623 | 250 | 13115641491
Proline pos | 8.15 | Target Compound | KD082321_082521_A3 | 33250426097 |  | Cal Std | C5H9NO2 | 2.101 | 125 | 15825209652
Proline pos | 8.15 | Target Compound | KD082321_082521_A4 | 21270758973 |  | Cal Std | C5H9NO2 | 1.097 | 62.5 | 19389504043
Proline pos | 8.15 | Target Compound | KD082321_082521_A5 | 12226895979 |  | Cal Std | C5H9NO2 | 0.57 | 31.25 | 21432902393
Proline pos | 8.15 | Target Compound | KD082321_082521_A6 | 7069206134 |  | Cal Std | C5H9NO2 | 0.302 | 15.625 | 23392839621
Proline pos | 8.15 | Target Compound | KD082321_082521_A7 | 3745856919 |  | Cal Std | C5H9NO2 | 0.154 | 7.813 | 24264341994
Proline pos | 8.15 | Target Compound | KD082321_082521_A8 | 2261293401 |  | Cal Std | C5H9NO2 | 0.088 | 3.906 | 25655211342 | Calibration Amount Diff -31.135% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_A9 | 1081011062 |  | Cal Std | C5H9NO2 | 0.045 | 1.953 | 23792479530 | Calibration Amount Diff -106.26% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_A10 | 575193950 |  | Cal Std | C5H9NO2 | 0.024 | 0.977 | 24154369633 | Calibration Amount Diff -258.24% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_A11 | 288361152 |  | Cal Std | C5H9NO2 | 0.012 | 0.488 | 24415892917 | Calibration Amount Diff -578.763% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD082321_082521_A12 | 158006457 |  | Cal Std | C5H9NO2 | 0.006 | 0.244 | 25042563917 | Calibration Amount Diff -1205.977% (absolute) > max calibration amount diff 20%

## Proline U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B1 | 9335264391 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0 | 8.157:Peak area 9335264391.296 is out of bounds (ISTD Minimum Recovery 10425868024.343 and ISTD Max Recovery 31277604073.028)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B2 | 12709846020 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B3 | 15853229804 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B4 | 18865278735 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B5 | 21648782186 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B6 | 23315044292 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B7 | 24671570209 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B8 | 24193049899 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B9 | 24729270944 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B10 | 24587337645 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B11 | 24757474355 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_B12 | 24593278365 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C1 | 9776428873 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0 | 8.133:Peak area 9776428873.334 is out of bounds (ISTD Minimum Recovery 10425868024.343 and ISTD Max Recovery 31277604073.028)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C2 | 12936079309 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C3 | 16212811484 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C4 | 19267215088 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C5 | 20951068276 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C6 | 22457604655 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C7 | 24903309134 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C8 | 24147991358 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C9 | 24760295333 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C10 | 24940176602 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C11 | 25738935464 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_C12 | 24953425334 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A1 | 9876773465 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0 | 8.157:Peak area 9876773464.547 is out of bounds (ISTD Minimum Recovery 10425868024.343 and ISTD Max Recovery 31277604073.028)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A2 | 13115641491 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A3 | 15825209652 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A4 | 19389504043 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A5 | 21432902393 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A6 | 23392839621 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A7 | 24264341994 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A8 | 25655211342 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A9 | 23792479530 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A10 | 24154369633 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A11 | 24415892917 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD082321_082521_A12 | 25042563917 |  | Cal Std | [13]C5H9[15]NO2 | 0 | 1 | 0

## Serine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Serine neg | 10.04 | Target Compound | KD082321_082521_B1 | 1192792233 |  | Cal Std | C3H7NO3 | 9.336 | 500 | 127756936
Serine neg | 10.04 | Target Compound | KD082321_082521_B2 | 597219498 |  | Cal Std | C3H7NO3 | 4.2 | 250 | 142199847
Serine neg | 10.04 | Target Compound | KD082321_082521_B3 | 334110834 |  | Cal Std | C3H7NO3 | 2.225 | 125 | 150140647
Serine neg | 10.04 | Target Compound | KD082321_082521_B4 | 178771950 |  | Cal Std | C3H7NO3 | 1.124 | 62.5 | 159020598
Serine neg | 10.04 | Target Compound | KD082321_082521_B5 | 96613226 |  | Cal Std | C3H7NO3 | 0.56 | 31.25 | 172545500
Serine neg | 10.04 | Target Compound | KD082321_082521_B6 | 47719500 |  | Cal Std | C3H7NO3 | 0.279 | 15.625 | 170874818
Serine neg | 10.04 | Target Compound | KD082321_082521_B7 | 23441518 |  | Cal Std | C3H7NO3 | 0.137 | 7.813 | 170628300
Serine neg | 10.04 | Target Compound | KD082321_082521_B8 | 13159959 |  | Cal Std | C3H7NO3 | 0.074 | 3.906 | 178430697
Serine neg | 10.04 | Target Compound | KD082321_082521_B9 | 5718269 |  | Cal Std | C3H7NO3 | 0.032 | 1.953 | 179533372
Serine neg | 10.04 | Target Compound | KD082321_082521_B10 | 2806057 |  | Cal Std | C3H7NO3 | 0.015 | 0.977 | 187221752 | Calibration Amount Diff -25.636% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD082321_082521_B11 | 1504830 |  | Cal Std | C3H7NO3 | 0.008 | 0.488 | 185174924 | Calibration Amount Diff -31.981% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD082321_082521_B12 | 685134 |  | Cal Std | C3H7NO3 | 0.004 | 0.244 | 190998450 | Calibration Amount Diff -70.954% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD082321_082521_C1 | 1120439061 |  | Cal Std | C3H7NO3 | 8.731 | 500 | 128330838
Serine neg | 10.04 | Target Compound | KD082321_082521_C2 | 611015171 |  | Cal Std | C3H7NO3 | 4.144 | 250 | 147448037
Serine neg | 10.04 | Target Compound | KD082321_082521_C3 | 357987512 |  | Cal Std | C3H7NO3 | 2.205 | 125 | 162373525
Serine neg | 10.04 | Target Compound | KD082321_082521_C4 | 191415818 |  | Cal Std | C3H7NO3 | 1.123 | 62.5 | 170401320
Serine neg | 10.04 | Target Compound | KD082321_082521_C5 | 98831777 |  | Cal Std | C3H7NO3 | 0.594 | 31.25 | 166450360
Serine neg | 10.04 | Target Compound | KD082321_082521_C6 | 54903261 |  | Cal Std | C3H7NO3 | 0.303 | 15.625 | 181138836
Serine neg | 10.04 | Target Compound | KD082321_082521_C7 | 26766259 |  | Cal Std | C3H7NO3 | 0.138 | 7.813 | 193421700
Serine neg | 10.04 | Target Compound | KD082321_082521_C8 | 14843833 |  | Cal Std | C3H7NO3 | 0.078 | 3.906 | 190473983
Serine neg | 10.04 | Target Compound | KD082321_082521_C9 | 6017533 |  | Cal Std | C3H7NO3 | 0.032 | 1.953 | 187253971
Serine neg | 10.04 | Target Compound | KD082321_082521_C10 | 2720029 |  | Cal Std | C3H7NO3 | 0.014 | 0.977 | 191559337 | Calibration Amount Diff -30.278% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD082321_082521_C11 | 1205581 |  | Cal Std | C3H7NO3 | 0.006 | 0.488 | 192799044 | Calibration Amount Diff -54.059% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD082321_082521_C12 | 883636 |  | Cal Std | C3H7NO3 | 0.005 | 0.244 | 191003750 | Calibration Amount Diff -46.462% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD082321_082521_A1 | 1120861482 |  | Cal Std | C3H7NO3 | 8.266 | 500 | 135593330
Serine neg | 10.04 | Target Compound | KD082321_082521_A2 | 560738452 |  | Cal Std | C3H7NO3 | 4.101 | 250 | 136740419
Serine neg | 10.04 | Target Compound | KD082321_082521_A3 | 332205097 |  | Cal Std | C3H7NO3 | 2.248 | 125 | 147802097
Serine neg | 10.04 | Target Compound | KD082321_082521_A4 | 178458870 |  | Cal Std | C3H7NO3 | 1.126 | 62.5 | 158529177
Serine neg | 10.04 | Target Compound | KD082321_082521_A5 | 93646457 |  | Cal Std | C3H7NO3 | 0.573 | 31.25 | 163370004
Serine neg | 10.04 | Target Compound | KD082321_082521_A6 | 50894467 |  | Cal Std | C3H7NO3 | 0.287 | 15.625 | 177363556
Serine neg | 10.04 | Target Compound | KD082321_082521_A7 | 26627871 |  | Cal Std | C3H7NO3 | 0.141 | 7.813 | 189081179
Serine neg | 10.04 | Target Compound | KD082321_082521_A8 | 12565978 |  | Cal Std | C3H7NO3 | 0.065 | 3.906 | 192477916
Serine neg | 10.04 | Target Compound | KD082321_082521_A9 | 5824578 |  | Cal Std | C3H7NO3 | 0.031 | 1.953 | 185230144
Serine neg | 10.04 | Target Compound | KD082321_082521_A10 | 3374472 |  | Cal Std | C3H7NO3 | 0.018 | 0.977 | 183768166
Serine neg | 10.04 | Target Compound | KD082321_082521_A11 | 1364269 |  | Cal Std | C3H7NO3 | 0.007 | 0.488 | 184576482 | Calibration Amount Diff -40.645% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD082321_082521_A12 | 951779 |  | Cal Std | C3H7NO3 | 0.005 | 0.244 | 184071179 | Calibration Amount Diff -33.629% (absolute) > max calibration amount diff 20%

## Serine U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B1 | 127756936 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B2 | 142199847 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B3 | 150140647 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B4 | 159020598 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B5 | 172545500 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B6 | 170874818 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B7 | 170628300 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B8 | 178430697 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B9 | 179533372 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B10 | 187221752 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B11 | 185174924 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_B12 | 190998450 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C1 | 128330838 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C2 | 147448037 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C3 | 162373525 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C4 | 170401320 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C5 | 166450360 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C6 | 181138836 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C7 | 193421700 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C8 | 190473983 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C9 | 187253971 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C10 | 191559337 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C11 | 192799044 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_C12 | 191003750 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A1 | 135593330 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A2 | 136740419 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A3 | 147802097 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A4 | 158529177 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A5 | 163370004 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A6 | 177363556 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A7 | 189081179 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A8 | 192477916 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A9 | 185230144 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A10 | 183768166 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A11 | 184576482 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD082321_082521_A12 | 184071179 |  | Cal Std | [13]C3H7[15]NO3 | 0 | 1 | 0

## Threonine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Threonine neg | 9.15 | Target Compound | KD082321_082521_B1 | 4887545479 |  | Cal Std | C4H9NO3 | 9.611 | 500 | 508521781
Threonine neg | 9.15 | Target Compound | KD082321_082521_B2 | 1936010536 |  | Cal Std | C4H9NO3 | 4.6 | 250 | 420840241
Threonine neg | 9.15 | Target Compound | KD082321_082521_B3 | 705304790 |  | Cal Std | C4H9NO3 | 2.2 | 125 | 320650001
Threonine neg | 9.15 | Target Compound | KD082321_082521_B4 | 558677499 |  | Cal Std | C4H9NO3 | 1.129 | 62.5 | 494964612
Threonine neg | 9.15 | Target Compound | KD082321_082521_B5 | 363930371 |  | Cal Std | C4H9NO3 | 0.587 | 31.25 | 619974022
Threonine neg | 9.15 | Target Compound | KD082321_082521_B6 | 219781793 |  | Cal Std | C4H9NO3 | 0.301 | 15.625 | 731355634
Threonine neg | 9.15 | Target Compound | KD082321_082521_B7 | 120123485 |  | Cal Std | C4H9NO3 | 0.164 | 7.813 | 734177995 | Calibration Amount Diff 21.149% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_B8 | 68697381 |  | Cal Std | C4H9NO3 | 0.09 | 3.906 | 765969141 | Calibration Amount Diff 39.137% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_B9 | 38287714 |  | Cal Std | C4H9NO3 | 0.047 | 1.953 | 819877710 | Calibration Amount Diff 58.268% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_B10 | 21394418 |  | Cal Std | C4H9NO3 | 0.024 | 0.977 | 883497686 | Calibration Amount Diff 90.905% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_B11 | 11416462 |  | Cal Std | C4H9NO3 | 0.013 | 0.488 | 864254117 | Calibration Amount Diff 159.239% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_B12 | 6228513 |  | Cal Std | C4H9NO3 | 0.007 | 0.244 | 907329735 | Calibration Amount Diff 276.703% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_C1 | 4639876045 |  | Cal Std | C4H9NO3 | 9.4 | 500 | 493587905
Threonine neg | 9.15 | Target Compound | KD082321_082521_C2 | 1768937907 |  | Cal Std | C4H9NO3 | 4.398 | 250 | 402203207
Threonine neg | 9.15 | Target Compound | KD082321_082521_C3 | 738108598 |  | Cal Std | C4H9NO3 | 2.158 | 125 | 342086693
Threonine neg | 9.15 | Target Compound | KD082321_082521_C4 | 566390587 |  | Cal Std | C4H9NO3 | 1.12 | 62.5 | 505832522
Threonine neg | 9.15 | Target Compound | KD082321_082521_C5 | 370908639 |  | Cal Std | C4H9NO3 | 0.596 | 31.25 | 621861392
Threonine neg | 9.15 | Target Compound | KD082321_082521_C6 | 226955018 |  | Cal Std | C4H9NO3 | 0.308 | 15.625 | 737983097
Threonine neg | 9.15 | Target Compound | KD082321_082521_C7 | 123801669 |  | Cal Std | C4H9NO3 | 0.151 | 7.813 | 819767936
Threonine neg | 9.15 | Target Compound | KD082321_082521_C8 | 77098725 |  | Cal Std | C4H9NO3 | 0.091 | 3.906 | 851019493 | Calibration Amount Diff 40.406% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_C9 | 36812656 |  | Cal Std | C4H9NO3 | 0.044 | 1.953 | 828357515 | Calibration Amount Diff 51.963% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_C10 | 21229362 |  | Cal Std | C4H9NO3 | 0.024 | 0.977 | 874458043 | Calibration Amount Diff 91.249% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_C11 | 11247000 |  | Cal Std | C4H9NO3 | 0.013 | 0.488 | 890529165 | Calibration Amount Diff 152.759% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_C12 | 5220201 |  | Cal Std | C4H9NO3 | 0.006 | 0.244 | 870741924 | Calibration Amount Diff 257.273% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_A1 | 4512637144 |  | Cal Std | C4H9NO3 | 8.756 | 500 | 515367407
Threonine neg | 9.15 | Target Compound | KD082321_082521_A2 | 1509111889 |  | Cal Std | C4H9NO3 | 4.309 | 250 | 350190389
Threonine neg | 9.15 | Target Compound | KD082321_082521_A3 | 765864214 |  | Cal Std | C4H9NO3 | 2.172 | 125 | 352605236
Threonine neg | 9.15 | Target Compound | KD082321_082521_A4 | 591686302 |  | Cal Std | C4H9NO3 | 1.112 | 62.5 | 532273159
Threonine neg | 9.15 | Target Compound | KD082321_082521_A5 | 360309547 |  | Cal Std | C4H9NO3 | 0.583 | 31.25 | 618189799
Threonine neg | 9.15 | Target Compound | KD082321_082521_A6 | 214061746 |  | Cal Std | C4H9NO3 | 0.302 | 15.625 | 707951334
Threonine neg | 9.15 | Target Compound | KD082321_082521_A7 | 120059746 |  | Cal Std | C4H9NO3 | 0.156 | 7.813 | 771633946
Threonine neg | 9.15 | Target Compound | KD082321_082521_A8 | 75075843 |  | Cal Std | C4H9NO3 | 0.089 | 3.906 | 844209649 | Calibration Amount Diff 38.081% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_A9 | 37165517 |  | Cal Std | C4H9NO3 | 0.046 | 1.953 | 809929037 | Calibration Amount Diff 56.002% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_A10 | 20570466 |  | Cal Std | C4H9NO3 | 0.025 | 0.977 | 831772178 | Calibration Amount Diff 93.781% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_A11 | 11079325 |  | Cal Std | C4H9NO3 | 0.013 | 0.488 | 827634192 | Calibration Amount Diff 161.218% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD082321_082521_A12 | 6609588 |  | Cal Std | C4H9NO3 | 0.008 | 0.244 | 847807090 | Calibration Amount Diff 297.515% (absolute) > max calibration amount diff 20%

## Threonine U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B1 | 508521781 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B2 | 420840241 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B3 | 320650001 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0 | 9.242:Peak area 320650001.423 is out of bounds (ISTD Minimum Recovery 337769513.666 and ISTD Max Recovery 1013308540.999)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B4 | 494964612 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B5 | 619974022 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B6 | 731355634 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B7 | 734177995 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B8 | 765969141 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B9 | 819877710 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B10 | 883497686 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B11 | 864254117 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_B12 | 907329735 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C1 | 493587905 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C2 | 402203207 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C3 | 342086693 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C4 | 505832522 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C5 | 621861392 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C6 | 737983097 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C7 | 819767936 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C8 | 851019493 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C9 | 828357515 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C10 | 874458043 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C11 | 890529165 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_C12 | 870741924 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A1 | 515367407 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A2 | 350190389 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A3 | 352605236 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A4 | 532273159 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A5 | 618189799 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A6 | 707951334 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A7 | 771633946 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A8 | 844209649 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A9 | 809929037 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A10 | 831772178 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A11 | 827634192 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD082321_082521_A12 | 847807090 |  | Cal Std | [13]C4H9[15]NO3 | 0 | 1 | 0

## Tryptophan pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B1 | 36954072369 |  | Cal Std | C11H12N2O2 | 6.894 | 500 | 5360352531
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B2 | 20206218393 |  | Cal Std | C11H12N2O2 | 3.234 | 250 | 6248328591
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B3 | 11733551084 |  | Cal Std | C11H12N2O2 | 1.748 | 125 | 6713972000
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B4 | 6399150914 |  | Cal Std | C11H12N2O2 | 0.906 | 62.5 | 7060785949
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B5 | 3486337511 |  | Cal Std | C11H12N2O2 | 0.475 | 31.25 | 7341739034
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B6 | 1806733533 |  | Cal Std | C11H12N2O2 | 0.238 | 15.625 | 7597569783
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B7 | 924748996 |  | Cal Std | C11H12N2O2 | 0.126 | 7.813 | 7311490442
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B8 | 480223734 |  | Cal Std | C11H12N2O2 | 0.068 | 3.906 | 7072405805
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B9 | 244503819 |  | Cal Std | C11H12N2O2 | 0.034 | 1.953 | 7106465486 | Calibration Amount Diff -58.665% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B10 | 128682182 |  | Cal Std | C11H12N2O2 | 0.018 | 0.977 | 7305820658 | Calibration Amount Diff -148.072% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B11 | 59910537 |  | Cal Std | C11H12N2O2 | 0.008 | 0.488 | 7244588200 | Calibration Amount Diff -341.848% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_B12 | 26540063 |  | Cal Std | C11H12N2O2 | 0.004 | 0.244 | 7259890528 | Calibration Amount Diff -727.493% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C1 | 35777829258 |  | Cal Std | C11H12N2O2 | 6.692 | 500 | 5346044270
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C2 | 18917574629 |  | Cal Std | C11H12N2O2 | 3.204 | 250 | 5903519419
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C3 | 12106239991 |  | Cal Std | C11H12N2O2 | 1.728 | 125 | 7006383195
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C4 | 6477371844 |  | Cal Std | C11H12N2O2 | 0.905 | 62.5 | 7160277180
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C5 | 3282989335 |  | Cal Std | C11H12N2O2 | 0.472 | 31.25 | 6953428945
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C6 | 1725097295 |  | Cal Std | C11H12N2O2 | 0.243 | 15.625 | 7085128664
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C7 | 846976510 |  | Cal Std | C11H12N2O2 | 0.116 | 7.813 | 7279484569
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C8 | 519698486 |  | Cal Std | C11H12N2O2 | 0.07 | 3.906 | 7434211470
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C9 | 244570708 |  | Cal Std | C11H12N2O2 | 0.034 | 1.953 | 7282506917 | Calibration Amount Diff -61.867% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C10 | 127782000 |  | Cal Std | C11H12N2O2 | 0.017 | 0.977 | 7337422122 | Calibration Amount Diff -149.618% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C11 | 61977277 |  | Cal Std | C11H12N2O2 | 0.008 | 0.488 | 7546417965 | Calibration Amount Diff -342.734% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_C12 | 28183363 |  | Cal Std | C11H12N2O2 | 0.004 | 0.244 | 7157597787 | Calibration Amount Diff -718.709% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A1 | 35664457405 |  | Cal Std | C11H12N2O2 | 6.318 | 500 | 5644679809
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A2 | 18941182775 |  | Cal Std | C11H12N2O2 | 3.108 | 250 | 6093836752
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A3 | 11517414266 |  | Cal Std | C11H12N2O2 | 1.749 | 125 | 6583293120
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A4 | 6355854662 |  | Cal Std | C11H12N2O2 | 0.913 | 62.5 | 6959347021
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A5 | 3215328836 |  | Cal Std | C11H12N2O2 | 0.467 | 31.25 | 6886379511
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A6 | 1705708677 |  | Cal Std | C11H12N2O2 | 0.238 | 15.625 | 7153287172
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A7 | 883990344 |  | Cal Std | C11H12N2O2 | 0.123 | 7.813 | 7212680963
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A8 | 534241186 |  | Cal Std | C11H12N2O2 | 0.069 | 3.906 | 7757482002
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A9 | 248772484 |  | Cal Std | C11H12N2O2 | 0.035 | 1.953 | 7188742502 | Calibration Amount Diff -57.886% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A10 | 125333362 |  | Cal Std | C11H12N2O2 | 0.018 | 0.977 | 7093810129 | Calibration Amount Diff -147.649% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A11 | 58648856 |  | Cal Std | C11H12N2O2 | 0.008 | 0.488 | 7020884015 | Calibration Amount Diff -340.542% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD082321_082521_A12 | 29237061 |  | Cal Std | C11H12N2O2 | 0.004 | 0.244 | 7296772461 | Calibration Amount Diff -716.55% (absolute) > max calibration amount diff 20%

## Tryptophan U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B1 | 5360352531 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B2 | 6248328591 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B3 | 6713972000 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B4 | 7060785949 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B5 | 7341739034 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B6 | 7597569783 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B7 | 7311490442 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B8 | 7072405805 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B9 | 7106465486 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B10 | 7305820658 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B11 | 7244588200 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_B12 | 7259890528 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C1 | 5346044270 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C2 | 5903519419 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C3 | 7006383195 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C4 | 7160277180 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C5 | 6953428945 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C6 | 7085128664 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C7 | 7279484569 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C8 | 7434211470 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C9 | 7282506917 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C10 | 7337422122 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C11 | 7546417965 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_C12 | 7157597787 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A1 | 5644679809 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A2 | 6093836752 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A3 | 6583293120 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A4 | 6959347021 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A5 | 6886379511 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A6 | 7153287172 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A7 | 7212680963 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A8 | 7757482002 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A9 | 7188742502 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A10 | 7093810129 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A11 | 7020884015 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD082321_082521_A12 | 7296772461 |  | Cal Std | [13]C11H12[15]N2O2 | 0 | 1 | 0

## Tyrosine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B1 | 16718112598 |  | Cal Std | C9H11NO3 | 8.641 | 500 | 1934772816
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B2 | 11684742288 |  | Cal Std | C9H11NO3 | 3.954 | 250 | 2954985257
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B3 | 7301145631 |  | Cal Std | C9H11NO3 | 2.143 | 125 | 3407128962
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B4 | 4093981110 |  | Cal Std | C9H11NO3 | 1.111 | 62.5 | 3684777324
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B5 | 2325043445 |  | Cal Std | C9H11NO3 | 0.571 | 31.25 | 4071252527
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B6 | 1198149953 |  | Cal Std | C9H11NO3 | 0.292 | 15.625 | 4108995823
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B7 | 662964534 |  | Cal Std | C9H11NO3 | 0.154 | 7.813 | 4302191641
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B8 | 349347428 |  | Cal Std | C9H11NO3 | 0.084 | 3.906 | 4173744742
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B9 | 172522762 |  | Cal Std | C9H11NO3 | 0.042 | 1.953 | 4112094140 | Calibration Amount Diff -23.34% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B10 | 90836336 |  | Cal Std | C9H11NO3 | 0.021 | 0.977 | 4241951375 | Calibration Amount Diff -74.984% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B11 | 43744555 |  | Cal Std | C9H11NO3 | 0.01 | 0.488 | 4211023959 | Calibration Amount Diff -187.71% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_B12 | 19080086 |  | Cal Std | C9H11NO3 | 0.005 | 0.244 | 4193329448 | Calibration Amount Diff -421.341% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C1 | 16313074451 |  | Cal Std | C9H11NO3 | 8.312 | 500 | 1962651524
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C2 | 10790466958 |  | Cal Std | C9H11NO3 | 3.89 | 250 | 2773775062
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C3 | 7347426035 |  | Cal Std | C9H11NO3 | 2.119 | 125 | 3467062372
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C4 | 4061396361 |  | Cal Std | C9H11NO3 | 1.099 | 62.5 | 3695513717
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C5 | 2196447770 |  | Cal Std | C9H11NO3 | 0.57 | 31.25 | 3854868198
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C6 | 1183694304 |  | Cal Std | C9H11NO3 | 0.296 | 15.625 | 4003039505
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C7 | 598510741 |  | Cal Std | C9H11NO3 | 0.144 | 7.813 | 4170363038
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C8 | 362413821 |  | Cal Std | C9H11NO3 | 0.086 | 3.906 | 4222946807
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C9 | 170773849 |  | Cal Std | C9H11NO3 | 0.041 | 1.953 | 4193236082 | Calibration Amount Diff -27.178% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C10 | 92411167 |  | Cal Std | C9H11NO3 | 0.022 | 0.977 | 4236320499 | Calibration Amount Diff -72.486% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C11 | 48154714 |  | Cal Std | C9H11NO3 | 0.011 | 0.488 | 4469069842 | Calibration Amount Diff -182.874% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_C12 | 20810446 |  | Cal Std | C9H11NO3 | 0.005 | 0.244 | 4252559867 | Calibration Amount Diff -412.755% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A1 | 16333167406 |  | Cal Std | C9H11NO3 | 7.972 | 500 | 2048816976
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A2 | 10997424896 |  | Cal Std | C9H11NO3 | 3.78 | 250 | 2909208070
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A3 | 6853270167 |  | Cal Std | C9H11NO3 | 2.145 | 125 | 3194922335
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A4 | 4005310124 |  | Cal Std | C9H11NO3 | 1.113 | 62.5 | 3597266314
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A5 | 2162125055 |  | Cal Std | C9H11NO3 | 0.566 | 31.25 | 3817263779
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A6 | 1146455623 |  | Cal Std | C9H11NO3 | 0.294 | 15.625 | 3893628006
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A7 | 620235044 |  | Cal Std | C9H11NO3 | 0.149 | 7.813 | 4160167505
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A8 | 371213565 |  | Cal Std | C9H11NO3 | 0.084 | 3.906 | 4398959770
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A9 | 167941848 |  | Cal Std | C9H11NO3 | 0.041 | 1.953 | 4048447547 | Calibration Amount Diff -24.814% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A10 | 87945177 |  | Cal Std | C9H11NO3 | 0.022 | 0.977 | 4001974630 | Calibration Amount Diff -71.478% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A11 | 43398335 |  | Cal Std | C9H11NO3 | 0.01 | 0.488 | 4151792988 | Calibration Amount Diff -186.901% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD082321_082521_A12 | 21368540 |  | Cal Std | C9H11NO3 | 0.005 | 0.244 | 4185440342 | Calibration Amount Diff -407.461% (absolute) > max calibration amount diff 20%

## Tyrosine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B1 | 1934772816 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B2 | 2954985257 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B3 | 3407128962 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B4 | 3684777324 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B5 | 4071252527 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B6 | 4108995823 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B7 | 4302191641 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B8 | 4173744742 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B9 | 4112094140 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B10 | 4241951375 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B11 | 4211023959 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_B12 | 4193329448 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C1 | 1962651524 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C2 | 2773775062 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C3 | 3467062372 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C4 | 3695513717 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C5 | 3854868198 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C6 | 4003039505 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C7 | 4170363038 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C8 | 4222946807 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C9 | 4193236082 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C10 | 4236320499 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C11 | 4469069842 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_C12 | 4252559867 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A1 | 2048816976 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A2 | 2909208070 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A3 | 3194922335 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A4 | 3597266314 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A5 | 3817263779 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A6 | 3893628006 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A7 | 4160167505 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A8 | 4398959770 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A9 | 4048447547 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A10 | 4001974630 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A11 | 4151792988 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD082321_082521_A12 | 4185440342 |  | Cal Std | [13]C9H11[15]NO3 | 0 | 1 | 0

## Valine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Valine pos | 7.85 | Target Compound | KD082321_082521_B1 | 16712007421 |  | Cal Std | C5H11NO2 | 9.275 | 500 | 1801892298
Valine pos | 7.85 | Target Compound | KD082321_082521_B2 | 10385366207 |  | Cal Std | C5H11NO2 | 4.255 | 250 | 2440469598
Valine pos | 7.85 | Target Compound | KD082321_082521_B3 | 6716025393 |  | Cal Std | C5H11NO2 | 2.279 | 125 | 2946874819
Valine pos | 7.85 | Target Compound | KD082321_082521_B4 | 4051177031 |  | Cal Std | C5H11NO2 | 1.169 | 62.5 | 3466825665
Valine pos | 7.85 | Target Compound | KD082321_082521_B5 | 2521615040 |  | Cal Std | C5H11NO2 | 0.606 | 31.25 | 4160832896
Valine pos | 7.85 | Target Compound | KD082321_082521_B6 | 1389061630 |  | Cal Std | C5H11NO2 | 0.308 | 15.625 | 4505079926
Valine pos | 7.85 | Target Compound | KD082321_082521_B7 | 807938332 |  | Cal Std | C5H11NO2 | 0.167 | 7.813 | 4841848294
Valine pos | 7.85 | Target Compound | KD082321_082521_B8 | 437346977 |  | Cal Std | C5H11NO2 | 0.093 | 3.906 | 4691095433
Valine pos | 7.85 | Target Compound | KD082321_082521_B9 | 258386523 |  | Cal Std | C5H11NO2 | 0.051 | 1.953 | 5079820854
Valine pos | 7.85 | Target Compound | KD082321_082521_B10 | 124822791 |  | Cal Std | C5H11NO2 | 0.024 | 0.977 | 5111989524 | Calibration Amount Diff -86.092% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_B11 | 76326639 |  | Cal Std | C5H11NO2 | 0.015 | 0.488 | 5049385016 | Calibration Amount Diff -181.395% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_B12 | 43951529 |  | Cal Std | C5H11NO2 | 0.009 | 0.244 | 5070760371 | Calibration Amount Diff -414.251% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_C1 | 16367361862 |  | Cal Std | C5H11NO2 | 8.871 | 500 | 1844973529
Valine pos | 7.85 | Target Compound | KD082321_082521_C2 | 9918237147 |  | Cal Std | C5H11NO2 | 4.211 | 250 | 2355363841
Valine pos | 7.85 | Target Compound | KD082321_082521_C3 | 6732682176 |  | Cal Std | C5H11NO2 | 2.262 | 125 | 2976312180
Valine pos | 7.85 | Target Compound | KD082321_082521_C4 | 4056325229 |  | Cal Std | C5H11NO2 | 1.168 | 62.5 | 3472223589
Valine pos | 7.85 | Target Compound | KD082321_082521_C5 | 2348681301 |  | Cal Std | C5H11NO2 | 0.607 | 31.25 | 3870600498
Valine pos | 7.85 | Target Compound | KD082321_082521_C6 | 1396074746 |  | Cal Std | C5H11NO2 | 0.313 | 15.625 | 4455835345
Valine pos | 7.85 | Target Compound | KD082321_082521_C7 | 690375323 |  | Cal Std | C5H11NO2 | 0.152 | 7.813 | 4552538011
Valine pos | 7.85 | Target Compound | KD082321_082521_C8 | 475775201 |  | Cal Std | C5H11NO2 | 0.095 | 3.906 | 5026508038
Valine pos | 7.85 | Target Compound | KD082321_082521_C9 | 236477709 |  | Cal Std | C5H11NO2 | 0.046 | 1.953 | 5110271014 | Calibration Amount Diff -28.902% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_C10 | 137046758 |  | Cal Std | C5H11NO2 | 0.027 | 0.977 | 5112693311 | Calibration Amount Diff -72.087% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_C11 | 72266411 |  | Cal Std | C5H11NO2 | 0.014 | 0.488 | 5173781253 | Calibration Amount Diff -194.88% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_C12 | 43337756 |  | Cal Std | C5H11NO2 | 0.008 | 0.244 | 5310912983 | Calibration Amount Diff -426.172% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_A1 | 16379233799 |  | Cal Std | C5H11NO2 | 8.316 | 500 | 1969610287
Valine pos | 7.85 | Target Compound | KD082321_082521_A2 | 9931865361 |  | Cal Std | C5H11NO2 | 4.083 | 250 | 2432687515
Valine pos | 7.85 | Target Compound | KD082321_082521_A3 | 6515574341 |  | Cal Std | C5H11NO2 | 2.263 | 125 | 2878666545
Valine pos | 7.85 | Target Compound | KD082321_082521_A4 | 4102256686 |  | Cal Std | C5H11NO2 | 1.178 | 62.5 | 3482472258
Valine pos | 7.85 | Target Compound | KD082321_082521_A5 | 2328998556 |  | Cal Std | C5H11NO2 | 0.594 | 31.25 | 3923715517
Valine pos | 7.85 | Target Compound | KD082321_082521_A6 | 1417733855 |  | Cal Std | C5H11NO2 | 0.313 | 15.625 | 4532741152
Valine pos | 7.85 | Target Compound | KD082321_082521_A7 | 779797693 |  | Cal Std | C5H11NO2 | 0.16 | 7.813 | 4878190806
Valine pos | 7.85 | Target Compound | KD082321_082521_A8 | 478138082 |  | Cal Std | C5H11NO2 | 0.091 | 3.906 | 5245363304
Valine pos | 7.85 | Target Compound | KD082321_082521_A9 | 247195317 |  | Cal Std | C5H11NO2 | 0.049 | 1.953 | 5023378400 | Calibration Amount Diff -20.292% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_A10 | 140292187 |  | Cal Std | C5H11NO2 | 0.028 | 0.977 | 5023085844 | Calibration Amount Diff -65.492% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_A11 | 92336309 |  | Cal Std | C5H11NO2 | 0.018 | 0.488 | 5093679812 | Calibration Amount Diff -146.027% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD082321_082521_A12 | 43982651 |  | Cal Std | C5H11NO2 | 0.009 | 0.244 | 5023193887 | Calibration Amount Diff -412.178% (absolute) > max calibration amount diff 20%

## Valine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B1 | 1801892298 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0 | 7.853:Peak area 1801892297.771 is out of bounds (ISTD Minimum Recovery 2054662133.471 and ISTD Max Recovery 6163986400.414)
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B2 | 2440469598 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B3 | 2946874819 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B4 | 3466825665 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B5 | 4160832896 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B6 | 4505079926 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B7 | 4841848294 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B8 | 4691095433 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B9 | 5079820854 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B10 | 5111989524 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B11 | 5049385016 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_B12 | 5070760371 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C1 | 1844973529 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0 | 7.856:Peak area 1844973529.360 is out of bounds (ISTD Minimum Recovery 2054662133.471 and ISTD Max Recovery 6163986400.414)
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C2 | 2355363841 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C3 | 2976312180 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C4 | 3472223589 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C5 | 3870600498 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C6 | 4455835345 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C7 | 4552538011 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C8 | 5026508038 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C9 | 5110271014 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C10 | 5112693311 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C11 | 5173781253 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_C12 | 5310912983 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A1 | 1969610287 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0 | 7.880:Peak area 1969610286.637 is out of bounds (ISTD Minimum Recovery 2054662133.471 and ISTD Max Recovery 6163986400.414)
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A2 | 2432687515 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A3 | 2878666545 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A4 | 3482472258 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A5 | 3923715517 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A6 | 4532741152 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A7 | 4878190806 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A8 | 5245363304 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A9 | 5023378400 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A10 | 5023085844 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A11 | 5093679812 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD082321_082521_A12 | 5023193887 |  | Cal Std | [13]C5H11[15]NO2 | 0 | 1 | 0