## Adenine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Adenine pos | 5.69 | Target Compound | KD_050321_050321_AAmix | 3032383 |  | Unknown | C5H5N5 | 47.7 | 0 | 63572
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D1 | 42335608368 |  | Cal Std | C5H5N5 | 66.225 | 100 | 639267813
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D2 | 32261477520 |  | Cal Std | C5H5N5 | 30.585 | 50 | 1054822521
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D3 | 23323510967 |  | Cal Std | C5H5N5 | 15.922 | 25 | 1464879705
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D4 | 14159548991 |  | Cal Std | C5H5N5 | 7.591 | 12.5 | 1865202572
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D5 | 8352251293 |  | Cal Std | C5H5N5 | 3.637 | 6.25 | 2296716150
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D6 | 4591197685 |  | Cal Std | C5H5N5 | 1.761 | 3.125 | 2607732133
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D7 | 2049728008 |  | Cal Std | C5H5N5 | 0.774 | 1.563 | 2649716419
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D8 | 1189352141 |  | Cal Std | C5H5N5 | 0.466 | 0.781 | 2554338772
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D9 | 576297901 |  | Cal Std | C5H5N5 | 0.228 | 0.391 | 2528481499
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D10 | 304121996 |  | Cal Std | C5H5N5 | 0.117 | 0.195 | 2600935629
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D11 | 141900343 |  | Cal Std | C5H5N5 | 0.058 | 0.098 | 2444467893
Adenine pos | 5.69 | Target Compound | KD_050321_050321_D12 | 71872959 |  | Cal Std | C5H5N5 | 0.03 | 0.049 | 2402667832
Adenine pos | 5.69 | Target Compound | KD_050321_050321_P1 | 9976297 | pig1 | Unknown | C5H5N5 | 0.004 | 0 | 2257551673
Adenine pos | 5.69 | Target Compound | KD_050321_050321_P2 | 7391583 | pig2 | Unknown | C5H5N5 | 0.003 | 0 | 2223769328
Adenine pos | 5.69 | Target Compound | KD_050321_050321_P3 | 7895622 | pig3 | Unknown | C5H5N5 | 0.004 | 0 | 2187776820
Adenine pos | 5.69 | Target Compound | KD_050321_050321_P12 | 392862 |  | Unknown | C5H5N5 | 3.889 | 0 | 101028
Adenine pos | 5.69 | Target Compound | KD_050321_050321_P13 | 32491705556 |  | Unknown | C5H5N5 | 0 | 0 | 0
Adenine pos | 5.69 | Target Compound | KD_050321_050321_P14 | 5495596741 |  | Unknown | C5H5N5 | 62780.023 | 0 | 87537
Adenine pos | 5.69 | Target Compound | KD_050321_050321_P15 | 617116448 |  | Unknown | C5H5N5 | 536.995 | 0 | 1149202

## Cytosine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_AAmix | 722716 |  | Unknown | C4H5N3O | 0 | 0 | 0
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D1 | 15110651694 |  | Cal Std | C4H5N3O | 577.641 | 100 | 26159222
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D2 | 9977841662 |  | Cal Std | C4H5N3O | 293.579 | 50 | 33986900
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D3 | 6692836154 |  | Cal Std | C4H5N3O | 136.069 | 25 | 49187197
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D4 | 4061447099 |  | Cal Std | C4H5N3O | 64.966 | 12.5 | 62516032
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D5 | 2460638474 |  | Cal Std | C4H5N3O | 31.321 | 6.25 | 78562875
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D6 | 1441104248 |  | Cal Std | C4H5N3O | 17.021 | 3.125 | 84663867
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D7 | 803683932 |  | Cal Std | C4H5N3O | 9.406 | 1.563 | 85445099
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D8 | 502540795 |  | Cal Std | C4H5N3O | 5.153 | 0.781 | 97527200
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D9 | 244391421 |  | Cal Std | C4H5N3O | 2.565 | 0.391 | 95297052
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D10 | 127026468 |  | Cal Std | C4H5N3O | 1.301 | 0.195 | 97642146
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D11 | 64119165 |  | Cal Std | C4H5N3O | 0.679 | 0.098 | 94449770
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_D12 | 32669607 |  | Cal Std | C4H5N3O | 0.351 | 0.049 | 93122180
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_P1 | 20954130 | pig1 | Unknown | C4H5N3O | 0.368 | 0 | 56926617
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_P2 | 93242766 | pig2 | Unknown | C4H5N3O | 1.633 | 0 | 57096724
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_P3 | 38086426 | pig3 | Unknown | C4H5N3O | 0.654 | 0 | 58271597
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_P12 | 1022992 |  | Unknown | C4H5N3O | 0 | 0 | 0
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_P13 | 14587263924 |  | Unknown | C4H5N3O | 0 | 0 | 0
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C4H5N3O | 0 | 0 | 0
Cytosine pos | 7.01 | Target Compound | KD_050321_050321_P15 | 715350 |  | Unknown | C4H5N3O | 0 | 0 | 0

## Guanine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Guanine pos | 7.83 | Target Compound | KD_050321_050321_AAmix | 88018 |  | Unknown | C5H5N5O | 0 | 0 | 0
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D1 | 7357190500 |  | Cal Std | C5H5N5O | 869.323 | 100 | 8463128
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D2 | 4822112791 |  | Cal Std | C5H5N5O | 455.654 | 50 | 10582832
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D3 | 3237257357 |  | Cal Std | C5H5N5O | 226.554 | 25 | 14289111
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D4 | 1898883659 |  | Cal Std | C5H5N5O | 79.155 | 12.5 | 23989572
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D5 | 1192402607 |  | Cal Std | C5H5N5O | 32.418 | 6.25 | 36782551
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D6 | 636239159 |  | Cal Std | C5H5N5O | 14.103 | 3.125 | 45114724
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D7 | 269322022 |  | Cal Std | C5H5N5O | 5.965 | 1.563 | 45150085
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D8 | 154773656 |  | Cal Std | C5H5N5O | 3.372 | 0.781 | 45898929
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D9 | 76013099 |  | Cal Std | C5H5N5O | 1.632 | 0.391 | 46572384
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D10 | 37071529 |  | Cal Std | C5H5N5O | 0.738 | 0.195 | 50207592
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D11 | 16066095 |  | Cal Std | C5H5N5O | 0.364 | 0.098 | 44162872
Guanine pos | 7.83 | Target Compound | KD_050321_050321_D12 | 8414503 |  | Cal Std | C5H5N5O | 0.17 | 0.049 | 49520833
Guanine pos | 7.83 | Target Compound | KD_050321_050321_P1 | 63371694 | pig1 | Unknown | C5H5N5O | 2.032 | 0 | 31193531
Guanine pos | 7.83 | Target Compound | KD_050321_050321_P2 | 53146471 | pig2 | Unknown | C5H5N5O | 1.434 | 0 | 37058276
Guanine pos | 7.83 | Target Compound | KD_050321_050321_P3 | 80589736 | pig3 | Unknown | C5H5N5O | 2.373 | 0 | 33967989
Guanine pos | 7.83 | Target Compound | KD_050321_050321_P12 | 6726971140 |  | Unknown | C5H5N5O | 0 | 0 | 0
Guanine pos | 7.83 | Target Compound | KD_050321_050321_P13 | 9143936 |  | Unknown | C5H5N5O | 0 | 0 | 0
Guanine pos | 7.83 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C5H5N5O | 0 | 0 | 0
Guanine pos | 7.83 | Target Compound | KD_050321_050321_P15 | 98698785 |  | Unknown | C5H5N5O | 0 | 0 | 0

## Thymine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Thymine neg | 4.04 | Target Compound | KD_050321_050321_AAmix | 3442414 |  | Unknown | C5H6N2O2 | 148.326 | 0 | 23208
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D1 | 5865324996 |  | Cal Std | C5H6N2O2 | 175.201 | 100 | 33477645
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D2 | 3072965206 |  | Cal Std | C5H6N2O2 | 83.207 | 50 | 36931732
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D3 | 1809549906 |  | Cal Std | C5H6N2O2 | 49.453 | 25 | 36591525
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D4 | 904714652 |  | Cal Std | C5H6N2O2 | 25.568 | 12.5 | 35384600
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D5 | 511141419 |  | Cal Std | C5H6N2O2 | 12.921 | 6.25 | 39557929
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D6 | 271470892 |  | Cal Std | C5H6N2O2 | 6.528 | 3.125 | 41582467
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D7 | 148081642 |  | Cal Std | C5H6N2O2 | 3.473 | 1.563 | 42635242
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D8 | 90564690 |  | Cal Std | C5H6N2O2 | 2.175 | 0.781 | 41646380
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D9 | 42498860 |  | Cal Std | C5H6N2O2 | 1.063 | 0.391 | 39984582
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D10 | 20539499 |  | Cal Std | C5H6N2O2 | 0.491 | 0.195 | 41817295
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D11 | 12482047 |  | Cal Std | C5H6N2O2 | 0.29 | 0.098 | 42975669
Thymine neg | 4.04 | Target Compound | KD_050321_050321_D12 | 6803666 |  | Cal Std | C5H6N2O2 | 0.172 | 0.049 | 39467864
Thymine neg | 4.04 | Target Compound | KD_050321_050321_P1 | 30513859 | pig1 | Unknown | C5H6N2O2 | 0.736 | 0 | 41455958
Thymine neg | 4.04 | Target Compound | KD_050321_050321_P2 | 18890701 | pig2 | Unknown | C5H6N2O2 | 0.477 | 0 | 39589999
Thymine neg | 4.04 | Target Compound | KD_050321_050321_P3 | 17685847 | pig3 | Unknown | C5H6N2O2 | 0.45 | 0 | 39333750
Thymine neg | 4.04 | Target Compound | KD_050321_050321_P12 | 0 |  | Unknown | C5H6N2O2 | 0 | 0 | 66980
Thymine neg | 4.04 | Target Compound | KD_050321_050321_P13 | 2529123934 |  | Unknown | C5H6N2O2 | 0 | 0 | 0
Thymine neg | 4.04 | Target Compound | KD_050321_050321_P14 | 3870516893 |  | Unknown | C5H6N2O2 | 0 | 0 | 0
Thymine neg | 4.04 | Target Compound | KD_050321_050321_P15 | 6651455 |  | Unknown | C5H6N2O2 | 304.851 | 0 | 21819

## Uracil neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Uracil neg | 4.89 | Target Compound | KD_050321_050321_AAmix | 0 |  | Unknown | C4H4N2O2 | 0 | 0 | 50092
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D1 | 4253330132 |  | Cal Std | C4H4N2O2 | 52.677 | 100 | 80743963
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D2 | 2322851394 |  | Cal Std | C4H4N2O2 | 27.372 | 50 | 84861217
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D3 | 1514253909 |  | Cal Std | C4H4N2O2 | 15.637 | 25 | 96837221
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D4 | 792296055 |  | Cal Std | C4H4N2O2 | 7.636 | 12.5 | 103756768
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D5 | 427230220 |  | Cal Std | C4H4N2O2 | 4.125 | 6.25 | 103569123
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D6 | 264073631 |  | Cal Std | C4H4N2O2 | 2.243 | 3.125 | 117732561
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D7 | 130832754 |  | Cal Std | C4H4N2O2 | 1.119 | 1.563 | 116918319
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D8 | 77427826 |  | Cal Std | C4H4N2O2 | 0.67 | 0.781 | 115627748
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D9 | 38044914 |  | Cal Std | C4H4N2O2 | 0.312 | 0.391 | 121931647
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D10 | 12373075 |  | Cal Std | C4H4N2O2 | 0.136 | 0.195 | 91086627
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D11 | 7037513 |  | Cal Std | C4H4N2O2 | 0.067 | 0.098 | 105568825
Uracil neg | 4.89 | Target Compound | KD_050321_050321_D12 | 4463441 |  | Cal Std | C4H4N2O2 | 0.041 | 0.049 | 107820162
Uracil neg | 4.89 | Target Compound | KD_050321_050321_P1 | 20336396 | pig1 | Unknown | C4H4N2O2 | 0.177 | 0 | 114931325
Uracil neg | 4.89 | Target Compound | KD_050321_050321_P2 | 12834819 | pig2 | Unknown | C4H4N2O2 | 0.114 | 0 | 112202632
Uracil neg | 4.89 | Target Compound | KD_050321_050321_P3 | 14686795 | pig3 | Unknown | C4H4N2O2 | 0.128 | 0 | 114724554
Uracil neg | 4.89 | Target Compound | KD_050321_050321_P12 | 566019 |  | Unknown | C4H4N2O2 | 4.526 | 0 | 125055
Uracil neg | 4.89 | Target Compound | KD_050321_050321_P13 | 2237059793 |  | Unknown | C4H4N2O2 | 0 | 0 | 0
Uracil neg | 4.89 | Target Compound | KD_050321_050321_P14 | 3211184693 |  | Unknown | C4H4N2O2 | 0 | 0 | 0
Uracil neg | 4.89 | Target Compound | KD_050321_050321_P15 | 1187354 |  | Unknown | C4H4N2O2 | 61.254 | 0 | 19384

## Cytidine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_AAmix | 0 |  | Unknown | C9H13N3O5 | 0 | 0 | 0
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D1 | 2991682286 |  | Cal Std | C9H13N3O5 | 257.01 | 100 | 11640322
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D2 | 1934045708 |  | Cal Std | C9H13N3O5 | 98.838 | 50 | 19567892
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D3 | 1263830872 |  | Cal Std | C9H13N3O5 | 47.867 | 25 | 26402949
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D4 | 731458762 |  | Cal Std | C9H13N3O5 | 18.982 | 12.5 | 38534282
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D5 | 439391873 |  | Cal Std | C9H13N3O5 | 9.415 | 6.25 | 46667279
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D6 | 225588943 |  | Cal Std | C9H13N3O5 | 4.792 | 3.125 | 47076947
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D7 | 113254303 |  | Cal Std | C9H13N3O5 | 2.298 | 1.563 | 49277356
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D8 | 61297398 |  | Cal Std | C9H13N3O5 | 1.295 | 0.781 | 47349026
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D9 | 28083750 |  | Cal Std | C9H13N3O5 | 0.566 | 0.391 | 49623980
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D10 | 12375954 |  | Cal Std | C9H13N3O5 | 0.262 | 0.195 | 47159592
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D11 | 5465547 |  | Cal Std | C9H13N3O5 | 0.108 | 0.098 | 50701018
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_D12 | 1778938 |  | Cal Std | C9H13N3O5 | 0.037 | 0.049 | 47553019
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_P1 | 54929725 | pig1 | Unknown | C9H13N3O5 | 1.449 | 0 | 37918780
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_P2 | 61455460 | pig2 | Unknown | C9H13N3O5 | 1.499 | 0 | 40993089
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_P3 | 57615105 | pig3 | Unknown | C9H13N3O5 | 1.641 | 0 | 35100209
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_P12 | 0 |  | Unknown | C9H13N3O5 | 0 | 0 | 0
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C9H13N3O5 | 0 | 0 | 0
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_P14 | 3531312132 |  | Unknown | C9H13N3O5 | 0 | 0 | 0
Cytidine pos | 7.63 | Target Compound | KD_050321_050321_P15 | 28115 |  | Unknown | C9H13N3O5 | 0 | 0 | 0

## Uridine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Uridine neg | 6.19 | Target Compound | KD_050321_050321_AAmix | 0 |  | Unknown | C9H12N2O6 | 0 | 0 | 0
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D1 | 1212409199 |  | Cal Std | C9H12N2O6 | 443.136 | 100 | 2735972
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D2 | 661286925 |  | Cal Std | C9H12N2O6 | 210.344 | 50 | 3143837
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D3 | 429034671 |  | Cal Std | C9H12N2O6 | 81.427 | 25 | 5268937
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D4 | 225857720 |  | Cal Std | C9H12N2O6 | 35.547 | 12.5 | 6353720
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D5 | 132447112 |  | Cal Std | C9H12N2O6 | 15.48 | 6.25 | 8556168
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D6 | 72526368 |  | Cal Std | C9H12N2O6 | 6.8 | 3.125 | 10665599
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D7 | 34183035 |  | Cal Std | C9H12N2O6 | 3.502 | 1.563 | 9759749
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D8 | 22684610 |  | Cal Std | C9H12N2O6 | 1.634 | 0.781 | 13883830
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D9 | 8051567 |  | Cal Std | C9H12N2O6 | 0.641 | 0.391 | 12555245
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D10 | 2712125 |  | Cal Std | C9H12N2O6 | 0.219 | 0.195 | 12399373
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D11 | 1224042 |  | Cal Std | C9H12N2O6 | 0.09 | 0.098 | 13656608
Uridine neg | 6.19 | Target Compound | KD_050321_050321_D12 | 438312 |  | Cal Std | C9H12N2O6 | 0.034 | 0.049 | 13033324
Uridine neg | 6.19 | Target Compound | KD_050321_050321_P1 | 19313413 | pig1 | Unknown | C9H12N2O6 | 9.03 | 0 | 2138840
Uridine neg | 6.19 | Target Compound | KD_050321_050321_P2 | 25825771 | pig2 | Unknown | C9H12N2O6 | 4.077 | 0 | 6333785
Uridine neg | 6.19 | Target Compound | KD_050321_050321_P3 | 18374055 | pig3 | Unknown | C9H12N2O6 | 3.989 | 0 | 4606505
Uridine neg | 6.19 | Target Compound | KD_050321_050321_P12 | 5305 |  | Unknown | C9H12N2O6 | 0 | 0 | 0
Uridine neg | 6.19 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C9H12N2O6 | 0 | 0 | 0
Uridine neg | 6.19 | Target Compound | KD_050321_050321_P14 | 1632086074 |  | Unknown | C9H12N2O6 | 0 | 0 | 0
Uridine neg | 6.19 | Target Compound | KD_050321_050321_P15 | 6236 |  | Unknown | C9H12N2O6 | 0 | 0 | 0

## Hypoxanthine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_AAmix | 282550 |  | Unknown | C5H4N4O | 0 | 0 | 0
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D1 | 11124468239 |  | Cal Std | C5H4N4O | 580.391 | 100 | 19167200
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D2 | 6971221914 |  | Cal Std | C5H4N4O | 243.535 | 50 | 28625191
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D3 | 4938222855 |  | Cal Std | C5H4N4O | 111.704 | 25 | 44208171
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D4 | 2864085293 |  | Cal Std | C5H4N4O | 53.178 | 12.5 | 53858203
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D5 | 1778447410 |  | Cal Std | C5H4N4O | 26.353 | 6.25 | 67485777
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D6 | 979656672 |  | Cal Std | C5H4N4O | 13.819 | 3.125 | 70891206
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D7 | 510587363 |  | Cal Std | C5H4N4O | 7.026 | 1.563 | 72667147
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D8 | 299069698 |  | Cal Std | C5H4N4O | 4.135 | 0.781 | 72319163
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D9 | 151221809 |  | Cal Std | C5H4N4O | 1.954 | 0.391 | 77401720
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D10 | 79847535 |  | Cal Std | C5H4N4O | 0.981 | 0.195 | 81392962
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D11 | 35454564 |  | Cal Std | C5H4N4O | 0.495 | 0.098 | 71602971
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_D12 | 19122556 |  | Cal Std | C5H4N4O | 0.242 | 0.049 | 78878951
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_P1 | 865419116 | pig1 | Unknown | C5H4N4O | 39.289 | 0 | 22026832
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_P2 | 1093700618 | pig2 | Unknown | C5H4N4O | 42.511 | 0 | 25727210
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_P3 | 1134893109 | pig3 | Unknown | C5H4N4O | 51.623 | 0 | 21984259
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_P12 | 3723547 |  | Unknown | C5H4N4O | 0 | 0 | 0
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_P13 | 13817953301 |  | Unknown | C5H4N4O | 0 | 0 | 0
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_P14 | 3760939 |  | Unknown | C5H4N4O | 0 | 0 | 0
Hypoxanthine pos | 6.45 | Target Compound | KD_050321_050321_P15 | 7279422 |  | Unknown | C5H4N4O | 0 | 0 | 0

## Adenosine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_AAmix | 192222 |  | Unknown | C10H13N5O4 | 0 | 0 | 0
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D1 | 14333698598 |  | Cal Std | C10H13N5O4 | 19.846 | 100 | 722239817
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D2 | 9421419829 |  | Cal Std | C10H13N5O4 | 9.785 | 50 | 962799228
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D3 | 6665311695 |  | Cal Std | C10H13N5O4 | 5.461 | 25 | 1220620877
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D4 | 4092383613 |  | Cal Std | C10H13N5O4 | 2.768 | 12.5 | 1478502569
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D5 | 2439485543 |  | Cal Std | C10H13N5O4 | 1.485 | 6.25 | 1642433405
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D6 | 1416268375 |  | Cal Std | C10H13N5O4 | 0.796 | 3.125 | 1780135959
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D7 | 786566968 |  | Cal Std | C10H13N5O4 | 0.412 | 1.563 | 1910559014
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D8 | 421267803 |  | Cal Std | C10H13N5O4 | 0.242 | 0.781 | 1742894159
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D9 | 217292818 |  | Cal Std | C10H13N5O4 | 0.115 | 0.391 | 1894708620
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D10 | 103530764 |  | Cal Std | C10H13N5O4 | 0.055 | 0.195 | 1889164593
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D11 | 53036555 |  | Cal Std | C10H13N5O4 | 0.027 | 0.098 | 1951705095
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_D12 | 19636163 |  | Cal Std | C10H13N5O4 | 0.012 | 0.049 | 1697092170
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_P1 | 1618202 | pig1 | Unknown | C10H13N5O4 | 0.001 | 0 | 1782302890
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_P2 | 1381965 | pig2 | Unknown | C10H13N5O4 | 0.001 | 0 | 1768756858
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_P3 | 1163982 | pig3 | Unknown | C10H13N5O4 | 0.001 | 0 | 1932080273
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_P12 | 48073 |  | Unknown | C10H13N5O4 | 12.811 | 0 | 3753
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_P13 | 5717 |  | Unknown | C10H13N5O4 | 0 | 0 | 0
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_P14 | 21563245875 |  | Unknown | C10H13N5O4 | 0 | 0 | 0
Adenosine pos | 5.5 | Target Compound | KD_050321_050321_P15 | 5093441 |  | Unknown | C10H13N5O4 | 18.678 | 0 | 272691

## Guanosine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_AAmix | 0 |  | Unknown | C10H13N5O5 | 0 | 0 | 0
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D1 | 1061182318 |  | Cal Std | C10H13N5O5 | 91.014 | 100 | 11659583
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D2 | 635993773 |  | Cal Std | C10H13N5O5 | 39.771 | 50 | 15991577
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D3 | 420322673 |  | Cal Std | C10H13N5O5 | 22.496 | 25 | 18684371
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D4 | 235681425 |  | Cal Std | C10H13N5O5 | 8.154 | 12.5 | 28902243
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D5 | 118742086 |  | Cal Std | C10H13N5O5 | 3.504 | 6.25 | 33891535
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D6 | 58526300 |  | Cal Std | C10H13N5O5 | 1.553 | 3.125 | 37686719
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D7 | 29159607 |  | Cal Std | C10H13N5O5 | 0.631 | 1.563 | 46202140
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D8 | 13124279 |  | Cal Std | C10H13N5O5 | 0.288 | 0.781 | 45508000
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D9 | 5048118 |  | Cal Std | C10H13N5O5 | 0.109 | 0.391 | 46167001
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D10 | 2387529 |  | Cal Std | C10H13N5O5 | 0.047 | 0.195 | 51152551
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D11 | 1058237 |  | Cal Std | C10H13N5O5 | 0.021 | 0.098 | 49636706
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_D12 | 383419 |  | Cal Std | C10H13N5O5 | 0.008 | 0.049 | 47988847
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_P1 | 783780 | pig1 | Unknown | C10H13N5O5 | 0.038 | 0 | 20797881
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_P2 | 1708815 | pig2 | Unknown | C10H13N5O5 | 0.082 | 0 | 20714674
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_P3 | 1756477 | pig3 | Unknown | C10H13N5O5 | 0.097 | 0 | 18132545
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_P12 | 0 |  | Unknown | C10H13N5O5 | 0 | 0 | 4180
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_P13 | 2936 |  | Unknown | C10H13N5O5 | 0.955 | 0 | 3076
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_P14 | 2105795016 |  | Unknown | C10H13N5O5 | 0 | 0 | 0
Guanosine pos | 8.2 | Target Compound | KD_050321_050321_P15 | 138493 |  | Unknown | C10H13N5O5 | 29.656 | 0 | 4670

## Inosine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Inosine pos | 7.15 | Target Compound | KD_050321_050321_AAmix | 0 |  | Unknown | C10H12N4O5 | 0 | 0 | 0
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D1 | 1304906205 |  | Cal Std | C10H12N4O5 | 4757.629 | 100 | 274277
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D2 | 791477915 |  | Cal Std | C10H12N4O5 | 2193.721 | 50 | 360792
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D3 | 533966822 |  | Cal Std | C10H12N4O5 | 648.001 | 25 | 824022
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D4 | 294217796 |  | Cal Std | C10H12N4O5 | 271.376 | 12.5 | 1084168
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D5 | 165229508 |  | Cal Std | C10H12N4O5 | 150.225 | 6.25 | 1099877
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D6 | 80730753 |  | Cal Std | C10H12N4O5 | 74.331 | 3.125 | 1086098
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D7 | 40525189 |  | Cal Std | C10H12N4O5 | 33.839 | 1.563 | 1197574
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D8 | 20168205 |  | Cal Std | C10H12N4O5 | 17.404 | 0.781 | 1158843
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D9 | 6968331 |  | Cal Std | C10H12N4O5 | 6.609 | 0.391 | 1054431
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D10 | 2360627 |  | Cal Std | C10H12N4O5 | 2.674 | 0.195 | 882675
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D11 | 1306112 |  | Cal Std | C10H12N4O5 | 1.128 | 0.098 | 1158367
Inosine pos | 7.15 | Target Compound | KD_050321_050321_D12 | 455509 |  | Cal Std | C10H12N4O5 | 0.548 | 0.049 | 831654
Inosine pos | 7.15 | Target Compound | KD_050321_050321_P1 | 8652684 | pig1 | Unknown | C10H12N4O5 | 8.541 | 0 | 1013039
Inosine pos | 7.15 | Target Compound | KD_050321_050321_P2 | 22327988 | pig2 | Unknown | C10H12N4O5 | 25.001 | 0 | 893074
Inosine pos | 7.15 | Target Compound | KD_050321_050321_P3 | 13059922 | pig3 | Unknown | C10H12N4O5 | 55.536 | 0 | 235159
Inosine pos | 7.15 | Target Compound | KD_050321_050321_P12 | 0 |  | Unknown | C10H12N4O5 | 0 | 0 | 0
Inosine pos | 7.15 | Target Compound | KD_050321_050321_P13 | 913845 |  | Unknown | C10H12N4O5 | 0 | 0 | 0
Inosine pos | 7.15 | Target Compound | KD_050321_050321_P14 | 9461312 |  | Unknown | C10H12N4O5 | 0 | 0 | 0
Inosine pos | 7.15 | Target Compound | KD_050321_050321_P15 | 1682963783 |  | Unknown | C10H12N4O5 | 0 | 0 | 0

## Deoxycytidine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_AAmix | 16481 |  | Unknown | C9H13N3O4 | 0 | 0 | 0
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D1 | 3047298679 |  | Cal Std | C9H13N3O4 | 492.454 | 100 | 6187985
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D2 | 1816822354 |  | Cal Std | C9H13N3O4 | 267.737 | 50 | 6785853
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D3 | 1315335887 |  | Cal Std | C9H13N3O4 | 135.78 | 25 | 9687251
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D4 | 736520899 |  | Cal Std | C9H13N3O4 | 54.753 | 12.5 | 13451747
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D5 | 395256699 |  | Cal Std | C9H13N3O4 | 25.52 | 6.25 | 15488307
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D6 | 221138032 |  | Cal Std | C9H13N3O4 | 11.235 | 3.125 | 19682955
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D7 | 120769382 |  | Cal Std | C9H13N3O4 | 4.758 | 1.563 | 25383625
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D8 | 70864131 |  | Cal Std | C9H13N3O4 | 2.532 | 0.781 | 27985839
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D9 | 33669429 |  | Cal Std | C9H13N3O4 | 1.139 | 0.391 | 29563357
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D10 | 14912205 |  | Cal Std | C9H13N3O4 | 0.514 | 0.195 | 29023945
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D11 | 6398383 |  | Cal Std | C9H13N3O4 | 0.225 | 0.098 | 28392741
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_D12 | 2384985 |  | Cal Std | C9H13N3O4 | 0.086 | 0.049 | 27704749
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_P1 | 12014812 | pig1 | Unknown | C9H13N3O4 | 0.533 | 0 | 22527091
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_P2 | 9632689 | pig2 | Unknown | C9H13N3O4 | 0.448 | 0 | 21493102
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_P3 | 13509053 | pig3 | Unknown | C9H13N3O4 | 0.617 | 0 | 21886179
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_P12 | 27520 |  | Unknown | C9H13N3O4 | 0 | 0 | 0
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C9H13N3O4 | 0 | 0 | 0
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_P14 | 3057265921 |  | Unknown | C9H13N3O4 | 0 | 0 | 0
Deoxycytidine pos | 6.62 | Target Compound | KD_050321_050321_P15 | 280411 |  | Unknown | C9H13N3O4 | 0 | 0 | 0

## Aspartate neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_AAmix | 984831504 |  | Unknown | C4H7NO4 | 0 | 0 | 0
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D1 | 871664720 |  | Cal Std | C4H7NO4 | 6.383 | 100 | 136554834
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D2 | 541868975 |  | Cal Std | C4H7NO4 | 3.107 | 50 | 174393560
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D3 | 320048525 |  | Cal Std | C4H7NO4 | 1.753 | 25 | 182530216
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D4 | 168160177 |  | Cal Std | C4H7NO4 | 0.84 | 12.5 | 200264463
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D5 | 83280347 |  | Cal Std | C4H7NO4 | 0.43 | 6.25 | 193493089
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D6 | 41832455 |  | Cal Std | C4H7NO4 | 0.221 | 3.125 | 189433017
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D7 | 20780968 |  | Cal Std | C4H7NO4 | 0.107 | 1.563 | 194588982
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D8 | 11561915 |  | Cal Std | C4H7NO4 | 0.06 | 0.781 | 191428579
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D9 | 5022765 |  | Cal Std | C4H7NO4 | 0.028 | 0.391 | 179021090
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D10 | 2787214 |  | Cal Std | C4H7NO4 | 0.015 | 0.195 | 189441149
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D11 | 999550 |  | Cal Std | C4H7NO4 | 0.006 | 0.098 | 176731730
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_D12 | 713435 |  | Cal Std | C4H7NO4 | 0.004 | 0.049 | 175909866
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_P1 | 33919104 | pig1 | Unknown | C4H7NO4 | 0.133 | 0 | 254091293
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_P2 | 33088399 | pig2 | Unknown | C4H7NO4 | 0.139 | 0 | 237813253
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_P3 | 43234670 | pig3 | Unknown | C4H7NO4 | 0.168 | 0 | 256678086
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_P12 | 152126 |  | Unknown | C4H7NO4 | 0 | 0 | 0
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_P13 | 62712 |  | Unknown | C4H7NO4 | 0 | 0 | 0
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C4H7NO4 | 0 | 0 | 0
Aspartate neg | 10.25 | Target Compound | KD_050321_050321_P15 | 404714 |  | Unknown | C4H7NO4 | 0 | 0 | 0

## Asparagine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_AAmix | 892002054 |  | Unknown | C4H8N2O3 | 0 | 0 | 0
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D1 | 834309490 |  | Cal Std | C4H8N2O3 | 130.436 | 100 | 6396319
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D2 | 492096167 |  | Cal Std | C4H8N2O3 | 61.706 | 50 | 7974820
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D3 | 300513507 |  | Cal Std | C4H8N2O3 | 28.335 | 25 | 10605819
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D4 | 158731935 |  | Cal Std | C4H8N2O3 | 12.491 | 12.5 | 12707811
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D5 | 84303151 |  | Cal Std | C4H8N2O3 | 5.879 | 6.25 | 14340251
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D6 | 44676367 |  | Cal Std | C4H8N2O3 | 2.946 | 3.125 | 15167551
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D7 | 21998741 |  | Cal Std | C4H8N2O3 | 1.435 | 1.563 | 15326593
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D8 | 12059145 |  | Cal Std | C4H8N2O3 | 0.781 | 0.781 | 15440611
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D9 | 5268106 |  | Cal Std | C4H8N2O3 | 0.351 | 0.391 | 15029005
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D10 | 1898147 |  | Cal Std | C4H8N2O3 | 0.126 | 0.195 | 15020849
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D11 | 884764 |  | Cal Std | C4H8N2O3 | 0.063 | 0.098 | 14015180
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_D12 | 0 |  | Cal Std | C4H8N2O3 | 0 | 0.049 | 15955843
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_P1 | 108849475 | pig1 | Unknown | C4H8N2O3 | 20.324 | 0 | 5355736
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_P2 | 87656560 | pig2 | Unknown | C4H8N2O3 | 12.488 | 0 | 7019308
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_P3 | 106524555 | pig3 | Unknown | C4H8N2O3 | 14.123 | 0 | 7542866
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_P12 | 0 |  | Unknown | C4H8N2O3 | 0 | 0 | 0
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_P13 | 136886 |  | Unknown | C4H8N2O3 | 0 | 0 | 0
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_P14 | 100564 |  | Unknown | C4H8N2O3 | 0 | 0 | 0
Asparagine pos | 9.84 | Target Compound | KD_050321_050321_P15 | 189089 |  | Unknown | C4H8N2O3 | 0 | 0 | 0

## Alanine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Alanine pos | 9.58 | Target Compound | KD_050321_050321_AAmix | 796332445 |  | Unknown | C3H7NO2 | 0 | 0 | 0
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D1 | 717370848 |  | Cal Std | C3H7NO2 | 3.291 | 100 | 217975764
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D2 | 421377391 |  | Cal Std | C3H7NO2 | 1.575 | 50 | 267507277
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D3 | 263745169 |  | Cal Std | C3H7NO2 | 0.868 | 25 | 303877702
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D4 | 141362185 |  | Cal Std | C3H7NO2 | 0.403 | 12.5 | 351116620
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D5 | 79160837 |  | Cal Std | C3H7NO2 | 0.217 | 6.25 | 364328075
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D6 | 43490147 |  | Cal Std | C3H7NO2 | 0.112 | 3.125 | 386914885
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D7 | 21608363 |  | Cal Std | C3H7NO2 | 0.056 | 1.563 | 383659482
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D8 | 12499693 |  | Cal Std | C3H7NO2 | 0.033 | 0.781 | 377073722 | Calibration Amount Diff 25.693% (absolute) > max calibration amount diff 20%
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D9 | 5911291 |  | Cal Std | C3H7NO2 | 0.015 | 0.391 | 388499159
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D10 | 3778320 |  | Cal Std | C3H7NO2 | 0.01 | 0.195 | 389212106 | Calibration Amount Diff 47.387% (absolute) > max calibration amount diff 20%
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D11 | 2086322 |  | Cal Std | C3H7NO2 | 0.006 | 0.098 | 375104132 | Calibration Amount Diff 68.022% (absolute) > max calibration amount diff 20%
Alanine pos | 9.58 | Target Compound | KD_050321_050321_D12 | 1832915 |  | Cal Std | C3H7NO2 | 0.005 | 0.049 | 369969310 | Calibration Amount Diff 199.323% (absolute) > max calibration amount diff 20%
Alanine pos | 9.58 | Target Compound | KD_050321_050321_P1 | 1328526529 | pig1 | Unknown | C3H7NO2 | 10.937 | 0 | 121469158
Alanine pos | 9.58 | Target Compound | KD_050321_050321_P2 | 722205067 | pig2 | Unknown | C3H7NO2 | 3.926 | 0 | 183932243
Alanine pos | 9.58 | Target Compound | KD_050321_050321_P3 | 975936185 | pig3 | Unknown | C3H7NO2 | 9.606 | 0 | 101597238
Alanine pos | 9.58 | Target Compound | KD_050321_050321_P12 | 1917120 |  | Unknown | C3H7NO2 | 0 | 0 | 0
Alanine pos | 9.58 | Target Compound | KD_050321_050321_P13 | 215790 |  | Unknown | C3H7NO2 | 0 | 0 | 0
Alanine pos | 9.58 | Target Compound | KD_050321_050321_P14 | 498743 |  | Unknown | C3H7NO2 | 0 | 0 | 0
Alanine pos | 9.58 | Target Compound | KD_050321_050321_P15 | 2900443 |  | Unknown | C3H7NO2 | 0 | 0 | 0

## alpha-Ketoglutarate neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_AAmix | 323209 |  | Unknown | C5H6O5 | 45.425 | 0 | 7115
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D1 | 1411085311 |  | Cal Std | C5H6O5 | 45.933 | 100 | 30720199
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D2 | 842452405 |  | Cal Std | C5H6O5 | 21.953 | 50 | 38375176
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D3 | 504901825 |  | Cal Std | C5H6O5 | 12.101 | 25 | 41723927
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D4 | 263702383 |  | Cal Std | C5H6O5 | 5.966 | 12.5 | 44199717
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D5 | 134430677 |  | Cal Std | C5H6O5 | 3.054 | 6.25 | 44012402
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D6 | 69322562 |  | Cal Std | C5H6O5 | 1.556 | 3.125 | 44541876
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D7 | 34790659 |  | Cal Std | C5H6O5 | 0.762 | 1.563 | 45637855
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D8 | 19031856 |  | Cal Std | C5H6O5 | 0.428 | 0.781 | 44427422
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D9 | 9120362 |  | Cal Std | C5H6O5 | 0.199 | 0.391 | 45901827
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D10 | 4368823 |  | Cal Std | C5H6O5 | 0.091 | 0.195 | 48106990
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D11 | 1892575 |  | Cal Std | C5H6O5 | 0.042 | 0.098 | 45334048
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_D12 | 914201 |  | Cal Std | C5H6O5 | 0.02 | 0.049 | 45780317
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_P1 | 153194744 | pig1 | Unknown | C5H6O5 | 3.635 | 0 | 42146212
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_P2 | 119346243 | pig2 | Unknown | C5H6O5 | 2.98 | 0 | 40045354
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_P3 | 170378610 | pig3 | Unknown | C5H6O5 | 3.904 | 0 | 43641246
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_P12 | 72286 |  | Unknown | C5H6O5 | 9.116 | 0 | 7930
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_P13 | 57273 |  | Unknown | C5H6O5 | 0 | 0 | 0
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C5H6O5 | 0 | 0 | 2969
alpha-Ketoglutarate neg | 10.64 | Target Compound | KD_050321_050321_P15 | 1566289564 |  | Unknown | C5H6O5 | 0 | 0 | 0

## Arginine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Arginine pos | 14.6 | Target Compound | KD_050321_050321_AAmix | 4255422428 |  | Unknown | C6H14N4O2 | 0 | 0 | 0
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D1 | 4110672825 |  | Cal Std | C6H14N4O2 | 40.465 | 100 | 101585962
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D2 | 2589609623 |  | Cal Std | C6H14N4O2 | 19.744 | 50 | 131160273
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D3 | 1895020092 |  | Cal Std | C6H14N4O2 | 11.408 | 25 | 166117121
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D4 | 1055944664 |  | Cal Std | C6H14N4O2 | 5.868 | 12.5 | 179956659
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D5 | 654079285 |  | Cal Std | C6H14N4O2 | 3.161 | 6.25 | 206931037
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D6 | 381837179 |  | Cal Std | C6H14N4O2 | 1.732 | 3.125 | 220498108
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D7 | 202437094 |  | Cal Std | C6H14N4O2 | 0.888 | 1.563 | 228044132
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D8 | 114633171 |  | Cal Std | C6H14N4O2 | 0.542 | 0.781 | 211467944 | Calibration Amount Diff 46.414% (absolute) > max calibration amount diff 20%
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D9 | 55649553 |  | Cal Std | C6H14N4O2 | 0.256 | 0.391 | 217468686 | Calibration Amount Diff 37.916% (absolute) > max calibration amount diff 20%
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D10 | 28029632 |  | Cal Std | C6H14N4O2 | 0.136 | 0.195 | 205447984 | Calibration Amount Diff 47.374% (absolute) > max calibration amount diff 20%
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D11 | 16247286 |  | Cal Std | C6H14N4O2 | 0.073 | 0.098 | 223357861 | Calibration Amount Diff 56.313% (absolute) > max calibration amount diff 20%
Arginine pos | 14.6 | Target Compound | KD_050321_050321_D12 | 8187941 |  | Cal Std | C6H14N4O2 | 0.041 | 0.049 | 199848808 | Calibration Amount Diff 76.064% (absolute) > max calibration amount diff 20%
Arginine pos | 14.6 | Target Compound | KD_050321_050321_P1 | 2134467493 | pig1 | Unknown | C6H14N4O2 | 14.21 | 0 | 150206432
Arginine pos | 14.6 | Target Compound | KD_050321_050321_P2 | 1361522831 | pig2 | Unknown | C6H14N4O2 | 7.803 | 0 | 174490512
Arginine pos | 14.6 | Target Compound | KD_050321_050321_P3 | 1850736381 | pig3 | Unknown | C6H14N4O2 | 11.429 | 0 | 161930091
Arginine pos | 14.6 | Target Compound | KD_050321_050321_P12 | 5524667 |  | Unknown | C6H14N4O2 | 0 | 0 | 0
Arginine pos | 14.6 | Target Compound | KD_050321_050321_P13 | 2704283 |  | Unknown | C6H14N4O2 | 0 | 0 | 0
Arginine pos | 14.6 | Target Compound | KD_050321_050321_P14 | 3183307 |  | Unknown | C6H14N4O2 | 0 | 0 | 0
Arginine pos | 14.6 | Target Compound | KD_050321_050321_P15 | 3206098 |  | Unknown | C6H14N4O2 | 0 | 0 | 0

## Cis-aconitate neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_AAmix | 231970884 |  | Unknown | C6H6O6 | 0 | 0 | 0
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D1 | 234941845 |  | Cal Std | C6H6O6 | 557.711 | 100 | 421261
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D2 | 142102056 |  | Cal Std | C6H6O6 | 140.654 | 50 | 1010298
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D3 | 93602053 |  | Cal Std | C6H6O6 | 40.324 | 25 | 2321231
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D4 | 62555640 |  | Cal Std | C6H6O6 | 11.905 | 12.5 | 5254644
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D5 | 34848764 |  | Cal Std | C6H6O6 | 3.991 | 6.25 | 8731139
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D6 | 19730381 |  | Cal Std | C6H6O6 | 2.002 | 3.125 | 9856322 | Calibration Amount Diff 35.255% (absolute) > max calibration amount diff 20%
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D7 | 9459971 |  | Cal Std | C6H6O6 | 0.733 | 1.563 | 12902056 | Calibration Amount Diff 31.049% (absolute) > max calibration amount diff 20%
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D8 | 0 |  | Cal Std | C6H6O6 | 0 | 0.781 | 12718507
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D9 | 393768 |  | Cal Std | C6H6O6 | 0.035 | 0.391 | 11292393 | Calibration Amount Diff -65.223% (absolute) > max calibration amount diff 20%
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D10 | 431783 |  | Cal Std | C6H6O6 | 0.034 | 0.195 | 12585172 | Calibration Amount Diff -31.36% (absolute) > max calibration amount diff 20%
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D11 | 178239 |  | Cal Std | C6H6O6 | 0.014 | 0.098 | 12316334 | Calibration Amount Diff -41.462% (absolute) > max calibration amount diff 20%
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_D12 | 239057 |  | Cal Std | C6H6O6 | 0.022 | 0.049 | 11016964 | Calibration Amount Diff 74.508% (absolute) > max calibration amount diff 20%
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_P1 | 40891287 | pig1 | Unknown | C6H6O6 | 5.368 | 0 | 7617728
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_P2 | 36996889 | pig2 | Unknown | C6H6O6 | 4.542 | 0 | 8146347
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_P3 | 56597263 | pig3 | Unknown | C6H6O6 | 7.852 | 0 | 7207870
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_P12 | 62533 |  | Unknown | C6H6O6 | 0 | 0 | 0
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_P13 | 105409 |  | Unknown | C6H6O6 | 0 | 0 | 0
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_P14 | 94389 |  | Unknown | C6H6O6 | 0 | 0 | 0
Cis-aconitate neg | 12 | Target Compound | KD_050321_050321_P15 | 23834 |  | Unknown | C6H6O6 | 0 | 0 | 0

## Deoxyadenosine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_AAmix | 123596 |  | Unknown | C10H13N5O3 | 3.271 | 0 | 37786
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D1 | 5718165456 |  | Cal Std | C10H13N5O3 | 9.431 | 100 | 606322793 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff -63.164% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D2 | 4558718161 |  | Cal Std | C10H13N5O3 | 7.07 | 50 | 644815922 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff -51.828% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D3 | 3735089481 |  | Cal Std | C10H13N5O3 | 4.967 | 25 | 751995116 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff -37.621% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D4 | 3181601047 |  | Cal Std | C10H13N5O3 | 3.813 | 12.5 | 834360232 | Rsquared 0.708 < R2Threshold 0.99
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D5 | 2403453363 |  | Cal Std | C10H13N5O3 | 2.963 | 6.25 | 811238901 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff 39.983% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D6 | 1771127720 |  | Cal Std | C10H13N5O3 | 2.139 | 3.125 | 828207672 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff 97.667% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D7 | 1050033610 |  | Cal Std | C10H13N5O3 | 1.524 | 1.563 | 688839294 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff 177.369% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D8 | 736126249 |  | Cal Std | C10H13N5O3 | 0.906 | 0.781 | 812590974 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff 224.998% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D9 | 357092626 |  | Cal Std | C10H13N5O3 | 0.449 | 0.391 | 795350747 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff 218.341% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D10 | 186605332 |  | Cal Std | C10H13N5O3 | 0.22 | 0.195 | 848220940 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff 211.153% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D11 | 84473124 |  | Cal Std | C10H13N5O3 | 0.111 | 0.098 | 758405819 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff 212.7% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_D12 | 39244531 |  | Cal Std | C10H13N5O3 | 0.048 | 0.049 | 813826392 | Rsquared 0.708 < R2Threshold 0.99. Calibration Amount Diff 170.381% (absolute) > max calibration amount diff 20%
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_P1 | 0 | pig1 | Unknown | C10H13N5O3 | 0 | 0 | 718462864
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_P2 | 0 | pig2 | Unknown | C10H13N5O3 | 0 | 0 | 727483012
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_P3 | 0 | pig3 | Unknown | C10H13N5O3 | 0 | 0 | 778958416
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_P12 | 75774 |  | Unknown | C10H13N5O3 | 2.663 | 0 | 28459
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C10H13N5O3 | 0 | 0 | 39719
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_P14 | 85463470 |  | Unknown | C10H13N5O3 | 1792.346 | 0 | 47682
Deoxyadenosine pos | 4.82 | Target Compound | KD_050321_050321_P15 | 30389811389 |  | Unknown | C10H13N5O3 | 0 | 0 | 0

## Deoxyguanosine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_AAmix | 192222 |  | Unknown | C10H13N5O4 | 0 | 0 | 0
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D1 | 14333698598 |  | Cal Std | C10H13N5O4 | 19.846 | 100 | 722239817
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D2 | 9421419829 |  | Cal Std | C10H13N5O4 | 9.785 | 50 | 962799228
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D3 | 6665311695 |  | Cal Std | C10H13N5O4 | 5.461 | 25 | 1220620877
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D4 | 4092383613 |  | Cal Std | C10H13N5O4 | 2.768 | 12.5 | 1478502569
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D5 | 2439485543 |  | Cal Std | C10H13N5O4 | 1.485 | 6.25 | 1642433405
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D6 | 1416268375 |  | Cal Std | C10H13N5O4 | 0.796 | 3.125 | 1780135959
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D7 | 786566968 |  | Cal Std | C10H13N5O4 | 0.412 | 1.563 | 1910559014
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D8 | 421267803 |  | Cal Std | C10H13N5O4 | 0.242 | 0.781 | 1742894159 | Calibration Amount Diff 37.644% (absolute) > max calibration amount diff 20%
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D9 | 217292818 |  | Cal Std | C10H13N5O4 | 0.115 | 0.391 | 1894708620 | Calibration Amount Diff 30.352% (absolute) > max calibration amount diff 20%
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D10 | 103530764 |  | Cal Std | C10H13N5O4 | 0.055 | 0.195 | 1889164593 | Calibration Amount Diff 24.854% (absolute) > max calibration amount diff 20%
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D11 | 53036555 |  | Cal Std | C10H13N5O4 | 0.027 | 0.098 | 1951705095 | Calibration Amount Diff 23.169% (absolute) > max calibration amount diff 20%
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_D12 | 19636163 |  | Cal Std | C10H13N5O4 | 0.012 | 0.049 | 1697092170
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_P1 | 1618202 | pig1 | Unknown | C10H13N5O4 | 0.001 | 0 | 1782302890
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_P2 | 1381965 | pig2 | Unknown | C10H13N5O4 | 0.001 | 0 | 1768756858
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_P3 | 1163982 | pig3 | Unknown | C10H13N5O4 | 0.001 | 0 | 1932080273
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_P12 | 48073 |  | Unknown | C10H13N5O4 | 12.811 | 0 | 3753
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_P13 | 5717 |  | Unknown | C10H13N5O4 | 0 | 0 | 0
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_P14 | 21563245875 |  | Unknown | C10H13N5O4 | 0 | 0 | 0
Deoxyguanosine pos | 5.5 | Target Compound | KD_050321_050321_P15 | 5093441 |  | Unknown | C10H13N5O4 | 18.678 | 0 | 272691

## Deoxythymidine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_AAmix | 223702 |  | Unknown | C10H14N2O5 | 0 | 0 | 0
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D1 | 1923847273 |  | Cal Std | C10H14N2O5 | 2851.985 | 100 | 674564 | Rsquared 0.978 < R2Threshold 0.99
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D2 | 900340619 |  | Cal Std | C10H14N2O5 | 1172.047 | 50 | 768178 | Rsquared 0.978 < R2Threshold 0.99
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D3 | 484198919 |  | Cal Std | C10H14N2O5 | 857.911 | 25 | 564393 | Rsquared 0.978 < R2Threshold 0.99. Calibration Amount Diff 22.571% (absolute) > max calibration amount diff 20%
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D4 | 236423448 |  | Cal Std | C10H14N2O5 | 441.769 | 12.5 | 535174 | Rsquared 0.978 < R2Threshold 0.99. Calibration Amount Diff 26.099% (absolute) > max calibration amount diff 20%
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D5 | 140107517 |  | Cal Std | C10H14N2O5 | 120.448 | 6.25 | 1163224 | Rsquared 0.978 < R2Threshold 0.99. Calibration Amount Diff -31.294% (absolute) > max calibration amount diff 20%
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D6 | 77189806 |  | Cal Std | C10H14N2O5 | 77.608 | 3.125 | 994608 | Rsquared 0.978 < R2Threshold 0.99
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D7 | 35789501 |  | Cal Std | C10H14N2O5 | 33.792 | 1.563 | 1059097 | Rsquared 0.978 < R2Threshold 0.99. Calibration Amount Diff -22.937% (absolute) > max calibration amount diff 20%
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D8 | 19603017 |  | Cal Std | C10H14N2O5 | 18.065 | 0.781 | 1085132 | Rsquared 0.978 < R2Threshold 0.99
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D9 | 5621351 |  | Cal Std | C10H14N2O5 | 6.437 | 0.391 | 873342 | Rsquared 0.978 < R2Threshold 0.99. Calibration Amount Diff -41.328% (absolute) > max calibration amount diff 20%
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D10 | 2485157 |  | Cal Std | C10H14N2O5 | 2.551 | 0.195 | 974347 | Rsquared 0.978 < R2Threshold 0.99. Calibration Amount Diff -53.382% (absolute) > max calibration amount diff 20%
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D11 | 1025987 |  | Cal Std | C10H14N2O5 | 0.858 | 0.098 | 1196296 | Rsquared 0.978 < R2Threshold 0.99. Calibration Amount Diff -68.809% (absolute) > max calibration amount diff 20%
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_D12 | 317103 |  | Cal Std | C10H14N2O5 | 0.293 | 0.049 | 1083792 | Rsquared 0.978 < R2Threshold 0.99. Calibration Amount Diff -78.718% (absolute) > max calibration amount diff 20%
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_P1 | 8191369 | pig1 | Unknown | C10H14N2O5 | 7.048 | 0 | 1162248
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_P2 | 3806737 | pig2 | Unknown | C10H14N2O5 | 4.636 | 0 | 821184
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_P3 | 2525446 | pig3 | Unknown | C10H14N2O5 | 3.807 | 0 | 663359
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_P12 | 119151 |  | Unknown | C10H14N2O5 | 0 | 0 | 0
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_P13 | 443128 |  | Unknown | C10H14N2O5 | 0 | 0 | 0
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_P14 | 2005535153 |  | Unknown | C10H14N2O5 | 0 | 0 | 0
Deoxythymidine neg | 4.03 | Target Compound | KD_050321_050321_P15 | 182120 |  | Unknown | C10H14N2O5 | 0 | 0 | 0

## Deoxyuridine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_AAmix | 14430 |  | Unknown | C9H12N2O5 | 0 | 0 | 0
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D1 | 1575557445 |  | Cal Std | C9H12N2O5 | 1588.352 | 100 | 991945
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D2 | 787465506 |  | Cal Std | C9H12N2O5 | 823.442 | 50 | 956309
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D3 | 509802123 |  | Cal Std | C9H12N2O5 | 474.783 | 25 | 1073757
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D4 | 269703420 |  | Cal Std | C9H12N2O5 | 193.213 | 12.5 | 1395888
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D5 | 135222702 |  | Cal Std | C9H12N2O5 | 112.397 | 6.25 | 1203084
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D6 | 78861100 |  | Cal Std | C9H12N2O5 | 58.635 | 3.125 | 1344957
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D7 | 35820636 |  | Cal Std | C9H12N2O5 | 24.859 | 1.563 | 1440931
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D8 | 22056074 |  | Cal Std | C9H12N2O5 | 13.351 | 0.781 | 1652054
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D9 | 9383679 |  | Cal Std | C9H12N2O5 | 8.277 | 0.391 | 1133689
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D10 | 1623120 |  | Cal Std | C9H12N2O5 | 1.484 | 0.195 | 1093492 | Calibration Amount Diff -57.14% (absolute) > max calibration amount diff 20%
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D11 | 751017 |  | Cal Std | C9H12N2O5 | 0.53 | 0.098 | 1417340 | Calibration Amount Diff -69.557% (absolute) > max calibration amount diff 20%
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_D12 | 371915 |  | Cal Std | C9H12N2O5 | 0.262 | 0.049 | 1420262 | Calibration Amount Diff -69.911% (absolute) > max calibration amount diff 20%
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_P1 | 3248951 | pig1 | Unknown | C9H12N2O5 | 2.34 | 0 | 1388219
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_P2 | 2611753 | pig2 | Unknown | C9H12N2O5 | 1.894 | 0 | 1378654
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_P3 | 1762149 | pig3 | Unknown | C9H12N2O5 | 1.368 | 0 | 1287883
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_P12 | 9225 |  | Unknown | C9H12N2O5 | 0 | 0 | 0
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_P13 | 5394 |  | Unknown | C9H12N2O5 | 0 | 0 | 0
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_P14 | 2002494619 |  | Unknown | C9H12N2O5 | 0 | 0 | 0
Deoxyuridine neg | 4.84 | Target Compound | KD_050321_050321_P15 | 0 |  | Unknown | C9H12N2O5 | 0 | 0 | 0

## Glutamate neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_AAmix | 621845914 |  | Unknown | C5H9NO4 | 0 | 0 | 0
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D1 | 571766567 |  | Cal Std | C5H9NO4 | 0.362 | 100 | 1580772420
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D2 | 270542524 |  | Cal Std | C5H9NO4 | 0.179 | 50 | 1509775170
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D3 | 165802200 |  | Cal Std | C5H9NO4 | 0.101 | 25 | 1643595884
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D4 | 83252823 |  | Cal Std | C5H9NO4 | 0.048 | 12.5 | 1720306795
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D5 | 44451154 |  | Cal Std | C5H9NO4 | 0.024 | 6.25 | 1823110292
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D6 | 21878850 |  | Cal Std | C5H9NO4 | 0.012 | 3.125 | 1810362502
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D7 | 10962062 |  | Cal Std | C5H9NO4 | 0.005 | 1.563 | 1999873231
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D8 | 5263382 |  | Cal Std | C5H9NO4 | 0.003 | 0.781 | 1894014824
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D9 | 2474279 |  | Cal Std | C5H9NO4 | 0.001 | 0.391 | 1928807377
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D10 | 1139043 |  | Cal Std | C5H9NO4 | 0.001 | 0.195 | 1843345301
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D11 | 699779 |  | Cal Std | C5H9NO4 | 0 | 0.098 | 1797927824
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_D12 | 561218 |  | Cal Std | C5H9NO4 | 0 | 0.049 | 1871831336 | Calibration Amount Diff 57.82% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_P1 | 609540446 | pig1 | Unknown | C5H9NO4 | 0.3 | 0 | 2029572357
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_P2 | 422339966 | pig2 | Unknown | C5H9NO4 | 0.223 | 0 | 1891884346
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_P3 | 508822461 | pig3 | Unknown | C5H9NO4 | 0.243 | 0 | 2095228732
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_P12 | 124338 |  | Unknown | C5H9NO4 | 28.784 | 0 | 4320
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_P13 | 55331 |  | Unknown | C5H9NO4 | 0 | 0 | 0
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_P14 | 37268 |  | Unknown | C5H9NO4 | 10.462 | 0 | 3562
Glutamate neg | 10.07 | Target Compound | KD_050321_050321_P15 | 111517 |  | Unknown | C5H9NO4 | 0 | 0 | 0

## Glutamine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_AAmix | 7001950857 |  | Unknown | C5H10N2O3 | 135658.555 | 0 | 51615
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D1 | 6314344815 |  | Cal Std | C5H10N2O3 | 62.428 | 100 | 101146775
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D2 | 3921170851 |  | Cal Std | C5H10N2O3 | 28.908 | 50 | 135644865
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D3 | 2540796651 |  | Cal Std | C5H10N2O3 | 15.967 | 25 | 159124599
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D4 | 1343084877 |  | Cal Std | C5H10N2O3 | 7.817 | 12.5 | 171812143
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D5 | 776715883 |  | Cal Std | C5H10N2O3 | 4.041 | 6.25 | 192208520
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D6 | 410215843 |  | Cal Std | C5H10N2O3 | 2.113 | 3.125 | 194179415
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D7 | 195777685 |  | Cal Std | C5H10N2O3 | 1.029 | 1.563 | 190272073
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D8 | 115173260 |  | Cal Std | C5H10N2O3 | 0.596 | 0.781 | 193119491 | Calibration Amount Diff 23.214% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D9 | 53594112 |  | Cal Std | C5H10N2O3 | 0.277 | 0.391 | 193200023
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D10 | 24888009 |  | Cal Std | C5H10N2O3 | 0.135 | 0.195 | 184664335
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D11 | 11779371 |  | Cal Std | C5H10N2O3 | 0.065 | 0.098 | 181906494
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_D12 | 5932269 |  | Cal Std | C5H10N2O3 | 0.03 | 0.049 | 195587954
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_P1 | 7828583032 | pig1 | Unknown | C5H10N2O3 | 97.017 | 0 | 80692719
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_P2 | 7552565485 | pig2 | Unknown | C5H10N2O3 | 70.184 | 0 | 107610908
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_P3 | 6459559168 | pig3 | Unknown | C5H10N2O3 | 74.195 | 0 | 87061589
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_P12 | 363411 |  | Unknown | C5H10N2O3 | 30.704 | 0 | 11836
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_P13 | 1417182 |  | Unknown | C5H10N2O3 | 454.258 | 0 | 3120
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C5H10N2O3 | 0 | 0 | 11090
Glutamine pos | 9.76 | Target Compound | KD_050321_050321_P15 | 0 |  | Unknown | C5H10N2O3 | 0 | 0 | 14396

## Glycine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Glycine pos | 10.21 | Target Compound | KD_050321_050321_AAmix | 19146949 |  | Unknown | C2H5NO2 | 0 | 0 | 0
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D1 | 16614572 |  | Cal Std | C2H5NO2 | 98.216 | 100 | 169164 | Rsquared 0.989 < R2Threshold 0.99
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D2 | 8644357 |  | Cal Std | C2H5NO2 | 19.122 | 50 | 452075 | Rsquared 0.989 < R2Threshold 0.99
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D3 | 4812341 |  | Cal Std | C2H5NO2 | 6.218 | 25 | 773995 | Rsquared 0.989 < R2Threshold 0.99
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D4 | 2306996 |  | Cal Std | C2H5NO2 | 3.522 | 12.5 | 655012 | Rsquared 0.989 < R2Threshold 0.99. Calibration Amount Diff 42.764% (absolute) > max calibration amount diff 20%
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D5 | 1320092 |  | Cal Std | C2H5NO2 | 1.415 | 6.25 | 932930 | Rsquared 0.989 < R2Threshold 0.99. Calibration Amount Diff 71.845% (absolute) > max calibration amount diff 20%
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D6 | 595633 |  | Cal Std | C2H5NO2 | 0.671 | 3.125 | 887221 | Rsquared 0.989 < R2Threshold 0.99. Calibration Amount Diff 122.201% (absolute) > max calibration amount diff 20%
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D7 | 250049 |  | Cal Std | C2H5NO2 | 0.266 | 1.563 | 939255 | Rsquared 0.989 < R2Threshold 0.99. Calibration Amount Diff 148.776% (absolute) > max calibration amount diff 20%
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D8 | 152793 |  | Cal Std | C2H5NO2 | 0.142 | 0.781 | 1077933 | Rsquared 0.989 < R2Threshold 0.99. Calibration Amount Diff 224.374% (absolute) > max calibration amount diff 20%
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D9 | 36271 |  | Cal Std | C2H5NO2 | 0.038 | 0.391 | 962960 | Rsquared 0.989 < R2Threshold 0.99. Calibration Amount Diff 134.769% (absolute) > max calibration amount diff 20%
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D10 | 13982 |  | Cal Std | C2H5NO2 | 0.014 | 0.195 | 1028527 | Rsquared 0.989 < R2Threshold 0.99. Calibration Amount Diff 93.44% (absolute) > max calibration amount diff 20%
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D11 | 0 |  | Cal Std | C2H5NO2 | 0 | 0.098 | 883621 | Rsquared 0.989 < R2Threshold 0.99
Glycine pos | 10.21 | Target Compound | KD_050321_050321_D12 | 16092 |  | Cal Std | C2H5NO2 | 0.02 | 0.049 | 801805 | Rsquared 0.989 < R2Threshold 0.99. Calibration Amount Diff 992.243% (absolute) > max calibration amount diff 20%
Glycine pos | 10.21 | Target Compound | KD_050321_050321_P1 | 128035101 | pig1 | Unknown | C2H5NO2 | 316.563 | 0 | 404454
Glycine pos | 10.21 | Target Compound | KD_050321_050321_P2 | 118899528 | pig2 | Unknown | C2H5NO2 | 185.262 | 0 | 641791
Glycine pos | 10.21 | Target Compound | KD_050321_050321_P3 | 158545298 | pig3 | Unknown | C2H5NO2 | 237.705 | 0 | 666984
Glycine pos | 10.21 | Target Compound | KD_050321_050321_P12 | 135768 |  | Unknown | C2H5NO2 | 0 | 0 | 0
Glycine pos | 10.21 | Target Compound | KD_050321_050321_P13 | 13708 |  | Unknown | C2H5NO2 | 0 | 0 | 0
Glycine pos | 10.21 | Target Compound | KD_050321_050321_P14 | 17819 |  | Unknown | C2H5NO2 | 0 | 0 | 0
Glycine pos | 10.21 | Target Compound | KD_050321_050321_P15 | 17427 |  | Unknown | C2H5NO2 | 0 | 0 | 0

## Isoleucine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_AAmix | 5534802033 |  | Unknown | C6H13NO2 | 0 | 0 | 0
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D1 | 4648657610 |  | Cal Std | C6H13NO2 | 38.551 | 100 | 120583155
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D2 | 2670299517 |  | Cal Std | C6H13NO2 | 17.942 | 50 | 148831810
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D3 | 1643354230 |  | Cal Std | C6H13NO2 | 9.321 | 25 | 176297586
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D4 | 889072814 |  | Cal Std | C6H13NO2 | 4.381 | 12.5 | 202942695
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D5 | 485458479 |  | Cal Std | C6H13NO2 | 2.232 | 6.25 | 217451895
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D6 | 262318043 |  | Cal Std | C6H13NO2 | 1.158 | 3.125 | 226519450
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D7 | 133861750 |  | Cal Std | C6H13NO2 | 0.552 | 1.563 | 242431443
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D8 | 74451015 |  | Cal Std | C6H13NO2 | 0.317 | 0.781 | 234977349
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D9 | 34090027 |  | Cal Std | C6H13NO2 | 0.145 | 0.391 | 235771538
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D10 | 14643054 |  | Cal Std | C6H13NO2 | 0.067 | 0.195 | 218599138
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D11 | 7478552 |  | Cal Std | C6H13NO2 | 0.033 | 0.098 | 229144820
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_D12 | 3262041 |  | Cal Std | C6H13NO2 | 0.014 | 0.049 | 230396666
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_P1 | 2338939739 | pig1 | Unknown | C6H13NO2 | 26.174 | 0 | 89360627
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_P2 | 2098697596 | pig2 | Unknown | C6H13NO2 | 16.96 | 0 | 123742294
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_P3 | 1966029057 | pig3 | Unknown | C6H13NO2 | 0 | 0 | 0
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_P12 | 650509 |  | Unknown | C6H13NO2 | 33.62 | 0 | 19349
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_P13 | 286555 |  | Unknown | C6H13NO2 | 0 | 0 | 0
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_P14 | 254710 |  | Unknown | C6H13NO2 | 0 | 0 | 0
Isoleucine pos | 7.15 | Target Compound | KD_050321_050321_P15 | 69509 |  | Unknown | C6H13NO2 | 0 | 0 | 0

## Leucine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Leucine pos | 6.84 | Target Compound | KD_050321_050321_AAmix | 4220867953 |  | Unknown | C6H13NO2 | 0 | 0 | 0
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D1 | 3966230745 |  | Cal Std | C6H13NO2 | 19.254 | 100 | 205991105
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D2 | 2240892878 |  | Cal Std | C6H13NO2 | 9.559 | 50 | 234429857
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D3 | 1357910023 |  | Cal Std | C6H13NO2 | 5.401 | 25 | 251401790
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D4 | 704302921 |  | Cal Std | C6H13NO2 | 2.597 | 12.5 | 271225290
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D5 | 406860385 |  | Cal Std | C6H13NO2 | 1.336 | 6.25 | 304464961
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D6 | 205591663 |  | Cal Std | C6H13NO2 | 0.666 | 3.125 | 308866174
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D7 | 108835073 |  | Cal Std | C6H13NO2 | 0.341 | 1.563 | 319091768
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D8 | 59851955 |  | Cal Std | C6H13NO2 | 0.191 | 0.781 | 314018202
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D9 | 28911225 |  | Cal Std | C6H13NO2 | 0.09 | 0.391 | 321488322
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D10 | 14947896 |  | Cal Std | C6H13NO2 | 0.047 | 0.195 | 319029968
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D11 | 6704293 |  | Cal Std | C6H13NO2 | 0.022 | 0.098 | 298596872
Leucine pos | 6.84 | Target Compound | KD_050321_050321_D12 | 2736286 |  | Cal Std | C6H13NO2 | 0.009 | 0.049 | 291789631
Leucine pos | 6.84 | Target Compound | KD_050321_050321_P1 | 4267083534 | pig1 | Unknown | C6H13NO2 | 14.768 | 0 | 288948876
Leucine pos | 6.84 | Target Compound | KD_050321_050321_P2 | 3483445985 | pig2 | Unknown | C6H13NO2 | 12.716 | 0 | 273947715
Leucine pos | 6.84 | Target Compound | KD_050321_050321_P3 | 4186424303 | pig3 | Unknown | C6H13NO2 | 11.287 | 0 | 370898386
Leucine pos | 6.84 | Target Compound | KD_050321_050321_P12 | 526628 |  | Unknown | C6H13NO2 | 17.422 | 0 | 30227
Leucine pos | 6.84 | Target Compound | KD_050321_050321_P13 | 306374 |  | Unknown | C6H13NO2 | 0 | 0 | 0
Leucine pos | 6.84 | Target Compound | KD_050321_050321_P14 | 255518 |  | Unknown | C6H13NO2 | 0 | 0 | 0
Leucine pos | 6.84 | Target Compound | KD_050321_050321_P15 | 757245 |  | Unknown | C6H13NO2 | 11.372 | 0 | 66587

## Lysine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Lysine pos | 14.07 | Target Compound | KD_050321_050321_AAmix | 1465891966 |  | Unknown | C6H14N2O2 | 0 | 0 | 0
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D1 | 1562092591 |  | Cal Std | C6H14N2O2 | 22.582 | 100 | 69173603
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D2 | 767145613 |  | Cal Std | C6H14N2O2 | 11.319 | 50 | 67777188
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D3 | 419329369 |  | Cal Std | C6H14N2O2 | 6.479 | 25 | 64725209
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D4 | 168688816 |  | Cal Std | C6H14N2O2 | 3.22 | 12.5 | 52380136
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D5 | 71277762 |  | Cal Std | C6H14N2O2 | 1.677 | 6.25 | 42497239
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D6 | 31171427 |  | Cal Std | C6H14N2O2 | 0.85 | 3.125 | 36690912
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D7 | 14736559 |  | Cal Std | C6H14N2O2 | 0.413 | 1.563 | 35643048
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D8 | 7372711 |  | Cal Std | C6H14N2O2 | 0.229 | 0.781 | 32226004
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D9 | 3255845 |  | Cal Std | C6H14N2O2 | 0.101 | 0.391 | 32243604
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D10 | 1393289 |  | Cal Std | C6H14N2O2 | 0.046 | 0.195 | 30345366
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D11 | 633646 |  | Cal Std | C6H14N2O2 | 0.022 | 0.098 | 28206679
Lysine pos | 14.07 | Target Compound | KD_050321_050321_D12 | 495167 |  | Cal Std | C6H14N2O2 | 0.016 | 0.049 | 31021430 | Calibration Amount Diff 25.956% (absolute) > max calibration amount diff 20%
Lysine pos | 14.07 | Target Compound | KD_050321_050321_P1 | 997747562 | pig1 | Unknown | C6H14N2O2 | 15.375 | 0 | 64895545
Lysine pos | 14.07 | Target Compound | KD_050321_050321_P2 | 556509238 | pig2 | Unknown | C6H14N2O2 | 8.958 | 0 | 62124925
Lysine pos | 14.07 | Target Compound | KD_050321_050321_P3 | 888030015 | pig3 | Unknown | C6H14N2O2 | 13.517 | 0 | 65695907
Lysine pos | 14.07 | Target Compound | KD_050321_050321_P12 | 315439 |  | Unknown | C6H14N2O2 | 0 | 0 | 0
Lysine pos | 14.07 | Target Compound | KD_050321_050321_P13 | 222108 |  | Unknown | C6H14N2O2 | 0 | 0 | 0
Lysine pos | 14.07 | Target Compound | KD_050321_050321_P14 | 273987 |  | Unknown | C6H14N2O2 | 0 | 0 | 0
Lysine pos | 14.07 | Target Compound | KD_050321_050321_P15 | 490544 |  | Unknown | C6H14N2O2 | 0 | 0 | 0

## Malate neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Malate neg | 10.92 | Target Compound | KD_050321_050321_AAmix | 47852424 |  | Unknown | C4H6O5 | 1957.722 | 0 | 24443
Malate neg | 10.92 | Target Compound | KD_050321_050321_D1 | 3407518882 |  | Cal Std | C4H6O5 | 64.732 | 100 | 52640004
Malate neg | 10.92 | Target Compound | KD_050321_050321_D2 | 2155580436 |  | Cal Std | C4H6O5 | 30.689 | 50 | 70239189
Malate neg | 10.92 | Target Compound | KD_050321_050321_D3 | 1354712068 |  | Cal Std | C4H6O5 | 17.194 | 25 | 78790153
Malate neg | 10.92 | Target Compound | KD_050321_050321_D4 | 837139504 |  | Cal Std | C4H6O5 | 8.303 | 12.5 | 100826164
Malate neg | 10.92 | Target Compound | KD_050321_050321_D5 | 421464387 |  | Cal Std | C4H6O5 | 4.511 | 6.25 | 93431543
Malate neg | 10.92 | Target Compound | KD_050321_050321_D6 | 244073461 |  | Cal Std | C4H6O5 | 2.301 | 3.125 | 106079771
Malate neg | 10.92 | Target Compound | KD_050321_050321_D7 | 122361405 |  | Cal Std | C4H6O5 | 1.165 | 1.563 | 105001918
Malate neg | 10.92 | Target Compound | KD_050321_050321_D8 | 72185051 |  | Cal Std | C4H6O5 | 0.685 | 0.781 | 105408238 | Calibration Amount Diff 28.852% (absolute) > max calibration amount diff 20%
Malate neg | 10.92 | Target Compound | KD_050321_050321_D9 | 33890752 |  | Cal Std | C4H6O5 | 0.333 | 0.391 | 101647360 | Calibration Amount Diff 25.266% (absolute) > max calibration amount diff 20%
Malate neg | 10.92 | Target Compound | KD_050321_050321_D10 | 22486734 |  | Cal Std | C4H6O5 | 0.212 | 0.195 | 105998848 | Calibration Amount Diff 59.796% (absolute) > max calibration amount diff 20%
Malate neg | 10.92 | Target Compound | KD_050321_050321_D11 | 12577944 |  | Cal Std | C4H6O5 | 0.126 | 0.098 | 100069965 | Calibration Amount Diff 88.373% (absolute) > max calibration amount diff 20%
Malate neg | 10.92 | Target Compound | KD_050321_050321_D12 | 6862911 |  | Cal Std | C4H6O5 | 0.071 | 0.049 | 96188876 | Calibration Amount Diff 113.848% (absolute) > max calibration amount diff 20%
Malate neg | 10.92 | Target Compound | KD_050321_050321_P1 | 298225746 | pig1 | Unknown | C4H6O5 | 2.031 | 0 | 146869853
Malate neg | 10.92 | Target Compound | KD_050321_050321_P2 | 261024446 | pig2 | Unknown | C4H6O5 | 1.946 | 0 | 134167788
Malate neg | 10.92 | Target Compound | KD_050321_050321_P3 | 338917560 | pig3 | Unknown | C4H6O5 | 2.58 | 0 | 131369589
Malate neg | 10.92 | Target Compound | KD_050321_050321_P12 | 558119 |  | Unknown | C4H6O5 | 0 | 0 | 0
Malate neg | 10.92 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C4H6O5 | 0 | 0 | 0
Malate neg | 10.92 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C4H6O5 | 0 | 0 | 0
Malate neg | 10.92 | Target Compound | KD_050321_050321_P15 | 3446075040 |  | Unknown | C4H6O5 | 0 | 0 | 0

## Methionine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Methionine pos | 7.22 | Target Compound | KD_050321_050321_AAmix | 7613375136 |  | Unknown | C5H11NO2S | 0 | 0 | 0
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D1 | 6219053854 |  | Cal Std | C5H11NO2S | 126.991 | 100 | 48972445
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D2 | 3549996955 |  | Cal Std | C5H11NO2S | 59.433 | 50 | 59730639
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D3 | 2228593718 |  | Cal Std | C5H11NO2S | 29.029 | 25 | 76770288
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D4 | 1189017324 |  | Cal Std | C5H11NO2S | 13.389 | 12.5 | 88805621
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D5 | 618499902 |  | Cal Std | C5H11NO2S | 6.948 | 6.25 | 89021688
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D6 | 323234942 |  | Cal Std | C5H11NO2S | 3.453 | 3.125 | 93623001
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D7 | 171117022 |  | Cal Std | C5H11NO2S | 1.692 | 1.563 | 101154527
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D8 | 91742647 |  | Cal Std | C5H11NO2S | 0.958 | 0.781 | 95736810
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D9 | 43260340 |  | Cal Std | C5H11NO2S | 0.448 | 0.391 | 96494583
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D10 | 18962660 |  | Cal Std | C5H11NO2S | 0.217 | 0.195 | 87193535
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D11 | 8966739 |  | Cal Std | C5H11NO2S | 0.099 | 0.098 | 90190156
Methionine pos | 7.22 | Target Compound | KD_050321_050321_D12 | 3682841 |  | Cal Std | C5H11NO2S | 0.041 | 0.049 | 90002882 | Calibration Amount Diff -23.999% (absolute) > max calibration amount diff 20%
Methionine pos | 7.22 | Target Compound | KD_050321_050321_P1 | 1348427304 | pig1 | Unknown | C5H11NO2S | 34.521 | 0 | 39060619
Methionine pos | 7.22 | Target Compound | KD_050321_050321_P2 | 629119971 | pig2 | Unknown | C5H11NO2S | 13.451 | 0 | 46770861
Methionine pos | 7.22 | Target Compound | KD_050321_050321_P3 | 1001602352 | pig3 | Unknown | C5H11NO2S | 28.48 | 0 | 35169064
Methionine pos | 7.22 | Target Compound | KD_050321_050321_P12 | 24171 |  | Unknown | C5H11NO2S | 0 | 0 | 0
Methionine pos | 7.22 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C5H11NO2S | 0 | 0 | 0
Methionine pos | 7.22 | Target Compound | KD_050321_050321_P14 | 5521 |  | Unknown | C5H11NO2S | 0 | 0 | 0
Methionine pos | 7.22 | Target Compound | KD_050321_050321_P15 | 0 |  | Unknown | C5H11NO2S | 0 | 0 | 0

## Phenylalanine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_AAmix | 8352374292 |  | Unknown | C9H11NO2 | 0 | 0 | 0
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D1 | 8215263877 |  | Cal Std | C9H11NO2 | 37.49 | 100 | 219133720
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D2 | 4297177287 |  | Cal Std | C9H11NO2 | 18.409 | 50 | 233426677
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D3 | 2542035449 |  | Cal Std | C9H11NO2 | 10.009 | 25 | 253971928
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D4 | 1239847113 |  | Cal Std | C9H11NO2 | 4.851 | 12.5 | 255609140
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D5 | 630633362 |  | Cal Std | C9H11NO2 | 2.53 | 6.25 | 249218111
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D6 | 327422750 |  | Cal Std | C9H11NO2 | 1.328 | 3.125 | 246555453
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D7 | 160781838 |  | Cal Std | C9H11NO2 | 0.657 | 1.563 | 244558966
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D8 | 85217156 |  | Cal Std | C9H11NO2 | 0.385 | 0.781 | 221462271 | Calibration Amount Diff 24.137% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D9 | 39868401 |  | Cal Std | C9H11NO2 | 0.175 | 0.391 | 227263614
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D10 | 19901912 |  | Cal Std | C9H11NO2 | 0.088 | 0.195 | 227041163
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D11 | 9440203 |  | Cal Std | C9H11NO2 | 0.044 | 0.098 | 216483493
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_D12 | 4240238 |  | Cal Std | C9H11NO2 | 0.019 | 0.049 | 227375756
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_P1 | 2341150833 | pig1 | Unknown | C9H11NO2 | 9.897 | 0 | 236550061
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_P2 | 2094697083 | pig2 | Unknown | C9H11NO2 | 8.645 | 0 | 242304933
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_P3 | 2696092310 | pig3 | Unknown | C9H11NO2 | 10.898 | 0 | 247402458
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_P12 | 438837 |  | Unknown | C9H11NO2 | 59.163 | 0 | 7417
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_P13 | 75256 |  | Unknown | C9H11NO2 | 0 | 0 | 0
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_P14 | 684366 |  | Unknown | C9H11NO2 | 78.019 | 0 | 8772
Phenylalanine pos | 6.24 | Target Compound | KD_050321_050321_P15 | 286091 |  | Unknown | C9H11NO2 | 93.042 | 0 | 3075

## Proline pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Proline pos | 8.25 | Target Compound | KD_050321_050321_AAmix | 14712974097 |  | Unknown | C5H9NO2 | 0 | 0 | 0
Proline pos | 8.25 | Target Compound | KD_050321_050321_D1 | 14635316244 |  | Cal Std | C5H9NO2 | 51.292 | 100 | 285333105
Proline pos | 8.25 | Target Compound | KD_050321_050321_D2 | 9191851749 |  | Cal Std | C5H9NO2 | 24.145 | 50 | 380701420
Proline pos | 8.25 | Target Compound | KD_050321_050321_D3 | 5920939017 |  | Cal Std | C5H9NO2 | 13.284 | 25 | 445704990
Proline pos | 8.25 | Target Compound | KD_050321_050321_D4 | 3412303356 |  | Cal Std | C5H9NO2 | 6.522 | 12.5 | 523207185
Proline pos | 8.25 | Target Compound | KD_050321_050321_D5 | 1923006333 |  | Cal Std | C5H9NO2 | 3.583 | 6.25 | 536711945
Proline pos | 8.25 | Target Compound | KD_050321_050321_D6 | 1028002600 |  | Cal Std | C5H9NO2 | 1.897 | 3.125 | 541827525
Proline pos | 8.25 | Target Compound | KD_050321_050321_D7 | 535931676 |  | Cal Std | C5H9NO2 | 0.963 | 1.563 | 556394308
Proline pos | 8.25 | Target Compound | KD_050321_050321_D8 | 302696588 |  | Cal Std | C5H9NO2 | 0.578 | 0.781 | 523678124 | Calibration Amount Diff 38.777% (absolute) > max calibration amount diff 20%
Proline pos | 8.25 | Target Compound | KD_050321_050321_D9 | 146601393 |  | Cal Std | C5H9NO2 | 0.279 | 0.391 | 524852429 | Calibration Amount Diff 33.911% (absolute) > max calibration amount diff 20%
Proline pos | 8.25 | Target Compound | KD_050321_050321_D10 | 76624701 |  | Cal Std | C5H9NO2 | 0.142 | 0.195 | 540777176 | Calibration Amount Diff 36.191% (absolute) > max calibration amount diff 20%
Proline pos | 8.25 | Target Compound | KD_050321_050321_D11 | 38566596 |  | Cal Std | C5H9NO2 | 0.073 | 0.098 | 528392093 | Calibration Amount Diff 39.582% (absolute) > max calibration amount diff 20%
Proline pos | 8.25 | Target Compound | KD_050321_050321_D12 | 19878623 |  | Cal Std | C5H9NO2 | 0.038 | 0.049 | 522469626 | Calibration Amount Diff 45.517% (absolute) > max calibration amount diff 20%
Proline pos | 8.25 | Target Compound | KD_050321_050321_P1 | 11231821779 | pig1 | Unknown | C5H9NO2 | 37.096 | 0 | 302779315
Proline pos | 8.25 | Target Compound | KD_050321_050321_P2 | 9110360801 | pig2 | Unknown | C5H9NO2 | 28.261 | 0 | 322361429
Proline pos | 8.25 | Target Compound | KD_050321_050321_P3 | 12516147444 | pig3 | Unknown | C5H9NO2 | 44.325 | 0 | 282372173
Proline pos | 8.25 | Target Compound | KD_050321_050321_P12 | 6900888 |  | Unknown | C5H9NO2 | 0 | 0 | 0
Proline pos | 8.25 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C5H9NO2 | 0 | 0 | 0
Proline pos | 8.25 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C5H9NO2 | 0 | 0 | 21915
Proline pos | 8.25 | Target Compound | KD_050321_050321_P15 | 1538527 |  | Unknown | C5H9NO2 | 0 | 0 | 0

## Serine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Serine neg | 10.14 | Target Compound | KD_050321_050321_AAmix | 146870820 |  | Unknown | C3H7NO3 | 0 | 0 | 0
Serine neg | 10.14 | Target Compound | KD_050321_050321_D1 | 104230742 |  | Cal Std | C3H7NO3 | 25.606 | 100 | 4070557
Serine neg | 10.14 | Target Compound | KD_050321_050321_D2 | 49549623 |  | Cal Std | C3H7NO3 | 10.372 | 50 | 4777025
Serine neg | 10.14 | Target Compound | KD_050321_050321_D3 | 25954902 |  | Cal Std | C3H7NO3 | 5.089 | 25 | 5100066
Serine neg | 10.14 | Target Compound | KD_050321_050321_D4 | 11855553 |  | Cal Std | C3H7NO3 | 2.021 | 12.5 | 5865531
Serine neg | 10.14 | Target Compound | KD_050321_050321_D5 | 5629253 |  | Cal Std | C3H7NO3 | 0.883 | 6.25 | 6374677
Serine neg | 10.14 | Target Compound | KD_050321_050321_D6 | 2597632 |  | Cal Std | C3H7NO3 | 0.371 | 3.125 | 7007625 | Calibration Amount Diff -24.437% (absolute) > max calibration amount diff 20%
Serine neg | 10.14 | Target Compound | KD_050321_050321_D7 | 1103248 |  | Cal Std | C3H7NO3 | 0.152 | 1.563 | 7253265 | Calibration Amount Diff -37.44% (absolute) > max calibration amount diff 20%
Serine neg | 10.14 | Target Compound | KD_050321_050321_D8 | 612075 |  | Cal Std | C3H7NO3 | 0.08 | 0.781 | 7612599 | Calibration Amount Diff -33.616% (absolute) > max calibration amount diff 20%
Serine neg | 10.14 | Target Compound | KD_050321_050321_D9 | 243146 |  | Cal Std | C3H7NO3 | 0.034 | 0.391 | 7181405 | Calibration Amount Diff -44.051% (absolute) > max calibration amount diff 20%
Serine neg | 10.14 | Target Compound | KD_050321_050321_D10 | 155896 |  | Cal Std | C3H7NO3 | 0.02 | 0.195 | 7826111 | Calibration Amount Diff -33.958% (absolute) > max calibration amount diff 20%
Serine neg | 10.14 | Target Compound | KD_050321_050321_D11 | 82757 |  | Cal Std | C3H7NO3 | 0.011 | 0.098 | 7228416 | Calibration Amount Diff -24.445% (absolute) > max calibration amount diff 20%
Serine neg | 10.14 | Target Compound | KD_050321_050321_D12 | 46274 |  | Cal Std | C3H7NO3 | 0.006 | 0.049 | 7629338
Serine neg | 10.14 | Target Compound | KD_050321_050321_P1 | 58366919 | pig1 | Unknown | C3H7NO3 | 9.521 | 0 | 6130314
Serine neg | 10.14 | Target Compound | KD_050321_050321_P2 | 53181328 | pig2 | Unknown | C3H7NO3 | 8.633 | 0 | 6160188
Serine neg | 10.14 | Target Compound | KD_050321_050321_P3 | 65027207 | pig3 | Unknown | C3H7NO3 | 10.603 | 0 | 6132979
Serine neg | 10.14 | Target Compound | KD_050321_050321_P12 | 918901 |  | Unknown | C3H7NO3 | 0 | 0 | 0
Serine neg | 10.14 | Target Compound | KD_050321_050321_P13 | 56252 |  | Unknown | C3H7NO3 | 0 | 0 | 0
Serine neg | 10.14 | Target Compound | KD_050321_050321_P14 | 113993 |  | Unknown | C3H7NO3 | 0 | 0 | 0
Serine neg | 10.14 | Target Compound | KD_050321_050321_P15 | 120398 |  | Unknown | C3H7NO3 | 0 | 0 | 0

## Succinate neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Succinate neg | 10.44 | Target Compound | KD_050321_050321_AAmix | 2855509 |  | Unknown | C4H6O4 | 0 | 0 | 0
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D1 | 609797133 |  | Cal Std | C4H6O4 | 11.38 | 100 | 53586278 | Rsquared 0.962 < R2Threshold 0.99
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D2 | 313611631 |  | Cal Std | C4H6O4 | 6.031 | 50 | 51999486 | Rsquared 0.962 < R2Threshold 0.99
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D3 | 178333476 |  | Cal Std | C4H6O4 | 3.483 | 25 | 51195302 | Rsquared 0.962 < R2Threshold 0.99
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D4 | 88176085 |  | Cal Std | C4H6O4 | 1.668 | 12.5 | 52877434 | Rsquared 0.962 < R2Threshold 0.99
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D5 | 42981544 |  | Cal Std | C4H6O4 | 0.87 | 6.25 | 49426889 | Rsquared 0.962 < R2Threshold 0.99
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D6 | 23733907 |  | Cal Std | C4H6O4 | 0.456 | 3.125 | 52089475 | Rsquared 0.962 < R2Threshold 0.99
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D7 | 12751102 |  | Cal Std | C4H6O4 | 0.247 | 1.563 | 51679600 | Rsquared 0.962 < R2Threshold 0.99
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D8 | 7412947 |  | Cal Std | C4H6O4 | 0.15 | 0.781 | 49557104 | Rsquared 0.962 < R2Threshold 0.99. Calibration Amount Diff 32.167% (absolute) > max calibration amount diff 20%
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D9 | 5078676 |  | Cal Std | C4H6O4 | 0.103 | 0.391 | 49475542 | Rsquared 0.962 < R2Threshold 0.99. Calibration Amount Diff 81.029% (absolute) > max calibration amount diff 20%
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D10 | 5673391 |  | Cal Std | C4H6O4 | 0.111 | 0.195 | 51218536 | Rsquared 0.962 < R2Threshold 0.99. Calibration Amount Diff 291.744% (absolute) > max calibration amount diff 20%
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D11 | 1593494 |  | Cal Std | C4H6O4 | 0.032 | 0.098 | 49090261 | Rsquared 0.962 < R2Threshold 0.99. Calibration Amount Diff 128.144% (absolute) > max calibration amount diff 20%
Succinate neg | 10.44 | Target Compound | KD_050321_050321_D12 | 3060205 |  | Cal Std | C4H6O4 | 0.06 | 0.049 | 50904136 | Rsquared 0.962 < R2Threshold 0.99. Calibration Amount Diff 745.421% (absolute) > max calibration amount diff 20%
Succinate neg | 10.44 | Target Compound | KD_050321_050321_P1 | 34179353 | pig1 | Unknown | C4H6O4 | 0.699 | 0 | 48895012
Succinate neg | 10.44 | Target Compound | KD_050321_050321_P2 | 12646668 | pig2 | Unknown | C4H6O4 | 0.261 | 0 | 48510557
Succinate neg | 10.44 | Target Compound | KD_050321_050321_P3 | 17585014 | pig3 | Unknown | C4H6O4 | 0.35 | 0 | 50203437
Succinate neg | 10.44 | Target Compound | KD_050321_050321_P12 | 1109087 |  | Unknown | C4H6O4 | 87.985 | 0 | 12605
Succinate neg | 10.44 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C4H6O4 | 0 | 0 | 0
Succinate neg | 10.44 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C4H6O4 | 0 | 0 | 0
Succinate neg | 10.44 | Target Compound | KD_050321_050321_P15 | 678470311 |  | Unknown | C4H6O4 | 0 | 0 | 0

## Threonine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Threonine neg | 9.23 | Target Compound | KD_050321_050321_AAmix | 341380962 |  | Unknown | C4H9NO3 | 0 | 0 | 0
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D1 | 280773701 |  | Cal Std | C4H9NO3 | 33.859 | 100 | 8292463
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D2 | 176502444 |  | Cal Std | C4H9NO3 | 14.854 | 50 | 11882497
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D3 | 114634557 |  | Cal Std | C4H9NO3 | 7.829 | 25 | 14642292
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D4 | 61268543 |  | Cal Std | C4H9NO3 | 3.711 | 12.5 | 16509652
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D5 | 34511316 |  | Cal Std | C4H9NO3 | 1.845 | 6.25 | 18707948
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D6 | 19477658 |  | Cal Std | C4H9NO3 | 0.926 | 3.125 | 21027584
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D7 | 9993708 |  | Cal Std | C4H9NO3 | 0.461 | 1.563 | 21659637
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D8 | 5668018 |  | Cal Std | C4H9NO3 | 0.264 | 0.781 | 21458615
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D9 | 2685043 |  | Cal Std | C4H9NO3 | 0.117 | 0.391 | 22914535
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D10 | 1360504 |  | Cal Std | C4H9NO3 | 0.056 | 0.195 | 24359034
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D11 | 555598 |  | Cal Std | C4H9NO3 | 0.023 | 0.098 | 24461878
Threonine neg | 9.23 | Target Compound | KD_050321_050321_D12 | 390919 |  | Cal Std | C4H9NO3 | 0.016 | 0.049 | 24277914
Threonine neg | 9.23 | Target Compound | KD_050321_050321_P1 | 202326251 | pig1 | Unknown | C4H9NO3 | 17.091 | 0 | 11837850
Threonine neg | 9.23 | Target Compound | KD_050321_050321_P2 | 186213896 | pig2 | Unknown | C4H9NO3 | 14.074 | 0 | 13230994
Threonine neg | 9.23 | Target Compound | KD_050321_050321_P3 | 235623095 | pig3 | Unknown | C4H9NO3 | 19.18 | 0 | 12284721
Threonine neg | 9.23 | Target Compound | KD_050321_050321_P12 | 265389 |  | Unknown | C4H9NO3 | 0 | 0 | 0
Threonine neg | 9.23 | Target Compound | KD_050321_050321_P13 | 0 |  | Unknown | C4H9NO3 | 0 | 0 | 0
Threonine neg | 9.23 | Target Compound | KD_050321_050321_P14 | 312782 |  | Unknown | C4H9NO3 | 0 | 0 | 0
Threonine neg | 9.23 | Target Compound | KD_050321_050321_P15 | 0 |  | Unknown | C4H9NO3 | 0 | 0 | 0

## Tryptophan pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_AAmix | 4526643214 |  | Unknown | C11H12N2O2 | 0 | 0 | 0
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D1 | 3332424768 |  | Cal Std | C11H12N2O2 | 298.84 | 100 | 11151183
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D2 | 1667406847 |  | Cal Std | C11H12N2O2 | 127.342 | 50 | 13093972
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D3 | 1012712068 |  | Cal Std | C11H12N2O2 | 61.458 | 25 | 16478163
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D4 | 479724254 |  | Cal Std | C11H12N2O2 | 26.668 | 12.5 | 17988691
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D5 | 235558962 |  | Cal Std | C11H12N2O2 | 13.309 | 6.25 | 17699186
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D6 | 120600962 |  | Cal Std | C11H12N2O2 | 6.307 | 3.125 | 19121354
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D7 | 58435326 |  | Cal Std | C11H12N2O2 | 2.993 | 1.563 | 19523206
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D8 | 35281535 |  | Cal Std | C11H12N2O2 | 1.609 | 0.781 | 21927761
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D9 | 15500379 |  | Cal Std | C11H12N2O2 | 0.703 | 0.391 | 22037927
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D10 | 6302871 |  | Cal Std | C11H12N2O2 | 0.314 | 0.195 | 20070440 | Calibration Amount Diff -23.737% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D11 | 2589670 |  | Cal Std | C11H12N2O2 | 0.123 | 0.098 | 21057869 | Calibration Amount Diff -40.552% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_D12 | 966284 |  | Cal Std | C11H12N2O2 | 0.046 | 0.049 | 20949958 | Calibration Amount Diff -55.401% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_P1 | 938952515 | pig1 | Unknown | C11H12N2O2 | 61.758 | 0 | 15203691
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_P2 | 836244655 | pig2 | Unknown | C11H12N2O2 | 56.066 | 0 | 14915295
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_P3 | 1129921879 | pig3 | Unknown | C11H12N2O2 | 84.773 | 0 | 13328740
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_P12 | 138236 |  | Unknown | C11H12N2O2 | 0 | 0 | 0
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_P13 | 13879 |  | Unknown | C11H12N2O2 | 0 | 0 | 0
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_P14 | 9801 |  | Unknown | C11H12N2O2 | 0 | 0 | 0
Tryptophan pos | 7.36 | Target Compound | KD_050321_050321_P15 | 9652 |  | Unknown | C11H12N2O2 | 1.44 | 0 | 6701

## Tyrosine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_AAmix | 2351328271 |  | Unknown | C9H11NO3 | 0 | 0 | 0
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D1 | 2202968342 |  | Cal Std | C9H11NO3 | 57.847 | 100 | 38082440
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D2 | 1136013741 |  | Cal Std | C9H11NO3 | 26.232 | 50 | 43307161
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D3 | 765247606 |  | Cal Std | C9H11NO3 | 13.277 | 25 | 57634937
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D4 | 420331133 |  | Cal Std | C9H11NO3 | 5.791 | 12.5 | 72585502
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D5 | 209013678 |  | Cal Std | C9H11NO3 | 2.894 | 6.25 | 72222969
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D6 | 109680502 |  | Cal Std | C9H11NO3 | 1.445 | 3.125 | 75909986
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D7 | 54328879 |  | Cal Std | C9H11NO3 | 0.68 | 1.563 | 79862215
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D8 | 31369981 |  | Cal Std | C9H11NO3 | 0.37 | 0.781 | 84896078
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D9 | 12432610 |  | Cal Std | C9H11NO3 | 0.158 | 0.391 | 78710765
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D10 | 5753914 |  | Cal Std | C9H11NO3 | 0.065 | 0.195 | 88179304 | Calibration Amount Diff -29.326% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D11 | 2396922 |  | Cal Std | C9H11NO3 | 0.029 | 0.098 | 82272337 | Calibration Amount Diff -37.202% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_D12 | 1166197 |  | Cal Std | C9H11NO3 | 0.014 | 0.049 | 82888022 | Calibration Amount Diff -39.342% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_P1 | 652752424 | pig1 | Unknown | C9H11NO3 | 15.224 | 0 | 42877112
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_P2 | 438971380 | pig2 | Unknown | C9H11NO3 | 8.764 | 0 | 50086201
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_P3 | 704712282 | pig3 | Unknown | C9H11NO3 | 14.178 | 0 | 49703701
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_P12 | 443044 |  | Unknown | C9H11NO3 | 4.689 | 0 | 94496
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_P13 | 30969 |  | Unknown | C9H11NO3 | 0.596 | 0 | 51990
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_P14 | 1350591 |  | Unknown | C9H11NO3 | 24.173 | 0 | 55871
Tyrosine pos | 8.39 | Target Compound | KD_050321_050321_P15 | 57182 |  | Unknown | C9H11NO3 | 0.98 | 0 | 58345

## Uric acid neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_AAmix | 68309 |  | Unknown | C5H4N4O3 | 0 | 0 | 0
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D1 | 3490642945 |  | Cal Std | C5H4N4O3 | 4295.262 | 100 | 812673 | Rsquared 0.906 < R2Threshold 0.99
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D2 | 2165023474 |  | Cal Std | C5H4N4O3 | 4435.46 | 50 | 488117 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff 70.453% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D3 | 1415501345 |  | Cal Std | C5H4N4O3 | 1284.589 | 25 | 1101910 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -20.802% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D4 | 640436707 |  | Cal Std | C5H4N4O3 | 549.04 | 12.5 | 1166467 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -34.599% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D5 | 320852606 |  | Cal Std | C5H4N4O3 | 282.937 | 6.25 | 1134008 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -33.375% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D6 | 169895062 |  | Cal Std | C5H4N4O3 | 134.548 | 3.125 | 1262706 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -37.033% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D7 | 72927918 |  | Cal Std | C5H4N4O3 | 82.45 | 1.563 | 884506 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -23.022% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D8 | 44618834 |  | Cal Std | C5H4N4O3 | 37.883 | 0.781 | 1177808 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -29.349% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D9 | 19599008 |  | Cal Std | C5H4N4O3 | 17.605 | 0.391 | 1113240 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -34.472% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D10 | 9978712 |  | Cal Std | C5H4N4O3 | 8.52 | 0.195 | 1171194 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -36.437% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D11 | 4447259 |  | Cal Std | C5H4N4O3 | 3.678 | 0.098 | 1209270 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -45.418% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_D12 | 1822515 |  | Cal Std | C5H4N4O3 | 1.649 | 0.049 | 1105411 | Rsquared 0.906 < R2Threshold 0.99. Calibration Amount Diff -51.065% (absolute) > max calibration amount diff 20%
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_P1 | 125994216 | pig1 | Unknown | C5H4N4O3 | 39.839 | 0 | 3162599
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_P2 | 52767834 | pig2 | Unknown | C5H4N4O3 | 18.271 | 0 | 2888124
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_P3 | 76419646 | pig3 | Unknown | C5H4N4O3 | 24.178 | 0 | 3160757
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_P12 | 1245281908 |  | Unknown | C5H4N4O3 | 0 | 0 | 0
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_P13 | 144307 |  | Unknown | C5H4N4O3 | 0 | 0 | 0
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_P14 | 56458 |  | Unknown | C5H4N4O3 | 0 | 0 | 0
Uric acid neg | 8.64 | Target Compound | KD_050321_050321_P15 | 94016 |  | Unknown | C5H4N4O3 | 0 | 0 | 0

## Valine pos
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Valine pos | 7.94 | Target Compound | KD_050321_050321_AAmix | 2864492371 |  | Unknown | C5H11NO2 | 0 | 0 | 0
Valine pos | 7.94 | Target Compound | KD_050321_050321_D1 | 2795100689 |  | Cal Std | C5H11NO2 | 15.295 | 100 | 182741955 | Rsquared 0.987 < R2Threshold 0.99
Valine pos | 7.94 | Target Compound | KD_050321_050321_D2 | 1674300950 |  | Cal Std | C5H11NO2 | 7.358 | 50 | 227557052 | Rsquared 0.987 < R2Threshold 0.99
Valine pos | 7.94 | Target Compound | KD_050321_050321_D3 | 1066706014 |  | Cal Std | C5H11NO2 | 4.126 | 25 | 258561190 | Rsquared 0.987 < R2Threshold 0.99
Valine pos | 7.94 | Target Compound | KD_050321_050321_D4 | 586478891 |  | Cal Std | C5H11NO2 | 2.013 | 12.5 | 291362094 | Rsquared 0.987 < R2Threshold 0.99
Valine pos | 7.94 | Target Compound | KD_050321_050321_D5 | 345007334 |  | Cal Std | C5H11NO2 | 0.965 | 6.25 | 357583415 | Rsquared 0.987 < R2Threshold 0.99
Valine pos | 7.94 | Target Compound | KD_050321_050321_D6 | 201746785 |  | Cal Std | C5H11NO2 | 0.549 | 3.125 | 367482535 | Rsquared 0.987 < R2Threshold 0.99
Valine pos | 7.94 | Target Compound | KD_050321_050321_D7 | 106978189 |  | Cal Std | C5H11NO2 | 0.293 | 1.563 | 365707915 | Rsquared 0.987 < R2Threshold 0.99
Valine pos | 7.94 | Target Compound | KD_050321_050321_D8 | 61593297 |  | Cal Std | C5H11NO2 | 0.172 | 0.781 | 358910246 | Rsquared 0.987 < R2Threshold 0.99. Calibration Amount Diff 33.512% (absolute) > max calibration amount diff 20%
Valine pos | 7.94 | Target Compound | KD_050321_050321_D9 | 46146349 |  | Cal Std | C5H11NO2 | 0.12 | 0.391 | 385208745 | Rsquared 0.987 < R2Threshold 0.99. Calibration Amount Diff 86.111% (absolute) > max calibration amount diff 20%
Valine pos | 7.94 | Target Compound | KD_050321_050321_D10 | 25737481 |  | Cal Std | C5H11NO2 | 0.068 | 0.195 | 379641536 | Rsquared 0.987 < R2Threshold 0.99. Calibration Amount Diff 111.128% (absolute) > max calibration amount diff 20%
Valine pos | 7.94 | Target Compound | KD_050321_050321_D11 | 15618998 |  | Cal Std | C5H11NO2 | 0.044 | 0.098 | 352944900 | Rsquared 0.987 < R2Threshold 0.99. Calibration Amount Diff 174.192% (absolute) > max calibration amount diff 20%
Valine pos | 7.94 | Target Compound | KD_050321_050321_D12 | 16333584 |  | Cal Std | C5H11NO2 | 0.043 | 0.049 | 376110176 | Rsquared 0.987 < R2Threshold 0.99. Calibration Amount Diff 438.151% (absolute) > max calibration amount diff 20%
Valine pos | 7.94 | Target Compound | KD_050321_050321_P1 | 4285377933 | pig1 | Unknown | C5H11NO2 | 14.144 | 0 | 302989023
Valine pos | 7.94 | Target Compound | KD_050321_050321_P2 | 3435300763 | pig2 | Unknown | C5H11NO2 | 11.019 | 0 | 311751282
Valine pos | 7.94 | Target Compound | KD_050321_050321_P3 | 4300284489 | pig3 | Unknown | C5H11NO2 | 14.479 | 0 | 296992157
Valine pos | 7.94 | Target Compound | KD_050321_050321_P12 | 3121712 |  | Unknown | C5H11NO2 | 0 | 0 | 0
Valine pos | 7.94 | Target Compound | KD_050321_050321_P13 | 9429686 |  | Unknown | C5H11NO2 | 0 | 0 | 0
Valine pos | 7.94 | Target Compound | KD_050321_050321_P14 | 0 |  | Unknown | C5H11NO2 | 0 | 0 | 0
Valine pos | 7.94 | Target Compound | KD_050321_050321_P15 | 0 |  | Unknown | C5H11NO2 | 0 | 0 | 0

## Xanthine neg
Compound | RT | Type | Filename | Area | Sample ID | Sample Type | Formula | Response Ratio | Theoretical Amt | ISTD Response | Flag Details
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_AAmix | 89376 |  | Unknown | C5H4N4O2 | 0 | 0 | 0
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D1 | 1789113221 |  | Cal Std | C5H4N4O2 | 163.71 | 100 | 10928559
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D2 | 1170922519 |  | Cal Std | C5H4N4O2 | 72.025 | 50 | 16257072
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D3 | 774094078 |  | Cal Std | C5H4N4O2 | 39.239 | 25 | 19727523
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D4 | 513883896 |  | Cal Std | C5H4N4O2 | 18.444 | 12.5 | 27861539
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D5 | 347730185 |  | Cal Std | C5H4N4O2 | 9.37 | 6.25 | 37109929
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D6 | 212510075 |  | Cal Std | C5H4N4O2 | 4.901 | 3.125 | 43361251
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D7 | 122476988 |  | Cal Std | C5H4N4O2 | 2.407 | 1.563 | 50882335
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D8 | 66838042 |  | Cal Std | C5H4N4O2 | 1.395 | 0.781 | 47898623 | Calibration Amount Diff 22.586% (absolute) > max calibration amount diff 20%
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D9 | 35546197 |  | Cal Std | C5H4N4O2 | 0.651 | 0.391 | 54612624
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D10 | 16819968 |  | Cal Std | C5H4N4O2 | 0.322 | 0.195 | 52244693
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D11 | 9004304 |  | Cal Std | C5H4N4O2 | 0.165 | 0.098 | 54528472
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_D12 | 4179149 |  | Cal Std | C5H4N4O2 | 0.075 | 0.049 | 55724396
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_P1 | 21585522 | pig1 | Unknown | C5H4N4O2 | 0.454 | 0 | 47507142
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_P2 | 25378938 | pig2 | Unknown | C5H4N4O2 | 0.516 | 0 | 49230783
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_P3 | 30191993 | pig3 | Unknown | C5H4N4O2 | 0.564 | 0 | 53497788
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_P12 | 1893404554 |  | Unknown | C5H4N4O2 | 0 | 0 | 0
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_P13 | 739906 |  | Unknown | C5H4N4O2 | 0 | 0 | 0
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_P14 | 416929 |  | Unknown | C5H4N4O2 | 0 | 0 | 0
Xanthine neg | 7.64 | Target Compound | KD_050321_050321_P15 | 212214 |  | Unknown | C5H4N4O2 | 92.617 | 0 | 2291