## Alanine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Alanine neg | 9.49 | Target Compound | KD100520_100520_A1 | 27031314 |  | 341782582 | C3H7NO2 | 162.49
Alanine neg | 9.49 | Target Compound | KD100520_100520_A2 | 13474082 |  | 162449157 | C3H7NO2 | 64.93
Alanine neg | 9.49 | Target Compound | KD100520_100520_A3 | 7167527 |  | 91616719 | C3H7NO2 | 55.77
Alanine neg | 9.49 | Target Compound | KD100520_100520_A4 | 3361236 |  | 44798690 | C3H7NO2 | 45.56
Alanine neg | 9.49 | Target Compound | KD100520_100520_A5 | 2442635 |  | 29732508 | C3H7NO2 | 51.19 | Calibration Amount Diff 36.447% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_A6 | 972129 |  | 12479348 | C3H7NO2 | 48.13
Alanine neg | 9.49 | Target Compound | KD100520_100520_A7 | 439159 |  | 5593201 | C3H7NO2 | 61.16
Alanine neg | 9.49 | Target Compound | KD100520_100520_A8 | 216373 |  | 2639503 | C3H7NO2 | 58.26 | Calibration Amount Diff -51.019% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_A9 | 106530 |  | 1074647 | C3H7NO2 | 63.89 | Calibration Amount Diff -132.616% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | C3H7NO2 | N/A
Alanine neg | 9.49 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | C3H7NO2 | N/A
Alanine neg | 9.49 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | C3H7NO2 | N/A
Alanine neg | 9.49 | Target Compound | KD100520_100520_B1 | 29581682 |  | 363785959 | C3H7NO2 | 223.80
Alanine neg | 9.49 | Target Compound | KD100520_100520_B2 | 13027084 |  | 168648467 | C3H7NO2 | 70.23
Alanine neg | 9.49 | Target Compound | KD100520_100520_B3 | 7848416 |  | 95903196 | C3H7NO2 | 55.19
Alanine neg | 9.49 | Target Compound | KD100520_100520_B4 | 3988334 |  | 50586813 | C3H7NO2 | 49.77
Alanine neg | 9.49 | Target Compound | KD100520_100520_B5 | 1708691 |  | 21975169 | C3H7NO2 | 46.77
Alanine neg | 9.49 | Target Compound | KD100520_100520_B6 | 887615 |  | 11426859 | C3H7NO2 | 56.19
Alanine neg | 9.49 | Target Compound | KD100520_100520_B7 | 419688 |  | 5302544 | C3H7NO2 | 56.51
Alanine neg | 9.49 | Target Compound | KD100520_100520_B8 | 199430 |  | 2556831 | C3H7NO2 | 55.46 | Calibration Amount Diff -51.158% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_B9 | 107459 |  | 1323664 | C3H7NO2 | 69.56 | Calibration Amount Diff -110.396% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_B10 | 31291 |  | 401864 | C3H7NO2 | 34.23 | Calibration Amount Diff -284.644% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | C3H7NO2 | N/A
Alanine neg | 9.49 | Target Compound | KD100520_100520_B12 | N/F |  | N/F | C3H7NO2 | N/A
Alanine neg | 9.49 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C3H7NO2 | N/A
Alanine neg | 9.49 | Target Compound | KD100520_100520_C1 | 24256459 |  | 315560049 | C3H7NO2 | 111.04
Alanine neg | 9.49 | Target Compound | KD100520_100520_C2 | 13518104 |  | 175432958 | C3H7NO2 | 52.62
Alanine neg | 9.49 | Target Compound | KD100520_100520_C3 | 7021228 |  | 90357153 | C3H7NO2 | 73.33
Alanine neg | 9.49 | Target Compound | KD100520_100520_C4 | 3478356 |  | 43279145 | C3H7NO2 | 58.47
Alanine neg | 9.49 | Target Compound | KD100520_100520_C5 | 1749827 |  | 22240805 | C3H7NO2 | 51.78
Alanine neg | 9.49 | Target Compound | KD100520_100520_C6 | 833778 |  | 10474000 | C3H7NO2 | 54.40
Alanine neg | 9.49 | Target Compound | KD100520_100520_C7 | 416193 |  | 5281732 | C3H7NO2 | 51.53
Alanine neg | 9.49 | Target Compound | KD100520_100520_C8 | 189978 |  | 2277989 | C3H7NO2 | 61.48 | Calibration Amount Diff -60.524% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_C9 | 88019 |  | 1006686 | C3H7NO2 | 56.09 | Calibration Amount Diff -137.617% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_C10 | 26687 |  | 334959 | C3H7NO2 | 35.80 | Calibration Amount Diff -294.02% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_C11 | 24246 |  | 275718 | C3H7NO2 | 33.51 | Calibration Amount Diff -516.12% (absolute) > max calibration amount diff 20%
Alanine neg | 9.49 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | C3H7NO2 | N/A

## Alanine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Alanine pos | 9.49 | Target Compound | KD100520_100520_A1 | 133906311 |  | 1710833085 | C3H7NO2 | 40.17
Alanine pos | 9.49 | Target Compound | KD100520_100520_A2 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_A3 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_A4 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_A5 | 20866301 |  | 242610268 | C3H7NO2 | 31.04
Alanine pos | 9.49 | Target Compound | KD100520_100520_A6 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_A7 | 5452765 |  | 69362110 | C3H7NO2 | 32.86 | Calibration Amount Diff 24.985% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD100520_100520_A8 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_A9 | 1445930 |  | 17551988 | C3H7NO2 | 78.31 | Calibration Amount Diff 79.275% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_B1 | 141497186 |  | 1790124936 | C3H7NO2 | 57.66
Alanine pos | 9.49 | Target Compound | KD100520_100520_B2 | 81074191 |  | 1017919181 | C3H7NO2 | 41.18
Alanine pos | 9.49 | Target Compound | KD100520_100520_B3 | 52459682 |  | 652770299 | C3H7NO2 | 33.30
Alanine pos | 9.49 | Target Compound | KD100520_100520_B4 | 32239244 |  | 422242228 | C3H7NO2 | 31.60
Alanine pos | 9.49 | Target Compound | KD100520_100520_B5 | 15829459 |  | 188691010 | C3H7NO2 | 31.64
Alanine pos | 9.49 | Target Compound | KD100520_100520_B6 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_B7 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_B8 | 2727561 |  | 35336816 | C3H7NO2 | 38.04 | Calibration Amount Diff 43.903% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD100520_100520_B9 | 1355290 |  | 16943221 | C3H7NO2 | 39.77 | Calibration Amount Diff 76.276% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD100520_100520_B10 | 656535 |  | 7313601 | C3H7NO2 | 30.85 | Calibration Amount Diff 125.198% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_B12 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_C1 | 142337224 |  | 1841708139 | C3H7NO2 | 50.70
Alanine pos | 9.49 | Target Compound | KD100520_100520_C2 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_C3 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_C4 | 29081263 |  | 380021718 | C3H7NO2 | 56.75
Alanine pos | 9.49 | Target Compound | KD100520_100520_C5 | 17782632 |  | 215220734 | C3H7NO2 | 66.83
Alanine pos | 9.49 | Target Compound | KD100520_100520_C6 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_C7 | 5301345 |  | 69629970 | C3H7NO2 | 30.29 | Calibration Amount Diff 24.35% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD100520_100520_C8 | 2414275 |  | 27278143 | C3H7NO2 | 33.94 | Calibration Amount Diff 20.032% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD100520_100520_C9 | 1197528 |  | 13848969 | C3H7NO2 | 31.97 | Calibration Amount Diff 54.947% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD100520_100520_C10 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_C11 | N/F |  | N/F | C3H7NO2 | N/A
Alanine pos | 9.49 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | C3H7NO2 | N/A

## Alanine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A1 | 2662806 |  | 34348814 | [13]C3H7NO2 | 3393.21 | Rsquared 0.321 < R2Threshold 0.99
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A2 | 2933888 |  | 38213395 | [13]C3H7NO2 | 1850.94 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -114.183% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A3 | 2535431 |  | 35210489 | [13]C3H7NO2 | 2082.62 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 170.026% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A4 | 2310327 |  | 33628548 | [13]C3H7NO2 | 1458.21 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 643.697% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A5 | 2564940 |  | 33663903 | [13]C3H7NO2 | 2603.34 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 684.64% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A6 | 2281201 |  | 32593892 | [13]C3H7NO2 | 1525.92 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -484.625% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A7 | 2177397 |  | 31228460 | [13]C3H7NO2 | 913.66 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -1499.018% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A8 | 2221610 |  | 32091240 | [13]C3H7NO2 | 887.01 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -4012.895% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A9 | 2312394 |  | 31234718 | [13]C3H7NO2 | 1186.92 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -12666.976% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A10 | 2237596 |  | 29927407 | [13]C3H7NO2 | 1107.93 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -24181.479% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A11 | 2106345 |  | 29508989 | [13]C3H7NO2 | 1214.63 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 31836.795% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_A12 | 2213528 |  | 29922268 | [13]C3H7NO2 | 1408.10 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 86459.226% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B1 | 2530722 |  | 32432610 | [13]C3H7NO2 | 3425.50 | Rsquared 0.321 < R2Threshold 0.99
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B2 | 2583673 |  | 35529473 | [13]C3H7NO2 | 2023.55 | Rsquared 0.321 < R2Threshold 0.99
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B3 | 2826793 |  | 37873121 | [13]C3H7NO2 | 2489.66 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 114.81% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B4 | 2584653 |  | 36240489 | [13]C3H7NO2 | 1194.88 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 194.696% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B5 | 2490315 |  | 34916696 | [13]C3H7NO2 | 1038.27 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 40.615% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B6 | 2231422 |  | 31769830 | [13]C3H7NO2 | 1389.21 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -209.126% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B7 | 2172679 |  | 31155692 | [13]C3H7NO2 | 1021.80 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -1159.36% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B8 | 2074541 |  | 30775298 | [13]C3H7NO2 | 1000.23 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -1673.014% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B9 | 2166173 |  | 30194360 | [13]C3H7NO2 | 978.28 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 1275.254% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B10 | 2104194 |  | 29965825 | [13]C3H7NO2 | 1810.94 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 4745.69% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B11 | 1915108 |  | 28171075 | [13]C3H7NO2 | 1273.03 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 26274.619% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_B12 | 1912427 |  | 27448830 | [13]C3H7NO2 | 1507.80 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 76048.121% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C3H7NO2 | N/A
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C1 | 2354982 |  | 31915740 | [13]C3H7NO2 | 5916.52 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 32.839% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C2 | 2573403 |  | 34916639 | [13]C3H7NO2 | 1231.34 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -91.024% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C3 | 2649754 |  | 36608553 | [13]C3H7NO2 | 1699.42 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -484.397% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C4 | 2555153 |  | 34603026 | [13]C3H7NO2 | 1629.23 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -537.902% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C5 | 2353846 |  | 32745942 | [13]C3H7NO2 | 1286.68 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -909.974% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C6 | 2254030 |  | 31352002 | [13]C3H7NO2 | 984.61 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff -2085.502% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C7 | 2056137 |  | 30110180 | [13]C3H7NO2 | 1094.68 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 3145.186% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C8 | 2059639 |  | 29183104 | [13]C3H7NO2 | 1208.19 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 3249.865% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C9 | 2097816 |  | 29428873 | [13]C3H7NO2 | 1001.00 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 7064.568% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C10 | 1947192 |  | 28613476 | [13]C3H7NO2 | 1120.47 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 7743.151% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C11 | 2135069 |  | 29643727 | [13]C3H7NO2 | 773.22 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 47227.168% (absolute) > max calibration amount diff 20%
Alanine U-13C neg | 9.49 | Target Compound | KD100520_100520_C12 | 1903806 |  | 26485018 | [13]C3H7NO2 | 878.80 | Rsquared 0.321 < R2Threshold 0.99. Calibration Amount Diff 41957.122% (absolute) > max calibration amount diff 20%

## Alanine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A1 | 13886959 |  | 178408114 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -240.868% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A2 | 17607941 |  | 237118646 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -287.345% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A3 | 19941389 |  | 282344353 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -59.961% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A4 | 22421216 |  | 325533034 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -23.956% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A5 | 25265575 |  | 342331115 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -3557.161% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A6 | 27640069 |  | 395889816 | [13]C3H7NO2 | 1028264.34 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 18237.673% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A7 | 29779123 |  | 417653028 | [13]C3H7NO2 | 111257.38 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 25062.539% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A8 | 26500611 |  | 392444843 | [13]C3H7NO2 | 142571.83 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 16031.179% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A9 | 27730909 |  | 390289804 | [13]C3H7NO2 | 43971.83 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 23301.002% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A10 | 27082670 |  | 386822635 | [13]C3H7NO2 | 66781.68 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 2312.595% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A11 | 25630115 |  | 366399603 | [13]C3H7NO2 | 48874.56 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -535689.431% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_A12 | 25717813 |  | 376555111 | [13]C3H7NO2 | 121828.05 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -900973.999% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B1 | 12833318 |  | 160452195 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -275.724% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B2 | 17060616 |  | 224676377 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 809.489% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B3 | 20885714 |  | 285011601 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 1355.943% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B4 | 23510939 |  | 328931504 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 1333.687% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B5 | 27489615 |  | 379888336 | [13]C3H7NO2 | 34326.60 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 3829.141% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B6 | 27652913 |  | 385395324 | [13]C3H7NO2 | 24365.80 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 2217.685% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B7 | 28302774 |  | 413982293 | [13]C3H7NO2 | 1903278.43 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 2571.644% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B8 | 27974427 |  | 400641758 | [13]C3H7NO2 | 122942.12 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 1975.488% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B9 | 25687800 |  | 381524656 | [13]C3H7NO2 | 59082.26 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -13890.577% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B10 | 27313165 |  | 381021763 | [13]C3H7NO2 | 451082.08 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -56796.109% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B11 | 23992558 |  | 352314107 | [13]C3H7NO2 | 194499.80 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -432203.07% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_B12 | 24198376 |  | 350466037 | [13]C3H7NO2 | 53587.60 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -952763.332% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C3H7NO2 | N/A
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C1 | 13859620 |  | 176531443 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -441.325% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C2 | 16956993 |  | 235647775 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 1544.697% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C3 | 19750064 |  | 284161020 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 2951.476% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C4 | 23473926 |  | 335677495 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 5729.097% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C5 | 27169942 |  | 366822069 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 8381.573% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C6 | 27863139 |  | 384417556 | [13]C3H7NO2 | 1351454.36 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff 14502.075% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C7 | 27774719 |  | 399210439 | [13]C3H7NO2 | INF | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -11818.6% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C8 | 27084059 |  | 381257636 | [13]C3H7NO2 | 89825.11 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -39407.795% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C9 | 26091801 |  | 370468761 | [13]C3H7NO2 | 43534.65 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -117471.876% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C10 | 24564854 |  | 356053300 | [13]C3H7NO2 | 129171.26 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -217338.95% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C11 | 25273284 |  | 371931531 | [13]C3H7NO2 | 53462.34 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -490433.816% (absolute) > max calibration amount diff 20%
Alanine U-13C pos | 9.49 | Target Compound | KD100520_100520_C12 | 24123101 |  | 338977387 | [13]C3H7NO2 | 268063.87 | Rsquared 0.006 < R2Threshold 0.99. Calibration Amount Diff -920607.661% (absolute) > max calibration amount diff 20%

## Alanine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A1 | 3712990 |  | 49739428 | [13]C3H7[15]NO2 | 6411.28
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A2 | 3849360 |  | 51046974 | [13]C3H7[15]NO2 | 2955.09
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A3 | 3489597 |  | 49325515 | [13]C3H7[15]NO2 | 2846.73
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A4 | 3199498 |  | 47905276 | [13]C3H7[15]NO2 | 2294.82
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A5 | 3438783 |  | 46597357 | [13]C3H7[15]NO2 | 1742.18
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A6 | 2949394 |  | 43407096 | [13]C3H7[15]NO2 | 1162.77
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A7 | 2883079 |  | 41336274 | [13]C3H7[15]NO2 | 714.19
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A8 | 2939604 |  | 42251464 | [13]C3H7[15]NO2 | 1360.79
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A9 | 3003969 |  | 40661879 | [13]C3H7[15]NO2 | 1326.09
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A10 | 2862263 |  | 39008058 | [13]C3H7[15]NO2 | 702.85
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A11 | 2919813 |  | 40380821 | [13]C3H7[15]NO2 | 912.60
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_A12 | 3015904 |  | 41237403 | [13]C3H7[15]NO2 | 823.09
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B1 | 3698605 |  | 46818546 | [13]C3H7[15]NO2 | 7953.21
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B2 | 3546448 |  | 49172463 | [13]C3H7[15]NO2 | 8032.03
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B3 | 3829219 |  | 52581829 | [13]C3H7[15]NO2 | 2969.87
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B4 | 3644036 |  | 49772182 | [13]C3H7[15]NO2 | 1243.51
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B5 | 3368895 |  | 47108675 | [13]C3H7[15]NO2 | 982.16
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B6 | 2969940 |  | 42536858 | [13]C3H7[15]NO2 | 1862.99
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B7 | 2843394 |  | 41375373 | [13]C3H7[15]NO2 | 1278.60
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B8 | 2745026 |  | 40978337 | [13]C3H7[15]NO2 | 865.11
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B9 | 2885362 |  | 40650126 | [13]C3H7[15]NO2 | 1587.66
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B10 | 2795634 |  | 40446379 | [13]C3H7[15]NO2 | 1020.88
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B11 | 2635346 |  | 38417031 | [13]C3H7[15]NO2 | 1434.93
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_B12 | 2596701 |  | 37705636 | [13]C3H7[15]NO2 | 1395.62
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C3H7[15]NO2 | N/A
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C1 | 3445577 |  | 46701273 | [13]C3H7[15]NO2 | 1678.08
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C2 | 3350147 |  | 46981592 | [13]C3H7[15]NO2 | 3151.13
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C3 | 3264873 |  | 46336458 | [13]C3H7[15]NO2 | 991.53
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C4 | 3267315 |  | 44892937 | [13]C3H7[15]NO2 | 1768.92
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C5 | 2986701 |  | 42589521 | [13]C3H7[15]NO2 | 2092.34
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C6 | 2856696 |  | 40496251 | [13]C3H7[15]NO2 | 1239.07
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C7 | 2806823 |  | 41723227 | [13]C3H7[15]NO2 | 788.84
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C8 | 2795305 |  | 39807951 | [13]C3H7[15]NO2 | 1165.14
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C9 | 2847263 |  | 40189570 | [13]C3H7[15]NO2 | 1044.78
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C10 | 2614350 |  | 38763813 | [13]C3H7[15]NO2 | 802.83
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C11 | 2952114 |  | 40957414 | [13]C3H7[15]NO2 | 649.45
Alanine U-13C, U-15N neg | 9.49 | Internal Standard | KD100520_100520_C12 | 2589038 |  | 35998414 | [13]C3H7[15]NO2 | 760.98

## Alanine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A1 | 20249861 |  | 260126913 | [13]C3H7[15]NO2 | INF
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A2 | 26126489 |  | 344146454 | [13]C3H7[15]NO2 | INF
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A3 | 28864291 |  | 405704678 | [13]C3H7[15]NO2 | INF
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A4 | 32898415 |  | 467785772 | [13]C3H7[15]NO2 | 7145.71
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A5 | 36140071 |  | 502819321 | [13]C3H7[15]NO2 | 9910.43
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A6 | 38858608 |  | 539674146 | [13]C3H7[15]NO2 | 10312.28
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A7 | 41602112 |  | 578828254 | [13]C3H7[15]NO2 | 15747.62
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A8 | 37364488 |  | 557695562 | [13]C3H7[15]NO2 | 16108.06
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A9 | 39902452 |  | 556462664 | [13]C3H7[15]NO2 | 13269.11
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A10 | 39746712 |  | 556113884 | [13]C3H7[15]NO2 | 12577.40
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A11 | 39276029 |  | 554878011 | [13]C3H7[15]NO2 | 14839.42
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_A12 | 39092185 |  | 565499985 | [13]C3H7[15]NO2 | 15405.67
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B1 | 18728638 |  | 234743837 | [13]C3H7[15]NO2 | 199808.17 | 9.490:Peak area 234743836.919 is out of bounds (ISTD Minimum Recovery 242509167.665 and ISTD Max Recovery 727527502.996)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B2 | 24106980 |  | 309614994 | [13]C3H7[15]NO2 | INF
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B3 | 29196604 |  | 396071124 | [13]C3H7[15]NO2 | 6350.38
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B4 | 33072304 |  | 465087933 | [13]C3H7[15]NO2 | 5732.52
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B5 | 39621603 |  | 533789697 | [13]C3H7[15]NO2 | 9544.18
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B6 | 39841417 |  | 550479848 | [13]C3H7[15]NO2 | 11477.69
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B7 | 40940018 |  | 593049297 | [13]C3H7[15]NO2 | 15601.24
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B8 | 41071086 |  | 575345080 | [13]C3H7[15]NO2 | 22077.78
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B9 | 37815338 |  | 551602503 | [13]C3H7[15]NO2 | 16052.15
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B10 | 40305489 |  | 553922405 | [13]C3H7[15]NO2 | 13058.40
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B11 | 36597387 |  | 528146818 | [13]C3H7[15]NO2 | 17056.67
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_B12 | 36895604 |  | 527658839 | [13]C3H7[15]NO2 | 13472.32
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C3H7[15]NO2 | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C1 | 20957925 |  | 262519474 | [13]C3H7[15]NO2 | 212210.16
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C2 | 22950243 |  | 314098182 | [13]C3H7[15]NO2 | INF
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C3 | 27014127 |  | 380780102 | [13]C3H7[15]NO2 | 148273.73
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C4 | 32131720 |  | 451196446 | [13]C3H7[15]NO2 | 18037.68
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C5 | 37124986 |  | 502014110 | [13]C3H7[15]NO2 | 10535.97
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C6 | 38308629 |  | 529669518 | [13]C3H7[15]NO2 | 9146.71
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C7 | 41057858 |  | 584472966 | [13]C3H7[15]NO2 | 21787.00
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C8 | 40888134 |  | 565043379 | [13]C3H7[15]NO2 | 11221.31
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C9 | 39334613 |  | 557413216 | [13]C3H7[15]NO2 | 23266.87
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C10 | 37532643 |  | 533899413 | [13]C3H7[15]NO2 | 22591.80
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C11 | 38252230 |  | 560748346 | [13]C3H7[15]NO2 | 13554.78
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD100520_100520_C12 | 37092724 |  | 509556903 | [13]C3H7[15]NO2 | 9529.76

## Arginine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Arginine neg | 14.24 | Target Compound | KD100520_100520_A1 | 61437177 |  | 1296243142 | C6H14N4O2 | 49103.57
Arginine neg | 14.24 | Target Compound | KD100520_100520_A2 | 32772030 |  | 682091227 | C6H14N4O2 | 76434.67
Arginine neg | 14.24 | Target Compound | KD100520_100520_A3 | 21976415 |  | 431662390 | C6H14N4O2 | 27981.55
Arginine neg | 14.24 | Target Compound | KD100520_100520_A4 | 9938236 |  | 216131027 | C6H14N4O2 | 13965.21
Arginine neg | 14.24 | Target Compound | KD100520_100520_A5 | 6994753 |  | 146687280 | C6H14N4O2 | 2260.01 | Calibration Amount Diff 43.155% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_A6 | 3166998 |  | 67983110 | C6H14N4O2 | 2001.20
Arginine neg | 14.24 | Target Compound | KD100520_100520_A7 | 1430632 |  | 30547481 | C6H14N4O2 | 1086.17 | Calibration Amount Diff -28.929% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_A8 | 743542 |  | 14694202 | C6H14N4O2 | 744.17 | Calibration Amount Diff -81.847% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_A9 | 416572 |  | 7255543 | C6H14N4O2 | 610.33 | Calibration Amount Diff -186.087% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_A10 | 160582 |  | 2860543 | C6H14N4O2 | 450.15 | Calibration Amount Diff -412.258% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_A11 | 76278 |  | 1352876 | C6H14N4O2 | 517.42 | Calibration Amount Diff -823.753% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_A12 | 18466 |  | 466312 | C6H14N4O2 | 1007.38 | Calibration Amount Diff -1659.348% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_B1 | 64189615 |  | 1395912094 | C6H14N4O2 | 28630.06
Arginine neg | 14.24 | Target Compound | KD100520_100520_B2 | 28776061 |  | 709512462 | C6H14N4O2 | 38227.27
Arginine neg | 14.24 | Target Compound | KD100520_100520_B3 | 18679963 |  | 424408802 | C6H14N4O2 | 25024.93
Arginine neg | 14.24 | Target Compound | KD100520_100520_B4 | 10491356 |  | 216056155 | C6H14N4O2 | 17643.50
Arginine neg | 14.24 | Target Compound | KD100520_100520_B5 | 5161819 |  | 108136311 | C6H14N4O2 | 3069.75
Arginine neg | 14.24 | Target Compound | KD100520_100520_B6 | 2877574 |  | 60666222 | C6H14N4O2 | 1847.75
Arginine neg | 14.24 | Target Compound | KD100520_100520_B7 | 1391005 |  | 29367184 | C6H14N4O2 | 1395.33 | Calibration Amount Diff -29.381% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_B8 | 678732 |  | 13867281 | C6H14N4O2 | 557.77 | Calibration Amount Diff -81.371% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_B9 | 345177 |  | 6692848 | C6H14N4O2 | 929.24 | Calibration Amount Diff -188.36% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_B10 | 159632 |  | 2836892 | C6H14N4O2 | 1200.96 | Calibration Amount Diff -416.804% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_B11 | 58996 |  | 1251571 | C6H14N4O2 | 376.79 | Calibration Amount Diff -830.477% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_B12 | 28578 |  | 545230 | C6H14N4O2 | 398.07 | Calibration Amount Diff -1649.779% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_blank | 1162 |  | 4391 | C6H14N4O2 | INF
Arginine neg | 14.24 | Target Compound | KD100520_100520_C1 | 62497496 |  | 1324785911 | C6H14N4O2 | 24389.43
Arginine neg | 14.24 | Target Compound | KD100520_100520_C2 | 33864584 |  | 710500296 | C6H14N4O2 | 32015.34
Arginine neg | 14.24 | Target Compound | KD100520_100520_C3 | 18090406 |  | 389619999 | C6H14N4O2 | 33087.16
Arginine neg | 14.24 | Target Compound | KD100520_100520_C4 | 10206409 |  | 209813719 | C6H14N4O2 | 7703.51
Arginine neg | 14.24 | Target Compound | KD100520_100520_C5 | 5261034 |  | 112556936 | C6H14N4O2 | 1720.96
Arginine neg | 14.24 | Target Compound | KD100520_100520_C6 | 2264332 |  | 49004942 | C6H14N4O2 | 1092.65
Arginine neg | 14.24 | Target Compound | KD100520_100520_C7 | 1255279 |  | 27539271 | C6H14N4O2 | 1781.61 | Calibration Amount Diff -23.996% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_C8 | 720673 |  | 13797624 | C6H14N4O2 | 775.33 | Calibration Amount Diff -87.931% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_C9 | 335506 |  | 6315148 | C6H14N4O2 | 614.05 | Calibration Amount Diff -198.905% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_C10 | 113644 |  | 2171788 | C6H14N4O2 | 371.27 | Calibration Amount Diff -434.125% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_C11 | 63243 |  | 1212521 | C6H14N4O2 | 387.42 | Calibration Amount Diff -833.353% (absolute) > max calibration amount diff 20%
Arginine neg | 14.24 | Target Compound | KD100520_100520_C12 | 29914 |  | 527287 | C6H14N4O2 | 7076.19 | Calibration Amount Diff -1647.309% (absolute) > max calibration amount diff 20%

## Arginine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Arginine pos | 14.24 | Target Compound | KD100520_100520_A1 | 556035497 |  | 14218521310 | C6H14N4O2 | 6369.04
Arginine pos | 14.24 | Target Compound | KD100520_100520_A2 | 334680770 |  | 9007950979 | C6H14N4O2 | 4752.99
Arginine pos | 14.24 | Target Compound | KD100520_100520_A3 | 249597966 |  | 6201870354 | C6H14N4O2 | 7513.06 | Calibration Amount Diff 20.542% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A4 | 147099341 |  | 3726942767 | C6H14N4O2 | 7303.69 | Calibration Amount Diff 23.035% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A5 | 109056796 |  | 2656336061 | C6H14N4O2 | 13595.79 | Calibration Amount Diff 61.49% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A6 | 49563943 |  | 1323637419 | C6H14N4O2 | 4401.87 | Calibration Amount Diff 25.257% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A7 | 23664800 |  | 634095001 | C6H14N4O2 | 10597.94 | Calibration Amount Diff -50.379% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A8 | 12867435 |  | 323993508 | C6H14N4O2 | 9951.46 | Calibration Amount Diff -171.47% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A9 | 7077007 |  | 173719434 | C6H14N4O2 | 5523.46 | Calibration Amount Diff -404.304% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A10 | 3135718 |  | 84962130 | C6H14N4O2 | 874.45 | Calibration Amount Diff -915.085% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A11 | 1832326 |  | 47487653 | C6H14N4O2 | 1067.94 | Calibration Amount Diff -1890.237% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_A12 | 829956 |  | 20819077 | C6H14N4O2 | 310.05 | Calibration Amount Diff -3928.259% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_B1 | 572820387 |  | 14956334203 | C6H14N4O2 | 5806.92
Arginine pos | 14.24 | Target Compound | KD100520_100520_B2 | 348526839 |  | 9419250528 | C6H14N4O2 | 5655.04
Arginine pos | 14.24 | Target Compound | KD100520_100520_B3 | 238553239 |  | 6215397462 | C6H14N4O2 | 13561.67
Arginine pos | 14.24 | Target Compound | KD100520_100520_B4 | 165024830 |  | 3908401259 | C6H14N4O2 | 8599.51 | Calibration Amount Diff 28.229% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_B5 | 83271094 |  | 2095230568 | C6H14N4O2 | 9495.54
Arginine pos | 14.24 | Target Compound | KD100520_100520_B6 | 49167319 |  | 1246949061 | C6H14N4O2 | 7603.24
Arginine pos | 14.24 | Target Compound | KD100520_100520_B7 | 24224387 |  | 636628733 | C6H14N4O2 | 7885.02 | Calibration Amount Diff -51.415% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_B8 | 13446863 |  | 338820670 | C6H14N4O2 | 12797.97 | Calibration Amount Diff -163.478% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_B9 | 7265418 |  | 183220780 | C6H14N4O2 | 2641.36 | Calibration Amount Diff -402.305% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_B10 | 3344738 |  | 86944644 | C6H14N4O2 | 1073.79 | Calibration Amount Diff -917.094% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_B11 | 1681680 |  | 46927227 | C6H14N4O2 | 656.71 | Calibration Amount Diff -1900.898% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_B12 | 823744 |  | 21197545 | C6H14N4O2 | 353.81 | Calibration Amount Diff -3928.598% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_blank | 50574 |  | 1457742 | C6H14N4O2 | 48.99
Arginine pos | 14.24 | Target Compound | KD100520_100520_C1 | 528445801 |  | 14656307535 | C6H14N4O2 | 6813.81
Arginine pos | 14.24 | Target Compound | KD100520_100520_C2 | 350141510 |  | 9484707129 | C6H14N4O2 | 6746.72
Arginine pos | 14.24 | Target Compound | KD100520_100520_C3 | 229909177 |  | 5904284619 | C6H14N4O2 | 12973.38 | Calibration Amount Diff 20.27% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_C4 | 139952618 |  | 3747073071 | C6H14N4O2 | 9691.33 | Calibration Amount Diff 24.884% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_C5 | 85852523 |  | 2240393969 | C6H14N4O2 | 11443.25 | Calibration Amount Diff 25.919% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_C6 | 42572075 |  | 1077833007 | C6H14N4O2 | 15452.77
Arginine pos | 14.24 | Target Compound | KD100520_100520_C7 | 27118869 |  | 646947046 | C6H14N4O2 | 27497.75 | Calibration Amount Diff -45.834% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_C8 | 12204372 |  | 316265974 | C6H14N4O2 | 8842.27 | Calibration Amount Diff -177.128% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_C9 | 6815715 |  | 171765007 | C6H14N4O2 | 3340.57 | Calibration Amount Diff -418.867% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_C10 | 2932835 |  | 76463535 | C6H14N4O2 | 1288.00 | Calibration Amount Diff -940.763% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_C11 | 1585066 |  | 42035480 | C6H14N4O2 | 728.81 | Calibration Amount Diff -1919.841% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD100520_100520_C12 | 832610 |  | 19716887 | C6H14N4O2 | 526.89 | Calibration Amount Diff -3932.206% (absolute) > max calibration amount diff 20%

## Arg pos calibration
Compound | Filename | Sample ID | Area | Area_IS | Response ratio | Conc. | Conc. Lin reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss | Conc. power reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss
Arginine pos | KD100520_100520_A1 | A1 | 14218521310 | 2180623550 | 6.5204 | 500 | 472.2851 | 27.7149 | 5.543 | 423.0428 | 76.9572 | 15.3914
Arginine pos | KD100520_100520_A2 | A2 | 9007950979 | 2750276091 | 3.2753 | 250 | 237.2358 | 12.7642 | 5.1057 | 199.8423 | 50.1577 | 20.0631
Arginine pos | KD100520_100520_A3 | A3 | 6201870354 | 2908277164 | 2.1325 | 125 | 154.4605 | 29.4605 | 23.5684 | 125.2279 | 0.22787 | 0.1823
Arginine pos | KD100520_100520_A4 | A4 | 3726942767 | 3235173762 | 1.152 | 62.5 | 83.4422 | 20.9422 | 33.5075 | 64.0345 | 1.5345 | 2.4553
Arginine pos | KD100520_100520_A5 | A5 | 2656336061 | 3317258941 | 0.80076 | 31.25 | 58.0008 | 26.7508 | 85.6026 | 43.0897 | 11.8397 | 37.8871
Arginine pos | KD100520_100520_A6 | A6 | 1323637419 | 3392154190 | 0.39021 | 15.625 | 28.2634 | 12.6384 | 80.8856 | 19.6931 | 4.0681 | 26.0359
Arginine pos | KD100520_100520_A7 | A7 | 634095001 | 3490972249 | 0.18164 | 7.8125 | 13.1564 | 5.3439 | 68.4024 | 8.5626 | 0.75012 | 9.6015
Arginine pos | KD100520_100520_A8 | A8 | 323993508 | 3483012107 | 0.093021 | 3.9062 | 6.7377 | 2.8315 | 72.4852 | 4.131 | 0.22476 | 5.7538
Arginine pos | KD100520_100520_A9 | A9 | 173719434 | 3396849427 | 0.051141 | 1.9531 | 3.7043 | 1.7511 | 89.6586 | 2.1531 | 0.20001 | 10.2407
Arginine pos | KD100520_100520_A10 | A10 | 84962130 | 3497361113 | 0.024293 | 0.97656 | 1.7596 | 0.78304 | 80.1836 | 0.95708 | 0.019484 | 1.9952
Arginine pos | KD100520_100520_A11 | A11 | 47487653 | 3386855996 | 0.014021 | 0.48828 | 1.0156 | 0.5273 | 107.9909 | 0.52596 | 0.037681 | 7.7171
Arginine pos | KD100520_100520_A12 | A12 | 20819077 | 3477555078 | 0.0059867 | 0.24414 | 0.43363 | 0.18949 | 77.6143 | 0.20816 | 0.035984 | 14.7391
Arginine pos | KD100520_100520_B1 | B1 | 14956334203 | 2116300841 | 7.0672 | 500 | 511.8919 | 11.8919 | 2.3784 | 461.8257 | 38.1743 | 7.6349
Arginine pos | KD100520_100520_B2 | B2 | 9419250528 | 2698253696 | 3.4909 | 250 | 252.8506 | 2.8506 | 1.1403 | 214.2104 | 35.7896 | 14.3158
Arginine pos | KD100520_100520_B3 | B3 | 6215397462 | 2942409085 | 2.1123 | 125 | 153.0017 | 28.0017 | 22.4014 | 123.9403 | 1.0597 | 0.8478
Arginine pos | KD100520_100520_B4 | B4 | 3908401259 | 3270209731 | 1.1952 | 62.5 | 86.5673 | 24.0673 | 38.5077 | 66.6511 | 4.1511 | 6.6417
Arginine pos | KD100520_100520_B5 | B5 | 2095230568 | 3471639106 | 0.60353 | 31.25 | 43.7147 | 12.4647 | 39.8871 | 31.6674 | 0.41743 | 1.3358
Arginine pos | KD100520_100520_B6 | B6 | 1246949061 | 3573769956 | 0.34892 | 15.625 | 25.2728 | 9.6478 | 61.7456 | 17.4345 | 1.8095 | 11.581
Arginine pos | KD100520_100520_B7 | B7 | 636628733 | 3525800033 | 0.18056 | 7.8125 | 13.0785 | 5.266 | 67.4052 | 8.5074 | 0.69491 | 8.8948
Arginine pos | KD100520_100520_B8 | B8 | 338820670 | 3486906301 | 0.097169 | 3.9062 | 7.0382 | 3.1319 | 80.1773 | 4.3321 | 0.42581 | 10.9008
Arginine pos | KD100520_100520_B9 | B9 | 183220780 | 3546646710 | 0.05166 | 1.9531 | 3.7419 | 1.7887 | 91.5831 | 2.1769 | 0.22382 | 11.4597
Arginine pos | KD100520_100520_B10 | B10 | 86944644 | 3617798622 | 0.024032 | 0.97656 | 1.7407 | 0.76416 | 78.2497 | 0.9459 | 0.030667 | 3.1403
Arginine pos | KD100520_100520_B11 | B11 | 46927227 | 3520466502 | 0.01333 | 0.48828 | 0.96551 | 0.47723 | 97.7357 | 0.49778 | 0.0094976 | 1.9451
Arginine pos | KD100520_100520_B12 | B12 | 21197545 | 3547282220 | 0.0059757 | 0.24414 | 0.43283 | 0.18869 | 77.2884 | 0.20774 | 0.0364 | 14.9095
Arginine pos | KD100520_100520_C1 | C1 | 14656307535 | 2329471061 | 6.2917 | 500 | 455.7196 | 44.2804 | 8.8561 | 406.9066 | 93.0934 | 18.6187
Arginine pos | KD100520_100520_C2 | C2 | 9484707129 | 2508720143 | 3.7807 | 250 | 273.8433 | 23.8433 | 9.5373 | 233.6515 | 16.3485 | 6.5394
Arginine pos | KD100520_100520_C3 | C3 | 5904284619 | 2774612942 | 2.128 | 125 | 154.1329 | 29.1329 | 23.3063 | 124.9386 | 0.061367 | 0.049093
Arginine pos | KD100520_100520_C4 | C4 | 3747073071 | 3209840787 | 1.1674 | 62.5 | 84.555 | 22.055 | 35.288 | 64.9652 | 2.4652 | 3.9444
Arginine pos | KD100520_100520_C5 | C5 | 2240393969 | 3430696373 | 0.65304 | 31.25 | 47.3012 | 16.0512 | 51.364 | 34.5074 | 3.2574 | 10.4237
Arginine pos | KD100520_100520_C6 | C6 | 1077833007 | 3303743825 | 0.32625 | 15.625 | 23.6306 | 8.0056 | 51.2361 | 16.2043 | 0.57932 | 3.7076
Arginine pos | KD100520_100520_C7 | C7 | 646947046 | 3471540009 | 0.18636 | 7.8125 | 13.4982 | 5.6857 | 72.7774 | 8.8052 | 0.99269 | 12.7065
Arginine pos | KD100520_100520_C8 | C8 | 316265974 | 3510791889 | 0.090084 | 3.9062 | 6.525 | 2.6187 | 67.039 | 3.9891 | 0.082891 | 2.122
Arginine pos | KD100520_100520_C9 | C9 | 171765007 | 3626659020 | 0.047362 | 1.9531 | 3.4305 | 1.4774 | 75.642 | 1.9804 | 0.027278 | 1.3966
Arginine pos | KD100520_100520_C10 | C10 | 76463535 | 3648175047 | 0.020959 | 0.97656 | 1.5181 | 0.54157 | 55.4566 | 0.81494 | 0.16163 | 16.5506
Arginine pos | KD100520_100520_C11 | C11 | 42035480 | 3473614690 | 0.012101 | 0.48828 | 0.87653 | 0.38824 | 79.5125 | 0.44802 | 0.040258 | 8.2448
Arginine pos | KD100520_100520_C12 | C12 | 19716887 | 3365387677 | 0.0058587 | 0.24414 | 0.42436 | 0.18022 | 73.8176 | 0.20331 | 0.040826 | 16.7223
|  | CAA |  |  | 1 | 50 | 72.432 |  |  | 54.888
[6 empty] | Mean |  |  | 55.3578 |  |  | 9.6304
[6 empty] | Sum |  | 396.4985 |  |  | 346.0268
Conclusion: Power regression is best over a wide range

## Arginine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A1 | 4228622 |  | 87048106 | [13]C6H14N4O2 | 13874.19 | Rsquared 0.16 < R2Threshold 0.99
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A2 | 4771081 |  | 96102869 | [13]C6H14N4O2 | 14616.10 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 91.64% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A3 | 5146051 |  | 98527988 | [13]C6H14N4O2 | 8914.59 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 211.434% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A4 | 4464186 |  | 98240208 | [13]C6H14N4O2 | 4120.74 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 605.258% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A5 | 4934351 |  | 100909218 | [13]C6H14N4O2 | 2390.69 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 677.575% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A6 | 4965839 |  | 109240733 | [13]C6H14N4O2 | 1993.30 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -1333.204% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A7 | 5182224 |  | 111355824 | [13]C6H14N4O2 | 2197.25 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -3340.745% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A8 | 5265320 |  | 110286590 | [13]C6H14N4O2 | 2310.21 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -10594.81% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A9 | 5402950 |  | 111505799 | [13]C6H14N4O2 | 2852.29 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -20943.826% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A10 | 5246862 |  | 110991361 | [13]C6H14N4O2 | 3754.61 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -54966.179% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A11 | 4897374 |  | 105956208 | [13]C6H14N4O2 | 2984.34 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 33359.026% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_A12 | 4416624 |  | 98668963 | [13]C6H14N4O2 | 1519.05 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 137417.176% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B1 | 4093301 |  | 86802005 | [13]C6H14N4O2 | 3776.87 | Rsquared 0.16 < R2Threshold 0.99
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B2 | 3821094 |  | 94293376 | [13]C6H14N4O2 | 3669.11 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 26.432% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B3 | 4554387 |  | 101137637 | [13]C6H14N4O2 | 2986.27 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 149.706% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B4 | 4727766 |  | 95920967 | [13]C6H14N4O2 | 2742.75 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 161.642% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B5 | 5034157 |  | 104268992 | [13]C6H14N4O2 | 1877.63 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 315.505% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B6 | 5102267 |  | 109746911 | [13]C6H14N4O2 | 3868.73 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -524.059% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B7 | 4903178 |  | 105516256 | [13]C6H14N4O2 | 2805.24 | Rsquared 0.16 < R2Threshold 0.99
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B8 | 4521908 |  | 100484732 | [13]C6H14N4O2 | 1572.83 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 949.96% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B9 | 4616961 |  | 102863344 | [13]C6H14N4O2 | 2821.37 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -6347.987% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B10 | 5117018 |  | 110922923 | [13]C6H14N4O2 | 5533.87 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 7791.475% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B11 | 4826793 |  | 103657988 | [13]C6H14N4O2 | 2232.10 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 100364.243% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_B12 | 4590761 |  | 99836744 | [13]C6H14N4O2 | 1876.42 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 172615.957% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H14N4O2 | N/A
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C1 | 4440471 |  | 89295553 | [13]C6H14N4O2 | 6833.99 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 65.291% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C2 | 4470690 |  | 90072123 | [13]C6H14N4O2 | 5109.69 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -193.499% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C3 | 4426245 |  | 95940282 | [13]C6H14N4O2 | 8873.46 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -454.17% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C4 | 4960921 |  | 101034742 | [13]C6H14N4O2 | 7108.69 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -963.841% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C5 | 4779654 |  | 103516134 | [13]C6H14N4O2 | 3037.89 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -2308.795% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C6 | 4457066 |  | 96985952 | [13]C6H14N4O2 | 2379.59 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff -3019.076% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C7 | 4012346 |  | 93121918 | [13]C6H14N4O2 | 3957.59 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 3136.943% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C8 | 4969326 |  | 103661835 | [13]C6H14N4O2 | 2195.84 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 8177.35% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C9 | 4837682 |  | 104588072 | [13]C6H14N4O2 | 1668.22 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 9619.952% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C10 | 4577682 |  | 103971775 | [13]C6H14N4O2 | 2441.16 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 29043.056% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C11 | 4739511 |  | 105751410 | [13]C6H14N4O2 | 1524.26 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 55568.508% (absolute) > max calibration amount diff 20%
Arginine U-13C neg | 14.24 | Target Compound | KD100520_100520_C12 | 4627166 |  | 94072972 | [13]C6H14N4O2 | 2161.62 | Rsquared 0.16 < R2Threshold 0.99. Calibration Amount Diff 133751.721% (absolute) > max calibration amount diff 20%

## Arginine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A1 | 46755860 |  | 1176600981 | [13]C6H14N4O2 | 18991.42 | Rsquared 0.33 < R2Threshold 0.99
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A2 | 57456159 |  | 1501873341 | [13]C6H14N4O2 | 7129.59 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 77.451% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A3 | 67417759 |  | 1599653519 | [13]C6H14N4O2 | 15147.08 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 199.206% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A4 | 72369279 |  | 1782572743 | [13]C6H14N4O2 | 10417.54 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 471.288% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A5 | 75551776 |  | 1844498881 | [13]C6H14N4O2 | 23080.75 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 758.956% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A6 | 73372001 |  | 1977632290 | [13]C6H14N4O2 | 4445.46 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -1421.523% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A7 | 74411205 |  | 2028502040 | [13]C6H14N4O2 | 13090.80 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -2307.665% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A8 | 78671411 |  | 2032794953 | [13]C6H14N4O2 | 6608.39 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -5670.183% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A9 | 80133244 |  | 1972661069 | [13]C6H14N4O2 | 7994.64 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -8626.833% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A10 | 73473014 |  | 2057522012 | [13]C6H14N4O2 | 4820.44 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -30796.234% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A11 | 71388294 |  | 1911275888 | [13]C6H14N4O2 | 28305.73 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 24990.399% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_A12 | 69715057 |  | 1949360767 | [13]C6H14N4O2 | 10273.81 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 77262.598% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B1 | 46313574 |  | 1158541646 | [13]C6H14N4O2 | 13810.65 | Rsquared 0.33 < R2Threshold 0.99
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B2 | 58862681 |  | 1507349314 | [13]C6H14N4O2 | 8509.71 | Rsquared 0.33 < R2Threshold 0.99
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B3 | 65531179 |  | 1661853101 | [13]C6H14N4O2 | 27431.04 | Rsquared 0.33 < R2Threshold 0.99
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B4 | 79274845 |  | 1848927504 | [13]C6H14N4O2 | 17402.25 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 65.945% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B5 | 78770024 |  | 1978189977 | [13]C6H14N4O2 | 16767.66 | Rsquared 0.33 < R2Threshold 0.99
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B6 | 80478564 |  | 2056059856 | [13]C6H14N4O2 | 7329.05 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -555.779% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B7 | 76495792 |  | 2025081425 | [13]C6H14N4O2 | 8981.36 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -795.423% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B8 | 78124561 |  | 2012289652 | [13]C6H14N4O2 | 19250.54 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -2725.336% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B9 | 79044114 |  | 2032608521 | [13]C6H14N4O2 | 8295.90 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -1751.712% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B10 | 78980902 |  | 2078447999 | [13]C6H14N4O2 | 8228.50 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -5923.386% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B11 | 70113627 |  | 1968513321 | [13]C6H14N4O2 | 10263.27 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 43605.201% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_B12 | 73805829 |  | 1971458645 | [13]C6H14N4O2 | 10795.45 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 111822.726% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_blank | 572 |  | 2166 | [13]C6H14N4O2 | INF
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C1 | 47216943 |  | 1240628822 | [13]C6H14N4O2 | 12511.66 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 36.272% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C2 | 54626778 |  | 1451842941 | [13]C6H14N4O2 | 12445.70 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -152.426% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C3 | 64655598 |  | 1621894788 | [13]C6H14N4O2 | 15306.44 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -286.971% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C4 | 71343948 |  | 1880384914 | [13]C6H14N4O2 | 18037.59 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -509.741% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C5 | 75895166 |  | 2004036479 | [13]C6H14N4O2 | 14042.29 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -825.375% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C6 | 76531191 |  | 1947391674 | [13]C6H14N4O2 | 6387.69 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff -2148.178% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C7 | 79486717 |  | 1953768861 | [13]C6H14N4O2 | 14410.64 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 1810.965% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C8 | 74507381 |  | 1970584349 | [13]C6H14N4O2 | 7093.75 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 4399.794% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C9 | 78082628 |  | 2024642459 | [13]C6H14N4O2 | 6200.71 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 11628.608% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C10 | 75416748 |  | 2037165515 | [13]C6H14N4O2 | 4738.19 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 23092.579% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C11 | 72746188 |  | 1947007107 | [13]C6H14N4O2 | 7069.94 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 38731.77% (absolute) > max calibration amount diff 20%
Arginine U-13C pos | 14.24 | Target Compound | KD100520_100520_C12 | 72968658 |  | 1882290414 | [13]C6H14N4O2 | 6241.84 | Rsquared 0.33 < R2Threshold 0.99. Calibration Amount Diff 86256.723% (absolute) > max calibration amount diff 20%

## Arginine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A1 | 7750529 |  | 165564957 | [13]C6H14[15]N4O2 | 280178.55
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A2 | 8738111 |  | 183028085 | [13]C6H14[15]N4O2 | 85638.94
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A3 | 9511137 |  | 186433367 | [13]C6H14[15]N4O2 | 10256.89
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A4 | 8582437 |  | 186580587 | [13]C6H14[15]N4O2 | 5395.87
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A5 | 8645918 |  | 188948777 | [13]C6H14[15]N4O2 | 5534.11
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A6 | 8820128 |  | 198390773 | [13]C6H14[15]N4O2 | 3259.72
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A7 | 9064586 |  | 201389834 | [13]C6H14[15]N4O2 | 8234.65
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A8 | 8998321 |  | 197327819 | [13]C6H14[15]N4O2 | 6410.14
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A9 | 9713647 |  | 199547945 | [13]C6H14[15]N4O2 | 3606.39
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A10 | 9541179 |  | 196898201 | [13]C6H14[15]N4O2 | 9509.72
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A11 | 8882238 |  | 197278642 | [13]C6H14[15]N4O2 | 5910.21
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_A12 | 8154463 |  | 185980359 | [13]C6H14[15]N4O2 | 2761.44
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B1 | 7694986 |  | 164394041 | [13]C6H14[15]N4O2 | 21542.61
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B2 | 7872962 |  | 177485109 | [13]C6H14[15]N4O2 | 24084.46
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B3 | 8344810 |  | 190314018 | [13]C6H14[15]N4O2 | 14419.10
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B4 | 8960242 |  | 178597259 | [13]C6H14[15]N4O2 | 8812.44
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B5 | 9190717 |  | 193678458 | [13]C6H14[15]N4O2 | 9319.39
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B6 | 9521018 |  | 201066717 | [13]C6H14[15]N4O2 | 12507.21
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B7 | 8869185 |  | 194329761 | [13]C6H14[15]N4O2 | 3793.36
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B8 | 8521770 |  | 185484624 | [13]C6H14[15]N4O2 | 4374.90
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B9 | 8580359 |  | 187725588 | [13]C6H14[15]N4O2 | 8690.81
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B10 | 9760869 |  | 205272791 | [13]C6H14[15]N4O2 | 21024.04
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B11 | 8635424 |  | 197576676 | [13]C6H14[15]N4O2 | 4163.21
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_B12 | 8682539 |  | 189347863 | [13]C6H14[15]N4O2 | 7018.97
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H14[15]N4O2 | N/A
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C1 | 8720022 |  | 174455034 | [13]C6H14[15]N4O2 | 32465.90
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C2 | 7607235 |  | 163116052 | [13]C6H14[15]N4O2 | 23332.67
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C3 | 7870221 |  | 171276698 | [13]C6H14[15]N4O2 | 20876.92
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C4 | 8733982 |  | 179188530 | [13]C6H14[15]N4O2 | 5863.47
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C5 | 8335013 |  | 181745331 | [13]C6H14[15]N4O2 | 3845.13
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C6 | 8004655 |  | 172986092 | [13]C6H14[15]N4O2 | 5905.18
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C7 | 7891367 |  | 174490552 | [13]C6H14[15]N4O2 | 15028.86
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C8 | 8976405 |  | 195220164 | [13]C6H14[15]N4O2 | 2760.49
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C9 | 8548354 |  | 195098312 | [13]C6H14[15]N4O2 | 9276.15
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C10 | 8422023 |  | 195263945 | [13]C6H14[15]N4O2 | 8114.74
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C11 | 9091107 |  | 198420923 | [13]C6H14[15]N4O2 | 7187.16
Arginine U-13C, U-15N neg | 14.24 | Internal Standard | KD100520_100520_C12 | 8408010 |  | 177203742 | [13]C6H14[15]N4O2 | 2980.34

## Arginine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A1 | 84667123 |  | 2180623550 | [13]C6H14[15]N4O2 | 42091.92
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A2 | 107218639 |  | 2750276091 | [13]C6H14[15]N4O2 | 25334.03
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A3 | 124341785 |  | 2908277164 | [13]C6H14[15]N4O2 | 11852.75
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A4 | 130664239 |  | 3235173762 | [13]C6H14[15]N4O2 | 23746.02
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A5 | 135902756 |  | 3317258941 | [13]C6H14[15]N4O2 | 14898.23
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A6 | 124744070 |  | 3392154190 | [13]C6H14[15]N4O2 | 11519.59
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A7 | 133649590 |  | 3490972249 | [13]C6H14[15]N4O2 | 11479.53
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A8 | 132384051 |  | 3483012107 | [13]C6H14[15]N4O2 | 11822.24
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A9 | 136662623 |  | 3396849427 | [13]C6H14[15]N4O2 | 9337.29
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A10 | 128611997 |  | 3497361113 | [13]C6H14[15]N4O2 | 10528.21
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A11 | 130001811 |  | 3386855996 | [13]C6H14[15]N4O2 | 20182.55
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_A12 | 128128256 |  | 3477555078 | [13]C6H14[15]N4O2 | 12147.20
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B1 | 81341478 |  | 2116300841 | [13]C6H14[15]N4O2 | 12712.72
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B2 | 106090886 |  | 2698253696 | [13]C6H14[15]N4O2 | 20947.16
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B3 | 116996513 |  | 2942409085 | [13]C6H14[15]N4O2 | 13111.40
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B4 | 139166775 |  | 3270209731 | [13]C6H14[15]N4O2 | 27623.44
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B5 | 133735973 |  | 3471639106 | [13]C6H14[15]N4O2 | 33514.37
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B6 | 139110502 |  | 3573769956 | [13]C6H14[15]N4O2 | 6492.94
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B7 | 133000583 |  | 3525800033 | [13]C6H14[15]N4O2 | 12239.68
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B8 | 135034185 |  | 3486906301 | [13]C6H14[15]N4O2 | 59177.56
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B9 | 134366890 |  | 3546646710 | [13]C6H14[15]N4O2 | 7990.19
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B10 | 139254402 |  | 3617798622 | [13]C6H14[15]N4O2 | 6889.42
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B11 | 129335524 |  | 3520466502 | [13]C6H14[15]N4O2 | 12224.36
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_B12 | 134443176 |  | 3547282220 | [13]C6H14[15]N4O2 | 24252.41
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H14[15]N4O2 | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C1 | 87278367 |  | 2329471061 | [13]C6H14[15]N4O2 | 13893.70
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C2 | 92977620 |  | 2508720143 | [13]C6H14[15]N4O2 | 18587.63
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C3 | 107918593 |  | 2774612942 | [13]C6H14[15]N4O2 | 19152.85
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C4 | 121388655 |  | 3209840787 | [13]C6H14[15]N4O2 | 35746.52
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C5 | 127759852 |  | 3430696373 | [13]C6H14[15]N4O2 | 10815.06
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C6 | 124208690 |  | 3303743825 | [13]C6H14[15]N4O2 | 16766.66
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C7 | 133853483 |  | 3471540009 | [13]C6H14[15]N4O2 | 12258.36
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C8 | 131985057 |  | 3510791889 | [13]C6H14[15]N4O2 | 17829.47
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C9 | 135206620 |  | 3626659020 | [13]C6H14[15]N4O2 | 10787.38
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C10 | 137572059 |  | 3648175047 | [13]C6H14[15]N4O2 | 9910.82
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C11 | 135285218 |  | 3473614690 | [13]C6H14[15]N4O2 | 16615.86
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD100520_100520_C12 | 135443693 |  | 3365387677 | [13]C6H14[15]N4O2 | 5787.90

## Asparagine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A1 | 48406695 |  | 724890830 | C4H8N2O3 | 16777.67 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -27.229% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A2 | 28749913 |  | 429717232 | C4H8N2O3 | 12130.11 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -30.579% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A3 | 23288663 |  | 326326944 | C4H8N2O3 | 21255.28 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A4 | 13253611 |  | 189634259 | C4H8N2O3 | 11741.79 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A5 | 9824477 |  | 138804962 | C4H8N2O3 | 7947.74 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A6 | 4714644 |  | 74291438 | C4H8N2O3 | 4482.45 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A7 | 2514918 |  | 37390630 | C4H8N2O3 | 2437.48 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -41.279% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A8 | 1478259 |  | 20585231 | C4H8N2O3 | 1870.23 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -74.977% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A9 | 771945 |  | 11434922 | C4H8N2O3 | 1103.15 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -133.332% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A10 | 308145 |  | 4612661 | C4H8N2O3 | 452.74 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -286.795% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A11 | 210775 |  | 2916253 | C4H8N2O3 | 357.63 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -550.057% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_A12 | 107206 |  | 1307270 | C4H8N2O3 | 133.32 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -1107.547% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B1 | 78098508 |  | 1154474761 | C4H8N2O3 | 353582.02 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff 32.51% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B2 | 47876017 |  | 705925906 | C4H8N2O3 | 21855.41 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff 21.11% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B3 | 34136581 |  | 489468925 | C4H8N2O3 | 34544.07 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff 34.433% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B4 | 20574904 |  | 314698560 | C4H8N2O3 | 13925.63 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff 39.492% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B5 | 12465656 |  | 181291445 | C4H8N2O3 | 10925.89 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff 26.711% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B6 | 6878205 |  | 102055059 | C4H8N2O3 | 8368.33 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff 32.011% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B7 | 4068060 |  | 57041289 | C4H8N2O3 | 6133.75 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B8 | 2137206 |  | 33301251 | C4H8N2O3 | 2425.88 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B9 | 1159325 |  | 18102712 | C4H8N2O3 | 1746.68 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -70.3% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B10 | 560083 |  | 8114431 | C4H8N2O3 | 754.53 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -225.407% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B11 | 292771 |  | 4323767 | C4H8N2O3 | 320.89 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -486.976% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_B12 | 132842 |  | 1870595 | C4H8N2O3 | 192.18 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -1060.257% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C4H8N2O3 | N/A
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C1 | 54181054 |  | 812571843 | C4H8N2O3 | INF | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C2 | 35106056 |  | 546210746 | C4H8N2O3 | 32149.42 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C3 | 22113401 |  | 352939849 | C4H8N2O3 | 11462.10 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C4 | 14209808 |  | 214941081 | C4H8N2O3 | 8566.75 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C5 | 9458312 |  | 135187883 | C4H8N2O3 | 9003.42 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C6 | 4830173 |  | 69952507 | C4H8N2O3 | 4588.45 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C7 | 3141837 |  | 45676437 | C4H8N2O3 | 4200.87 | Rsquared 0.929 < R2Threshold 0.99
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C8 | 1545057 |  | 22929144 | C4H8N2O3 | 1624.36 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -52.081% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C9 | 892285 |  | 12743141 | C4H8N2O3 | 1289.58 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -115.183% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C10 | 364909 |  | 5422588 | C4H8N2O3 | 836.13 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -269.65% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C11 | 191442 |  | 2855601 | C4H8N2O3 | 657.88 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -546.654% (absolute) > max calibration amount diff 20%
Asparagine neg | 9.77 | Target Compound | KD100520_100520_C12 | 77579 |  | 1230533 | C4H8N2O3 | 114.41 | Rsquared 0.929 < R2Threshold 0.99. Calibration Amount Diff -1098.328% (absolute) > max calibration amount diff 20%

## Asn_neg_calibration
Compound | Filename | Sample ID | Area | Area_IS | Response ratio | Conc. (Asn/Gln in cali B is double, hence correction) | Conc. Lin reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss | Conc. power reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss
Asparagine neg | KD100520_100520_A1 | A1 | 724890830 | 110598825 | 6.5542 | 500 | 460.9791 | 39.0209 | 7.8042 | 441.4636 | 58.5364 | 11.7073
Asparagine neg | KD100520_100520_A2 | A2 | 429717232 | 136364337 | 3.1512 | 250 | 221.6364 | 28.3636 | 11.3454 | 209.1061 | 40.8939 | 16.3576
Asparagine neg | KD100520_100520_A3 | A3 | 326326944 | 170484520 | 1.9141 | 125 | 134.6254 | 9.6254 | 7.7003 | 125.7291 | 0.72911 | 0.58329
Asparagine neg | KD100520_100520_A4 | A4 | 189634259 | 198636820 | 0.95468 | 62.5 | 67.1454 | 4.6454 | 7.4326 | 61.8247 | 0.67532 | 1.0805
Asparagine neg | KD100520_100520_A5 | A5 | 138804962 | 221660126 | 0.62621 | 31.25 | 44.043 | 12.793 | 40.9375 | 40.2056 | 8.9556 | 28.6578
Asparagine neg | KD100520_100520_A6 | A6 | 74291438 | 257182831 | 0.28887 | 15.625 | 20.3168 | 4.6918 | 30.0277 | 18.2562 | 2.6312 | 16.8398
Asparagine neg | KD100520_100520_A7 | A7 | 37390630 | 287922240 | 0.12986 | 7.8125 | 9.1337 | 1.3212 | 16.9114 | 8.0746 | 0.26205 | 3.3543
Asparagine neg | KD100520_100520_A8 | A8 | 20585231 | 315235089 | 0.065301 | 3.9062 | 4.5928 | 0.68658 | 17.5764 | 4.0037 | 0.097448 | 2.4947
Asparagine neg | KD100520_100520_A9 | A9 | 11434922 | 316027535 | 0.036183 | 1.9531 | 2.5449 | 0.59176 | 30.2979 | 2.1919 | 0.23876 | 12.2245
Asparagine neg | KD100520_100520_A10 | A10 | 4612661 | 303671023 | 0.01519 | 0.97656 | 1.0683 | 0.091772 | 9.3975 | 0.904 | 0.072564 | 7.4306
Asparagine neg | KD100520_100520_A11 | A11 | 2916253 | 341069571 | 0.0085503 | 0.48828 | 0.60137 | 0.11309 | 23.1605 | 0.50293 | 0.014652 | 3.0007
Asparagine neg | KD100520_100520_A12 | A12 | 1307270 | 338421953 | 0.0038628 | 0.24414 | 0.27169 | 0.027545 | 11.2822 | 0.22356 | 0.02058 | 8.4296
Asparagine neg | KD100520_100520_B1 | B1 | 1154474761 | 97050995 | 11.8955 | 1000 | 836.6496 | 163.3504 | 16.335 | 811.0321 | 188.9679 | 18.8968
Asparagine neg | KD100520_100520_B2 | B2 | 705925906 | 129242539 | 5.462 | 500 | 384.1606 | 115.8394 | 23.1679 | 366.5316 | 133.4684 | 26.6937
Asparagine neg | KD100520_100520_B3 | B3 | 489468925 | 160338517 | 3.0527 | 250 | 214.7071 | 35.2929 | 14.1172 | 202.4373 | 47.5627 | 19.0251
Asparagine neg | KD100520_100520_B4 | B4 | 314698560 | 195852051 | 1.6068 | 125 | 113.0123 | 11.9877 | 9.5901 | 105.1681 | 19.8319 | 15.8655
Asparagine neg | KD100520_100520_B5 | B5 | 181291445 | 239834523 | 0.7559 | 62.5 | 53.1649 | 9.3351 | 14.9362 | 48.7194 | 13.7806 | 22.0489
Asparagine neg | KD100520_100520_B6 | B6 | 102055059 | 244929703 | 0.41667 | 31.25 | 29.3057 | 1.9443 | 6.2217 | 26.5309 | 4.7191 | 15.101
Asparagine neg | KD100520_100520_B7 | B7 | 57041289 | 275385618 | 0.20713 | 15.625 | 14.5682 | 1.0568 | 6.7632 | 13.0022 | 2.6228 | 16.7862
Asparagine neg | KD100520_100520_B8 | B8 | 33301251 | 296992372 | 0.11213 | 7.8125 | 7.8863 | 0.07382 | 0.9449 | 6.951 | 0.86153 | 11.0276
Asparagine neg | KD100520_100520_B9 | B9 | 18102712 | 311064359 | 0.058196 | 3.9062 | 4.0931 | 0.18685 | 4.7834 | 3.5597 | 0.34655 | 8.8718
Asparagine neg | KD100520_100520_B10 | B10 | 8114431 | 313122792 | 0.025915 | 1.9531 | 1.8226 | 0.13048 | 6.6805 | 1.5592 | 0.39395 | 20.1701
Asparagine neg | KD100520_100520_B11 | B11 | 4323767 | 307631148 | 0.014055 | 0.97656 | 0.98853 | 0.01197 | 1.2258 | 0.83515 | 0.14141 | 14.4808
Asparagine neg | KD100520_100520_B12 | B12 | 1870595 | 315649031 | 0.0059262 | 0.48828 | 0.41681 | 0.071475 | 14.638 | 0.34598 | 0.1423 | 29.1425
Asparagine neg | KD100520_100520_C1 | C1 | 812571843 | 95574289 | 8.502 | 500 | 597.9706 | 97.9706 | 19.5941 | 575.7032 | 75.7032 | 15.1406
Asparagine neg | KD100520_100520_C2 | C2 | 546210746 | 118013263 | 4.6284 | 250 | 325.5282 | 75.5282 | 30.2113 | 309.5422 | 59.5422 | 23.8169
Asparagine neg | KD100520_100520_C3 | C3 | 352939849 | 147914067 | 2.3861 | 125 | 167.8226 | 42.8226 | 34.2581 | 157.4388 | 32.4388 | 25.9511
Asparagine neg | KD100520_100520_C4 | C4 | 214941081 | 178081086 | 1.207 | 62.5 | 84.8908 | 22.3908 | 35.8253 | 78.5388 | 16.0388 | 25.662
Asparagine neg | KD100520_100520_C5 | C5 | 135187883 | 209766920 | 0.64447 | 31.25 | 45.3273 | 14.0773 | 45.0474 | 41.4023 | 10.1523 | 32.4873
Asparagine neg | KD100520_100520_C6 | C6 | 69952507 | 230147668 | 0.30395 | 15.625 | 21.3774 | 5.7524 | 36.8157 | 19.2292 | 3.6042 | 23.067
Asparagine neg | KD100520_100520_C7 | C7 | 45676437 | 275250957 | 0.16594 | 7.8125 | 11.6714 | 3.8589 | 49.3938 | 10.3697 | 2.5572 | 32.7322
Asparagine neg | KD100520_100520_C8 | C8 | 22929144 | 282052749 | 0.081294 | 3.9062 | 5.7176 | 1.8114 | 46.3715 | 5.0065 | 1.1003 | 28.1676
Asparagine neg | KD100520_100520_C9 | C9 | 12743141 | 299688531 | 0.042521 | 1.9531 | 2.9906 | 1.0375 | 53.1213 | 2.5843 | 0.63119 | 32.317
Asparagine neg | KD100520_100520_C10 | C10 | 5422588 | 298189689 | 0.018185 | 0.97656 | 1.279 | 0.30245 | 30.9704 | 1.0862 | 0.10968 | 11.2315
Asparagine neg | KD100520_100520_C11 | C11 | 2855601 | 322764807 | 0.0088473 | 0.48828 | 0.62226 | 0.13398 | 27.4384 | 0.52077 | 0.032484 | 6.6527
Asparagine neg | KD100520_100520_C12 | C12 | 1230533 | 288511613 | 0.0042651 | 0.24414 | 0.29998 | 0.055837 | 22.8709 | 0.24734 | 0.0032005 | 1.3109
|  | CAA |  |  | 1 | 50 | 70.333 |  |  | 64.821
[6 empty] | Mean |  |  | 21.2554 |  |  | 16.3558
[6 empty] | Sum |  | 706.9952 |  |  | 727.8803
Conclusion: Power regression is best over a wide range

## Asparagine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A1 | 109197700 |  | 1430516182 | C4H8N2O3 | INF | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -24.47% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A2 | 57252574 |  | 835556216 | C4H8N2O3 | INF | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -33.514% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A3 | 42866690 |  | 604238687 | C4H8N2O3 | INF | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -20.514% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A4 | 23945654 |  | 337295213 | C4H8N2O3 | 6230.51 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -23.204% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A5 | 15685569 |  | 222024110 | C4H8N2O3 | 3882.61 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A6 | 7974025 |  | 112427515 | C4H8N2O3 | 3093.08 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A7 | 3374328 |  | 47114109 | C4H8N2O3 | 1139.69 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -27.264% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A8 | 1718486 |  | 23796751 | C4H8N2O3 | 726.02 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -31.853% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A9 | 689465 |  | 10769472 | C4H8N2O3 | 250.38 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -31.689% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A10 | 305864 |  | 4021537 | C4H8N2O3 | 197.15 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -33.379% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A11 | 175897 |  | 2061431 | C4H8N2O3 | 118.71 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_A12 | 61258 |  | 1045699 | C4H8N2O3 | 56.85 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 39.156% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B1 | 164740511 |  | 2349167352 | C4H8N2O3 | 21984.35 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 34.586% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B2 | 102847565 |  | 1435372336 | C4H8N2O3 | INF | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 20.092% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B3 | 64786265 |  | 917467496 | C4H8N2O3 | 9958.78 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 27.796% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B4 | 36813976 |  | 548674522 | C4H8N2O3 | 38277.44 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 32.196% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B5 | 18907383 |  | 282217063 | C4H8N2O3 | 6863.26 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B6 | 11101976 |  | 168902988 | C4H8N2O3 | 3264.41 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 29.439% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B7 | 5306071 |  | 81623770 | C4H8N2O3 | 2126.69 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 24.078% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B8 | 2943736 |  | 42185314 | C4H8N2O3 | 1189.93 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 24.21% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B9 | 1318214 |  | 19290409 | C4H8N2O3 | 618.23 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B10 | 518575 |  | 8293404 | C4H8N2O3 | 285.37 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B11 | 252974 |  | 3171031 | C4H8N2O3 | 165.05 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_B12 | 104777 |  | 1679388 | C4H8N2O3 | 72.79 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 67.394% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C4H8N2O3 | N/A
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C1 | 155698202 |  | 1991797620 | C4H8N2O3 | 772569.88 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C2 | 85093741 |  | 1243439137 | C4H8N2O3 | 18998.45 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C3 | 52491605 |  | 756019388 | C4H8N2O3 | 14807.87 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C4 | 29227327 |  | 427460959 | C4H8N2O3 | 5748.89 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C5 | 17026260 |  | 248183231 | C4H8N2O3 | 5025.11 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C6 | 8142410 |  | 117725272 | C4H8N2O3 | 2393.10 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C7 | 4473206 |  | 65867692 | C4H8N2O3 | 1823.06 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C8 | 1970939 |  | 28884984 | C4H8N2O3 | 739.93 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C9 | 1030054 |  | 13560794 | C4H8N2O3 | 365.52 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -20.27% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C10 | 311367 |  | 4396412 | C4H8N2O3 | 186.31 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff -29.538% (absolute) > max calibration amount diff 20%
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C11 | 155488 |  | 2313422 | C4H8N2O3 | 139.61 | Rsquared 0.928 < R2Threshold 0.99
Asparagine pos | 9.77 | Target Compound | KD100520_100520_C12 | 88254 |  | 1443233 | C4H8N2O3 | 53.28 | Rsquared 0.928 < R2Threshold 0.99. Calibration Amount Diff 56.446% (absolute) > max calibration amount diff 20%

## Asn_pos_calibration
Compound | Filename | Sample ID | Area | Area_IS | Response ratio | Conc. (Asn/Gln in cali B is double, hence correction) | Conc. Lin reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss | Conc. power reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss
Asparagine pos | KD100520_100520_A1 | A1 | 1430516182 | 196041843 | 7.297 | 434.3497 | 445.8317 | 11.4821 | 2.6435 | 371.9609 | 62.3888 | 14.3637
Asparagine pos | KD100520_100520_A2 | A2 | 835556216 | 260371485 | 3.2091 | 217.1748 | 196.0691 | 21.1057 | 9.7183 | 175.0238 | 42.1511 | 19.4088
Asparagine pos | KD100520_100520_A3 | A3 | 604238687 | 315287579 | 1.9165 | 108.5874 | 117.0924 | 8.505 | 7.8324 | 109.0541 | 0.46664 | 0.42973
Asparagine pos | KD100520_100520_A4 | A4 | 337295213 | 365246345 | 0.92347 | 54.2937 | 56.4224 | 2.1287 | 3.9206 | 55.8033 | 1.5096 | 2.7804
Asparagine pos | KD100520_100520_A5 | A5 | 222024110 | 374658178 | 0.5926 | 27.1469 | 36.2069 | 9.0601 | 33.3744 | 37.1412 | 9.9944 | 36.816
Asparagine pos | KD100520_100520_A6 | A6 | 112427515 | 429165852 | 0.26197 | 13.5734 | 16.0057 | 2.4323 | 17.9193 | 17.5596 | 3.9862 | 29.3678
Asparagine pos | KD100520_100520_A7 | A7 | 47114109 | 447179009 | 0.10536 | 6.7867 | 6.4372 | 0.34952 | 5.1501 | 7.6119 | 0.8252 | 12.1591
Asparagine pos | KD100520_100520_A8 | A8 | 23796751 | 506861349 | 0.046949 | 3.3934 | 2.8685 | 0.52485 | 15.4671 | 3.6253 | 0.23194 | 6.8351
Asparagine pos | KD100520_100520_A9 | A9 | 10769472 | 506095862 | 0.02128 | 1.6967 | 1.3001 | 0.39654 | 23.3717 | 1.7537 | 0.057042 | 3.362
Asparagine pos | KD100520_100520_A10 | A10 | 4021537 | 498341641 | 0.0080698 | 0.84834 | 0.49305 | 0.35529 | 41.8804 | 0.72031 | 0.12803 | 15.0916
Asparagine pos | KD100520_100520_A11 | A11 | 2061431 | 514311542 | 0.0040081 | 0.42417 | 0.24489 | 0.17928 | 42.2662 | 0.37897 | 0.045195 | 10.6549
Asparagine pos | KD100520_100520_A12 | A12 | 1045699 | 510023616 | 0.0020503 | 0.21208 | 0.12527 | 0.086816 | 40.9345 | 0.20485 | 0.007231 | 3.4095
Asparagine pos | KD100520_100520_B1 | B1 | 2349167352 | 180619541 | 13.0062 | 868.6994 | 794.6506 | 74.0488 | 8.5241 | 632.1859 | 236.5135 | 27.2262
Asparagine pos | KD100520_100520_B2 | B2 | 1435372336 | 247467439 | 5.8002 | 434.3497 | 354.3835 | 79.9662 | 18.4106 | 301.304 | 133.0457 | 30.631
Asparagine pos | KD100520_100520_B3 | B3 | 917467496 | 297487374 | 3.0841 | 217.1748 | 188.4296 | 28.7452 | 13.236 | 168.7553 | 48.4195 | 22.2952
Asparagine pos | KD100520_100520_B4 | B4 | 548674522 | 344442616 | 1.5929 | 108.5874 | 97.3251 | 11.2623 | 10.3717 | 92.0337 | 16.5537 | 15.2446
Asparagine pos | KD100520_100520_B5 | B5 | 282217063 | 394489754 | 0.7154 | 54.2937 | 43.7094 | 10.5843 | 19.4946 | 44.1477 | 10.146 | 18.6872
Asparagine pos | KD100520_100520_B6 | B6 | 168902988 | 436985381 | 0.38652 | 27.1469 | 23.6155 | 3.5313 | 13.0083 | 25.0921 | 2.0548 | 7.5692
Asparagine pos | KD100520_100520_B7 | B7 | 81623770 | 446236079 | 0.18292 | 13.5734 | 11.1758 | 2.3976 | 17.6641 | 12.6287 | 0.94473 | 6.9601
Asparagine pos | KD100520_100520_B8 | B8 | 42185314 | 472461478 | 0.089288 | 6.7867 | 5.4553 | 1.3314 | 19.6173 | 6.5393 | 0.24736 | 3.6448
Asparagine pos | KD100520_100520_B9 | B9 | 19290409 | 485621008 | 0.039723 | 3.3934 | 2.427 | 0.96635 | 28.4777 | 3.1098 | 0.28356 | 8.3562
Asparagine pos | KD100520_100520_B10 | B10 | 8293404 | 524612112 | 0.015809 | 1.6967 | 0.96588 | 0.7308 | 43.0725 | 1.3351 | 0.36157 | 21.3106
Asparagine pos | KD100520_100520_B11 | B11 | 3171031 | 498156258 | 0.0063655 | 0.84834 | 0.38892 | 0.45942 | 54.155 | 0.57939 | 0.26895 | 31.7033
Asparagine pos | KD100520_100520_B12 | B12 | 1679388 | 496501281 | 0.0033824 | 0.42417 | 0.20666 | 0.21751 | 51.2788 | 0.32431 | 0.099856 | 23.5416
Asparagine pos | KD100520_100520_C1 | C1 | 1991797620 | 222756136 | 8.9416 | 434.3497 | 546.3142 | 111.9646 | 25.7775 | 448.2333 | 13.8836 | 3.1964
Asparagine pos | KD100520_100520_C2 | C2 | 1243439137 | 257785639 | 4.8235 | 217.1748 | 294.7086 | 77.5338 | 35.7011 | 254.3987 | 37.2238 | 17.14
Asparagine pos | KD100520_100520_C3 | C3 | 756019388 | 309063488 | 2.4462 | 108.5874 | 149.4556 | 40.8682 | 37.6362 | 136.4278 | 27.8404 | 25.6387
Asparagine pos | KD100520_100520_C4 | C4 | 427460959 | 365227427 | 1.1704 | 54.2937 | 71.5089 | 17.2152 | 31.7075 | 69.3584 | 15.0647 | 27.7467
Asparagine pos | KD100520_100520_C5 | C5 | 248183231 | 403636009 | 0.61487 | 27.1469 | 37.5673 | 10.4204 | 38.3853 | 38.4199 | 11.273 | 41.526
Asparagine pos | KD100520_100520_C6 | C6 | 117725272 | 431389487 | 0.2729 | 13.5734 | 16.6735 | 3.1001 | 22.8394 | 18.2309 | 4.6574 | 34.3129
Asparagine pos | KD100520_100520_C7 | C7 | 65867692 | 460284716 | 0.1431 | 6.7867 | 8.7432 | 1.9565 | 28.8289 | 10.0815 | 3.2948 | 48.548
Asparagine pos | KD100520_100520_C8 | C8 | 28884984 | 474832522 | 0.060832 | 3.3934 | 3.7167 | 0.32335 | 9.529 | 4.5982 | 1.2048 | 35.5059
Asparagine pos | KD100520_100520_C9 | C9 | 13560794 | 529896983 | 0.025591 | 1.6967 | 1.5636 | 0.1331 | 7.8445 | 2.0773 | 0.38061 | 22.4329
Asparagine pos | KD100520_100520_C10 | C10 | 4396412 | 499857968 | 0.0087953 | 0.84834 | 0.53738 | 0.31096 | 36.6555 | 0.77952 | 0.068815 | 8.1117
Asparagine pos | KD100520_100520_C11 | C11 | 2313422 | 541445159 | 0.0042727 | 0.42417 | 0.26105 | 0.16312 | 38.4557 | 0.40187 | 0.022301 | 5.2576
Asparagine pos | KD100520_100520_C12 | C12 | 1443233 | 503574826 | 0.002866 | 0.21208 | 0.17511 | 0.036979 | 17.4362 | 0.27857 | 0.066482 | 31.3468
|  | CAA |  |  | 1 | 50 | 70.333 |  |  | 60.033
[6 empty] | Mean |  |  | 24.2385 |  |  | 18.6837
[6 empty] | Sum |  | 534.8735 |  |  | 685.7114
Conclusion: Power regression is best over a wide range

## Asn2Asp
Asn Response ratio | Asp Response ratio | Ratio
7.297 | 7.6536 | 0.95341
3.2091 | 3.626 | 0.88503
1.9165 | 2.2363 | 0.85697
0.92347 | 1.1414 | 0.8091
0.5926 | 0.75424 | 0.7857
0.26197 | 0.35182 | 0.74461
0.10536 | 0.16062 | 0.65595
0.046949 | 0.082791 | 0.56708
0.02128 | 0.045452 | 0.46817
0.0080698 | 0.020685 | 0.39014
0.0040081 | 0.012922 | 0.31018
0.0020503 | 0.0064773 | 0.31654
13.0062 | 8.266 | 1.5735
5.8002 | 3.9091 | 1.4838
3.0841 | 2.1947 | 1.4052
1.5929 | 1.1897 | 1.339
0.7154 | 0.56532 | 1.2655
0.38652 | 0.31444 | 1.2292
0.18292 | 0.15945 | 1.1472
0.089288 | 0.085575 | 1.0434
0.039723 | 0.04497 | 0.88333
0.015809 | 0.020971 | 0.75384
0.0063655 | 0.012089 | 0.52655
0.0033824 | 0.0058 | 0.58318
8.9416 | 7.4994 | 1.1923
4.8235 | 4.2204 | 1.1429
2.4462 | 2.2006 | 1.1116
1.1704 | 1.1471 | 1.0203
0.61487 | 0.61759 | 0.9956
0.2729 | 0.29634 | 0.92089
0.1431 | 0.16533 | 0.86558
0.060832 | 0.079071 | 0.76934
0.025591 | 0.041241 | 0.62053
0.0087953 | 0.018977 | 0.46347
0.0042727 | 0.010913 | 0.39151
0.002866 | 0.0055527 | 0.51614

## Asparagine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A1 | 205503 |  | 2688017 | [13]C4H8N2O3 | 5307.90 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -28.611% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A2 | 260541 |  | 3796738 | [13]C4H8N2O3 | 3191.25 | Rsquared 0.872 < R2Threshold 0.99
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A3 | 457554 |  | 5479936 | [13]C4H8N2O3 | 1919.79 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 57.559% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A4 | 548585 |  | 7009300 | [13]C4H8N2O3 | 1215.55 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 112.46% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A5 | 652713 |  | 8300384 | [13]C4H8N2O3 | 1137.48 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 183.888% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A6 | 727126 |  | 10993009 | [13]C4H8N2O3 | 3030.95 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -224.184% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A7 | 875050 |  | 12427444 | [13]C4H8N2O3 | 2117.45 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -457.688% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A8 | 1057499 |  | 14230065 | [13]C4H8N2O3 | 4201.68 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -1849.327% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A9 | 985749 |  | 14332322 | [13]C4H8N2O3 | 3937.98 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -3818.508% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A10 | 911140 |  | 13792008 | [13]C4H8N2O3 | 1836.48 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -7671.285% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A11 | 979547 |  | 14338964 | [13]C4H8N2O3 | 1823.19 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -1136.993% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_A12 | 1054293 |  | 14354453 | [13]C4H8N2O3 | 1828.59 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -5307.942% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B1 | 143338 |  | 1916941 | [13]C4H8N2O3 | 3915.18 | Rsquared 0.872 < R2Threshold 0.99
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B2 | 231646 |  | 3196043 | [13]C4H8N2O3 | INF | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 39.308% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B3 | 398637 |  | 4857587 | [13]C4H8N2O3 | 869.28 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 87.724% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B4 | 529555 |  | 7099411 | [13]C4H8N2O3 | 1710.13 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 81.052% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B5 | 710940 |  | 9559863 | [13]C4H8N2O3 | 1723.47 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 26.243% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B6 | 719810 |  | 9985763 | [13]C4H8N2O3 | 1933.45 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 33.665% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B7 | 869167 |  | 11850921 | [13]C4H8N2O3 | 2105.02 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -424.098% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B8 | 867123 |  | 12622424 | [13]C4H8N2O3 | 3534.66 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -469.738% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B9 | 897077 |  | 13762672 | [13]C4H8N2O3 | 2585.05 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -2660.936% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B10 | 981783 |  | 13782448 | [13]C4H8N2O3 | 1487.57 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -4743.599% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B11 | 901131 |  | 12987431 | [13]C4H8N2O3 | 1849.25 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -1874.615% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_B12 | 902218 |  | 13327210 | [13]C4H8N2O3 | 1696.95 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -3683.247% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C4H8N2O3 | N/A
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C1 | 128955 |  | 1691109 | [13]C4H8N2O3 | INF | Rsquared 0.872 < R2Threshold 0.99
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C2 | 220028 |  | 2984853 | [13]C4H8N2O3 | INF | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 34.708% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C3 | 351471 |  | 4561720 | [13]C4H8N2O3 | 3326.00 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 78.833% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C4 | 487911 |  | 6420503 | [13]C4H8N2O3 | 1242.71 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 87.42% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C5 | 649706 |  | 8522831 | [13]C4H8N2O3 | 1388.56 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -24.03% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C6 | 736326 |  | 9710368 | [13]C4H8N2O3 | 1678.70 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -152.075% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C7 | 810452 |  | 11290657 | [13]C4H8N2O3 | 1800.94 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 102.114% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C8 | 790085 |  | 11587047 | [13]C4H8N2O3 | 1784.43 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 272.076% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C9 | 901077 |  | 12887618 | [13]C4H8N2O3 | 2703.81 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -1364.645% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C10 | 843179 |  | 12644688 | [13]C4H8N2O3 | 2345.95 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -1377.68% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C11 | 953903 |  | 13471299 | [13]C4H8N2O3 | 2148.33 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff 134.298% (absolute) > max calibration amount diff 20%
Asparagine U-13C neg | 9.77 | Target Compound | KD100520_100520_C12 | 804665 |  | 12251761 | [13]C4H8N2O3 | 2028.18 | Rsquared 0.872 < R2Threshold 0.99. Calibration Amount Diff -5722.447% (absolute) > max calibration amount diff 20%

## Asparagine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A1 | 392062 |  | 4221287 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A2 | 516664 |  | 6636253 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 38.399% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A3 | 623688 |  | 8073644 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 170.79% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A4 | 796564 |  | 10039708 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 252.682% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A5 | 793997 |  | 9935069 | [13]C4H8N2O3 | 5870.04 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 800.228% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A6 | 1076139 |  | 13188293 | [13]C4H8N2O3 | 3671.53 | Rsquared 0.536 < R2Threshold 0.99
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A7 | 1134477 |  | 15067103 | [13]C4H8N2O3 | 746.77 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -2266.386% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A8 | 1308950 |  | 17842547 | [13]C4H8N2O3 | 1240.05 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -6858.055% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A9 | 1112851 |  | 16838567 | [13]C4H8N2O3 | 1144.25 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -7409.335% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A10 | 1327454 |  | 16900857 | [13]C4H8N2O3 | 1090.52 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -18841.72% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A11 | 1247272 |  | 16967308 | [13]C4H8N2O3 | 1104.11 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -25734.607% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_A12 | 1077500 |  | 16612827 | [13]C4H8N2O3 | 639.74 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -40621.048% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B1 | 418339 |  | 4636291 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -33.075% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B2 | 480732 |  | 7142567 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -46.367% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B3 | 615133 |  | 7906625 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 122.026% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B4 | 617513 |  | 8777646 | [13]C4H8N2O3 | 22467.99 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 454% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B5 | 880345 |  | 11141347 | [13]C4H8N2O3 | 12158.54 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 453.679% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B6 | 963149 |  | 13531039 | [13]C4H8N2O3 | 1161.43 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -86.53% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B7 | 943868 |  | 13087692 | [13]C4H8N2O3 | 773.37 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 1241.274% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B8 | 1093621 |  | 14277827 | [13]C4H8N2O3 | 695.32 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 1150.567% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B9 | 966447 |  | 15262314 | [13]C4H8N2O3 | 830.41 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -1483.813% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B10 | 1191965 |  | 16827874 | [13]C4H8N2O3 | 942.30 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -7033.029% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B11 | 1092092 |  | 14915368 | [13]C4H8N2O3 | 736.49 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 13498.786% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_B12 | 1055047 |  | 15836033 | [13]C4H8N2O3 | 817.70 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -23188.763% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C4H8N2O3 | N/A
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C1 | 561121 |  | 5376554 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C2 | 532037 |  | 7453149 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -47.612% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C3 | 636933 |  | 8490711 | [13]C4H8N2O3 | INF | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 77.1% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C4 | 713245 |  | 10112337 | [13]C4H8N2O3 | 4366.00 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 232.561% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C5 | 892268 |  | 11940745 | [13]C4H8N2O3 | 15266.78 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 184.329% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C6 | 1033527 |  | 14211813 | [13]C4H8N2O3 | 778.76 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -882.111% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C7 | 983383 |  | 13690180 | [13]C4H8N2O3 | 920.48 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 908.739% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C8 | 926031 |  | 13641284 | [13]C4H8N2O3 | 882.95 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 3548.143% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C9 | 1313126 |  | 17089875 | [13]C4H8N2O3 | 932.02 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -4129.392% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C10 | 1158081 |  | 16120062 | [13]C4H8N2O3 | 710.29 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -8141.605% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C11 | 1272682 |  | 17365673 | [13]C4H8N2O3 | 1280.81 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -13929.143% (absolute) > max calibration amount diff 20%
Asparagine U-13C pos | 9.77 | Target Compound | KD100520_100520_C12 | 1173826 |  | 16458653 | [13]C4H8N2O3 | 1031.27 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -43476.641% (absolute) > max calibration amount diff 20%

## Asparagine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A1 | 7849386 |  | 110598825 | [13]C4H8[15]N2O3 | 11558.00 | 9.797:Peak area 110598824.946 is out of bounds (ISTD Minimum Recovery 119865613.303 and ISTD Max Recovery 359596839.910)
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A2 | 9106748 |  | 136364337 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A3 | 12514992 |  | 170484520 | [13]C4H8[15]N2O3 | 17709.36
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A4 | 14205075 |  | 198636820 | [13]C4H8[15]N2O3 | 101065.86
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A5 | 15334618 |  | 221660126 | [13]C4H8[15]N2O3 | 56208.56
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A6 | 16797032 |  | 257182831 | [13]C4H8[15]N2O3 | 45287.29
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A7 | 19572253 |  | 287922240 | [13]C4H8[15]N2O3 | 13030.81
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A8 | 22810099 |  | 315235089 | [13]C4H8[15]N2O3 | 35854.14
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A9 | 21762689 |  | 316027535 | [13]C4H8[15]N2O3 | 23643.81
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A10 | 19439934 |  | 303671023 | [13]C4H8[15]N2O3 | 8269.92
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A11 | 22876742 |  | 341069571 | [13]C4H8[15]N2O3 | 17238.72
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_A12 | 24393689 |  | 338421953 | [13]C4H8[15]N2O3 | 10334.68
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B1 | 6623961 |  | 97050995 | [13]C4H8[15]N2O3 | 22159.26 | 9.783:Peak area 97050995.488 is out of bounds (ISTD Minimum Recovery 119865613.303 and ISTD Max Recovery 359596839.910)
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B2 | 9099611 |  | 129242539 | [13]C4H8[15]N2O3 | 70772.81
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B3 | 11516954 |  | 160338517 | [13]C4H8[15]N2O3 | 236601.38
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B4 | 13217767 |  | 195852051 | [13]C4H8[15]N2O3 | 12456.16
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B5 | 16783279 |  | 239834523 | [13]C4H8[15]N2O3 | 18367.18
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B6 | 16921585 |  | 244929703 | [13]C4H8[15]N2O3 | 26240.76
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B7 | 19633711 |  | 275385618 | [13]C4H8[15]N2O3 | 24407.76
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B8 | 19773415 |  | 296992372 | [13]C4H8[15]N2O3 | 10855.73
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B9 | 19938366 |  | 311064359 | [13]C4H8[15]N2O3 | 18809.21
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B10 | 21142402 |  | 313122792 | [13]C4H8[15]N2O3 | 16229.04
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B11 | 20161902 |  | 307631148 | [13]C4H8[15]N2O3 | 12936.17
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_B12 | 20625745 |  | 315649031 | [13]C4H8[15]N2O3 | 14528.07
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C4H8[15]N2O3 | N/A
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C1 | 6938899 |  | 95574289 | [13]C4H8[15]N2O3 | 12377.23 | 9.785:Peak area 95574288.970 is out of bounds (ISTD Minimum Recovery 119865613.303 and ISTD Max Recovery 359596839.910)
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C2 | 8030118 |  | 118013263 | [13]C4H8[15]N2O3 | 20699.12 | 9.767:Peak area 118013262.910 is out of bounds (ISTD Minimum Recovery 119865613.303 and ISTD Max Recovery 359596839.910)
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C3 | 9776040 |  | 147914067 | [13]C4H8[15]N2O3 | 11983.60
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C4 | 11624565 |  | 178081086 | [13]C4H8[15]N2O3 | 20628.82
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C5 | 14620396 |  | 209766920 | [13]C4H8[15]N2O3 | 19560.04
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C6 | 15991553 |  | 230147668 | [13]C4H8[15]N2O3 | 21448.95
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C7 | 19177402 |  | 275250957 | [13]C4H8[15]N2O3 | 32452.18
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C8 | 18975114 |  | 282052749 | [13]C4H8[15]N2O3 | 13926.57
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C9 | 20631693 |  | 299688531 | [13]C4H8[15]N2O3 | 22728.28
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C10 | 19384781 |  | 298189689 | [13]C4H8[15]N2O3 | 21564.59
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C11 | 21509888 |  | 322764807 | [13]C4H8[15]N2O3 | 34798.30
Asparagine U-13C, U-15N neg | 9.77 | Internal Standard | KD100520_100520_C12 | 18184174 |  | 288511613 | [13]C4H8[15]N2O3 | 11484.65

## Asparagine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A1 | 16305481 |  | 196041843 | [13]C4H8[15]N2O3 | INF | 9.755:Peak area 196041842.664 is out of bounds (ISTD Minimum Recovery 204839096.951 and ISTD Max Recovery 614517290.852)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A2 | 19028826 |  | 260371485 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A3 | 23457219 |  | 315287579 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A4 | 26953721 |  | 365246345 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A5 | 26864839 |  | 374658178 | [13]C4H8[15]N2O3 | 109100.02
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A6 | 30148717 |  | 429165852 | [13]C4H8[15]N2O3 | 726364.63
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A7 | 28459499 |  | 447179009 | [13]C4H8[15]N2O3 | 110062.85
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A8 | 32816885 |  | 506861349 | [13]C4H8[15]N2O3 | 27823.89
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A9 | 30661224 |  | 506095862 | [13]C4H8[15]N2O3 | 57293.01
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A10 | 32314118 |  | 498341641 | [13]C4H8[15]N2O3 | 3534317.58
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A11 | 34431110 |  | 514311542 | [13]C4H8[15]N2O3 | 32795.89
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_A12 | 29899639 |  | 510023616 | [13]C4H8[15]N2O3 | 3039234.44
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B1 | 13795203 |  | 180619541 | [13]C4H8[15]N2O3 | 10685.82 | 9.769:Peak area 180619541.276 is out of bounds (ISTD Minimum Recovery 204839096.951 and ISTD Max Recovery 614517290.852)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B2 | 19089520 |  | 247467439 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B3 | 21910652 |  | 297487374 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B4 | 23063611 |  | 344442616 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B5 | 25539762 |  | 394489754 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B6 | 27796129 |  | 436985381 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B7 | 29333107 |  | 446236079 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B8 | 31575359 |  | 472461478 | [13]C4H8[15]N2O3 | 2306449.04
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B9 | 28854261 |  | 485621008 | [13]C4H8[15]N2O3 | 113244.41
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B10 | 33549655 |  | 524612112 | [13]C4H8[15]N2O3 | 4866985.76
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B11 | 32314025 |  | 498156258 | [13]C4H8[15]N2O3 | 75856.33
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_B12 | 30476242 |  | 496501281 | [13]C4H8[15]N2O3 | 40690.98
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C4H8[15]N2O3 | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C1 | 18989349 |  | 222756136 | [13]C4H8[15]N2O3 | 74687.15
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C2 | 18048181 |  | 257785639 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C3 | 22043643 |  | 309063488 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C4 | 24872776 |  | 365227427 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C5 | 27535119 |  | 403636009 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C6 | 28676582 |  | 431389487 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C7 | 29174331 |  | 460284716 | [13]C4H8[15]N2O3 | 739568.29
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C8 | 29525845 |  | 474832522 | [13]C4H8[15]N2O3 | INF
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C9 | 35400015 |  | 529896983 | [13]C4H8[15]N2O3 | 38237.02
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C10 | 32810613 |  | 499857968 | [13]C4H8[15]N2O3 | 31967.08
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C11 | 35673818 |  | 541445159 | [13]C4H8[15]N2O3 | 611525.26
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD100520_100520_C12 | 32877859 |  | 503574826 | [13]C4H8[15]N2O3 | 3802917.31

## Aspartate neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A1 | 142090865 |  | 2147930837 | C4H7NO4 | 35448.77
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A2 | 86004036 |  | 1294415195 | C4H7NO4 | 33153.62
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A3 | 59958054 |  | 910411991 | C4H7NO4 | 27319.08
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A4 | 32088417 |  | 521857751 | C4H7NO4 | 20503.89
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A5 | 27946617 |  | 412161691 | C4H7NO4 | 18780.63 | Calibration Amount Diff 40.947% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A6 | 13142921 |  | 196523991 | C4H7NO4 | 5655.77
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A7 | 5480455 |  | 86124068 | C4H7NO4 | 3765.48 | Calibration Amount Diff -24.225% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A8 | 2758304 |  | 45984669 | C4H7NO4 | 2053.27 | Calibration Amount Diff -76.406% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A9 | 1586588 |  | 23901896 | C4H7NO4 | 1106.19 | Calibration Amount Diff -175.604% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A10 | 740761 |  | 11289211 | C4H7NO4 | 676.00 | Calibration Amount Diff -413.947% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A11 | 449357 |  | 6854127 | C4H7NO4 | 324.67 | Calibration Amount Diff -830.701% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_A12 | 214577 |  | 3655621 | C4H7NO4 | 158.59 | Calibration Amount Diff -1731.04% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B1 | 155251037 |  | 2351051145 | C4H7NO4 | 28606.50
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B2 | 84034280 |  | 1327840381 | C4H7NO4 | 24567.99
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B3 | 65222072 |  | 961041717 | C4H7NO4 | 23834.54
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B4 | 37864533 |  | 582448533 | C4H7NO4 | 15552.16
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B5 | 19765249 |  | 314915065 | C4H7NO4 | 12403.77
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B6 | 11292554 |  | 171280638 | C4H7NO4 | 5374.98
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B7 | 5867761 |  | 88581805 | C4H7NO4 | 5138.77 | Calibration Amount Diff -25.19% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B8 | 3266178 |  | 47199591 | C4H7NO4 | 2212.90 | Calibration Amount Diff -71.828% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B9 | 1640504 |  | 25661736 | C4H7NO4 | 1093.39 | Calibration Amount Diff -177.19% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B10 | 841219 |  | 11943980 | C4H7NO4 | 566.82 | Calibration Amount Diff -412.066% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B11 | 399943 |  | 6623166 | C4H7NO4 | 324.89 | Calibration Amount Diff -841.662% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_B12 | 176556 |  | 3057799 | C4H7NO4 | 148.84 | Calibration Amount Diff -1748.867% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C4H7NO4 | N/A
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C1 | 141769657 |  | 2138084572 | C4H7NO4 | 43417.24
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C2 | 99834508 |  | 1442565850 | C4H7NO4 | 40379.75
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C3 | 59285125 |  | 890883894 | C4H7NO4 | 22868.90
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C4 | 35770017 |  | 561035183 | C4H7NO4 | 15943.50
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C5 | 21726795 |  | 327798771 | C4H7NO4 | 15917.71
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C6 | 10397548 |  | 155286937 | C4H7NO4 | 5000.20
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C7 | 6568123 |  | 94766616 | C4H7NO4 | 4024.16 | Calibration Amount Diff -20.358% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C8 | 3081234 |  | 45680965 | C4H7NO4 | 1743.87 | Calibration Amount Diff -82.523% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C9 | 1618594 |  | 24645962 | C4H7NO4 | 1084.80 | Calibration Amount Diff -189.453% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C10 | 694371 |  | 10885397 | C4H7NO4 | 553.38 | Calibration Amount Diff -425.171% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C11 | 399134 |  | 6233618 | C4H7NO4 | 345.46 | Calibration Amount Diff -857.138% (absolute) > max calibration amount diff 20%
Aspartate neg | 10.26 | Target Compound | KD100520_100520_C12 | 192930 |  | 2910486 | C4H7NO4 | 172.25 | Calibration Amount Diff -1755.378% (absolute) > max calibration amount diff 20%

## Asp_neg_calibratio
Compound | Filename | Sample ID | Area | Area_IS | Area_IS2 | Response ratio | Response ratio 2 | Conc. | Conc. Lin reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss | Conc. power reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss | Conc. power reg. IS2 | Lin. Reg. l1 loss IS2 | Lin. Reg. c1 loss IS2
Aspartate neg | KD100520_100520_A1 | A1 | 2147930837 | 280643581 | 139365647 | 7.6536 | 15.4122 | 500 | 483.4007 | 16.5993 | 3.3199 | 470.146 | 29.854 | 5.9708 | 493.1715 | 6.8285 | 1.3657
Aspartate neg | KD100520_100520_A2 | A2 | 1294415195 | 356986169 | 180939922 | 3.626 | 7.1538 | 250 | 229.0152 | 20.9848 | 8.3939 | 213.3226 | 36.6774 | 14.6709 | 219.3342 | 30.6658 | 12.2663
Aspartate neg | KD100520_100520_A3 | A3 | 910411991 | 407101577 | 207620179 | 2.2363 | 4.385 | 125 | 141.2464 | 16.2464 | 12.9971 | 127.9436 | 2.9436 | 2.3549 | 130.8264 | 5.8264 | 4.6611
Aspartate neg | KD100520_100520_A4 | A4 | 521857751 | 457223332 | 234269110 | 1.1414 | 2.2276 | 62.5 | 72.0885 | 9.5885 | 15.3416 | 62.8092 | 0.30916 | 0.49465 | 64.0001 | 1.5001 | 2.4002
Aspartate neg | KD100520_100520_A5 | A5 | 412161691 | 546458563 | 279183435 | 0.75424 | 1.4763 | 31.25 | 47.6379 | 16.3879 | 52.4412 | 40.5238 | 9.2738 | 29.6763 | 41.4544 | 10.2044 | 32.654
Aspartate neg | KD100520_100520_A6 | A6 | 196523991 | 558592743 | 299509680 | 0.35182 | 0.65615 | 15.625 | 22.2209 | 6.5959 | 42.214 | 18.0875 | 2.4625 | 15.7598 | 17.6109 | 1.9859 | 12.7097
Aspartate neg | KD100520_100520_A7 | A7 | 86124068 | 536194956 | 289102108 | 0.16062 | 0.2979 | 7.8125 | 10.1448 | 2.3323 | 29.8536 | 7.8918 | 0.079322 | 1.0153 | 7.6515 | 0.16097 | 2.0604
Aspartate neg | KD100520_100520_A8 | A8 | 45984669 | 555430116 | 301279313 | 0.082791 | 0.15263 | 3.9062 | 5.2291 | 1.3228 | 33.8646 | 3.9149 | 0.008673 | 0.22203 | 3.777 | 0.12929 | 3.3099
Aspartate neg | KD100520_100520_A9 | A9 | 23901896 | 525869578 | 285650813 | 0.045452 | 0.083675 | 1.9531 | 2.8708 | 0.91763 | 46.9828 | 2.0761 | 0.12294 | 6.2946 | 2.0024 | 0.049293 | 2.5238
Aspartate neg | KD100520_100520_A10 | A10 | 11289211 | 545779050 | 294719737 | 0.020685 | 0.038305 | 0.97656 | 1.3064 | 0.32988 | 33.7793 | 0.90276 | 0.073805 | 7.5576 | 0.87763 | 0.098934 | 10.1308
Aspartate neg | KD100520_100520_A11 | A11 | 6854127 | 530421817 | 273647607 | 0.012922 | 0.025047 | 0.48828 | 0.81616 | 0.32787 | 67.1486 | 0.54884 | 0.060559 | 12.4024 | 0.56046 | 0.072174 | 14.7812
Aspartate neg | KD100520_100520_A12 | A12 | 3655621 | 564373570 | 289213058 | 0.0064773 | 0.01264 | 0.24414 | 0.40911 | 0.16497 | 67.5701 | 0.26435 | 0.020206 | 8.2762 | 0.27226 | 0.028117 | 11.5166
Aspartate neg | KD100520_100520_B1 | B1 | 2351051145 | 284424100 | 144541742 | 8.266 | 16.2656 | 500 | 522.0809 | 22.0809 | 4.4162 | 510.0297 | 10.0297 | 2.0059 | 522.0424 | 22.0424 | 4.4085
Aspartate neg | KD100520_100520_B2 | B2 | 1327840381 | 339683581 | 174195129 | 3.9091 | 7.6227 | 250 | 246.8956 | 3.1044 | 1.2417 | 230.9794 | 19.0206 | 7.6083 | 234.5377 | 15.4623 | 6.1849
Aspartate neg | KD100520_100520_B3 | B3 | 961041717 | 437887371 | 229464360 | 2.1947 | 4.1882 | 125 | 138.6187 | 13.6187 | 10.895 | 125.4272 | 0.42719 | 0.34175 | 124.6359 | 0.36408 | 0.29127
Aspartate neg | KD100520_100520_B4 | B4 | 582448533 | 489589565 | 255908952 | 1.1897 | 2.276 | 62.5 | 75.1394 | 12.6394 | 20.223 | 65.6244 | 3.1244 | 4.999 | 65.469 | 2.969 | 4.7504
Aspartate neg | KD100520_100520_B5 | B5 | 314915065 | 557059510 | 295553335 | 0.56532 | 1.0655 | 31.25 | 35.7054 | 4.4554 | 14.2573 | 29.8713 | 1.3787 | 4.4118 | 29.3807 | 1.8693 | 5.9818
Aspartate neg | KD100520_100520_B6 | B6 | 171280638 | 544716890 | 289023169 | 0.31444 | 0.59262 | 15.625 | 19.86 | 4.235 | 27.1041 | 16.0611 | 0.4361 | 2.791 | 15.8157 | 0.1907 | 1.2205
Aspartate neg | KD100520_100520_B7 | B7 | 88581805 | 555556265 | 292369008 | 0.15945 | 0.30298 | 7.8125 | 10.0707 | 2.2582 | 28.9046 | 7.8308 | 0.018331 | 0.23463 | 7.7893 | 0.023222 | 0.29724
Aspartate neg | KD100520_100520_B8 | B8 | 47199591 | 551557126 | 291575920 | 0.085575 | 0.16188 | 3.9062 | 5.4049 | 1.4987 | 38.3661 | 4.0543 | 0.14807 | 3.7905 | 4.0189 | 0.11265 | 2.8839
Aspartate neg | KD100520_100520_B9 | B9 | 25661736 | 570644259 | 301015922 | 0.04497 | 0.08525 | 1.9531 | 2.8403 | 0.88717 | 45.4229 | 2.0528 | 0.099641 | 5.1016 | 2.0422 | 0.08911 | 4.5624
Aspartate neg | KD100520_100520_B10 | B10 | 11943980 | 569554225 | 303880078 | 0.020971 | 0.039305 | 0.97656 | 1.3245 | 0.34795 | 35.6301 | 0.91597 | 0.060588 | 6.2042 | 0.90183 | 0.074728 | 7.6521
Aspartate neg | KD100520_100520_B11 | B11 | 6623166 | 547859161 | 281761361 | 0.012089 | 0.023506 | 0.48828 | 0.76355 | 0.27527 | 56.3755 | 0.51149 | 0.023211 | 4.7537 | 0.52412 | 0.035836 | 7.3392
Aspartate neg | KD100520_100520_B12 | B12 | 3057799 | 527204029 | 272054171 | 0.0058 | 0.01124 | 0.24414 | 0.36633 | 0.12219 | 50.0487 | 0.2352 | 0.008941 | 3.6622 | 0.24052 | 0.0036214 | 1.4833
Aspartate neg | KD100520_100520_C1 | C1 | 2138084572 | 285101688 | 140146969 | 7.4994 | 15.256 | 500 | 473.6605 | 26.3395 | 5.2679 | 460.1312 | 39.8688 | 7.9738 | 487.897 | 12.103 | 2.4206
Aspartate neg | KD100520_100520_C2 | C2 | 1442565850 | 341805349 | 184882070 | 4.2204 | 7.8026 | 250 | 266.5624 | 16.5624 | 6.625 | 250.4855 | 0.48547 | 0.19419 | 240.3853 | 9.6147 | 3.8459
Aspartate neg | KD100520_100520_C3 | C3 | 890883894 | 404842018 | 218820476 | 2.2006 | 4.0713 | 125 | 138.9881 | 13.9881 | 11.1905 | 125.7808 | 0.78076 | 0.62461 | 120.9664 | 4.0336 | 3.2269
Aspartate neg | KD100520_100520_C4 | C4 | 561035183 | 489106907 | 265578022 | 1.1471 | 2.1125 | 62.5 | 72.4483 | 9.9483 | 15.9173 | 63.1409 | 0.64086 | 1.0254 | 60.5144 | 1.9856 | 3.177
Aspartate neg | KD100520_100520_C5 | C5 | 327798771 | 530772114 | 289379369 | 0.61759 | 1.1328 | 31.25 | 39.0069 | 7.7569 | 24.8221 | 32.8006 | 1.5506 | 4.9619 | 31.3419 | 0.091852 | 0.29393
Aspartate neg | KD100520_100520_C6 | C6 | 155286937 | 524013365 | 285861620 | 0.29634 | 0.54322 | 15.625 | 18.7169 | 3.0919 | 19.7884 | 15.0849 | 0.5401 | 3.4567 | 14.4274 | 1.1976 | 7.6649
Aspartate neg | KD100520_100520_C7 | C7 | 94766616 | 573211439 | 294164178 | 0.16533 | 0.32216 | 7.8125 | 10.442 | 2.6295 | 33.6573 | 8.1366 | 0.32406 | 4.1479 | 8.3106 | 0.49813 | 6.3761
Aspartate neg | KD100520_100520_C8 | C8 | 45680965 | 577724082 | 298140586 | 0.079071 | 0.15322 | 3.9062 | 4.9941 | 1.0878 | 27.8489 | 3.7291 | 0.17718 | 4.5359 | 3.7923 | 0.11393 | 2.9165
Aspartate neg | KD100520_100520_C9 | C9 | 24645962 | 597610019 | 308326995 | 0.041241 | 0.079934 | 1.9531 | 2.6048 | 0.65165 | 33.3644 | 1.8732 | 0.079969 | 4.0944 | 1.908 | 0.045092 | 2.3087
Aspartate neg | KD100520_100520_C10 | C10 | 10885397 | 573603632 | 296331052 | 0.018977 | 0.036734 | 0.97656 | 1.1986 | 0.22204 | 22.7367 | 0.82413 | 0.15244 | 15.6094 | 0.83967 | 0.13689 | 14.0174
Aspartate neg | KD100520_100520_C11 | C11 | 6233618 | 571195357 | 295335400 | 0.010913 | 0.021107 | 0.48828 | 0.68928 | 0.201 | 41.1652 | 0.45902 | 0.029264 | 5.9932 | 0.4678 | 0.020477 | 4.1936
Aspartate neg | KD100520_100520_C12 | C12 | 2910486 | 524159301 | 271417017 | 0.0055527 | 0.010723 | 0.24414 | 0.35071 | 0.10657 | 43.6496 | 0.2246 | 0.019538 | 8.0028 | 0.22887 | 0.015272 | 6.2552
Aspartate neg |  | CAA |  |  |  | 1 |  | 50 | 63.16 |  |  | 54.611
[8 empty] | Mean |  |  | 28.6896 |  |  | 5.8672 |  |  | 5.9481
[8 empty] | Sum |  | 239.9072 |  |  | 161.3105 |  |  | 130.6431
Conclusion: Power regression is best over a wide range
[18 empty] | Before norm: | After norm:

## Aspartate pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A1 | 53549104 |  | 748665391 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A2 | 29681327 |  | 424203181 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A3 | 22936067 |  | 287977111 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A4 | 12858239 |  | 172705395 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A5 | 9221627 |  | 120085593 | C4H7NO4 | 5944.64 | Calibration Amount Diff 28.081% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A6 | 4390000 |  | 56849725 | C4H7NO4 | 10061.42
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A7 | 2143019 |  | 26112803 | C4H7NO4 | 14503.79 | Calibration Amount Diff 21.752% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A8 | 1135280 |  | 12244585 | C4H7NO4 | 5067.24 | Calibration Amount Diff 46.661% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A9 | 505431 |  | 6132883 | C4H7NO4 | 587.59 | Calibration Amount Diff 117.349% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A10 | 193398 |  | 2345006 | C4H7NO4 | 228.58 | Calibration Amount Diff 233.788% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A11 | 109039 |  | 1248809 | C4H7NO4 | 173.85 | Calibration Amount Diff 519.886% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_A12 | 54812 |  | 688995 | C4H7NO4 | 130.05 | Calibration Amount Diff 1080.424% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B1 | 58972977 |  | 849803183 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B2 | 33377228 |  | 479400813 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B3 | 22503796 |  | 296244911 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B4 | 13776902 |  | 176526502 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B5 | 7240974 |  | 95627525 | C4H7NO4 | 108171.06
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B6 | 4319840 |  | 55529812 | C4H7NO4 | 6989.22
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B7 | 1911230 |  | 24631279 | C4H7NO4 | 2071.27
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B8 | 1042805 |  | 12720514 | C4H7NO4 | 1361.95 | Calibration Amount Diff 50.43% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B9 | 448842 |  | 5839965 | C4H7NO4 | 821.74 | Calibration Amount Diff 113.597% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B10 | 183938 |  | 2317488 | C4H7NO4 | 342.17 | Calibration Amount Diff 234.21% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B11 | 98378 |  | 1239426 | C4H7NO4 | 239.07 | Calibration Amount Diff 517.774% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_B12 | 51534 |  | 630168 | C4H7NO4 | 178.95 | Calibration Amount Diff 1078.867% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C4H7NO4 | N/A
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C1 | 60270223 |  | 853127861 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C2 | 36346371 |  | 486139252 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C3 | 22065953 |  | 290951174 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C4 | 13014884 |  | 172358862 | C4H7NO4 | INF
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C5 | 8109909 |  | 105128384 | C4H7NO4 | 9773.63
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C6 | 3704819 |  | 47314373 | C4H7NO4 | 3696.55
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C7 | 2396666 |  | 28584028 | C4H7NO4 | 5013.54 | Calibration Amount Diff 24.471% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C8 | 1007846 |  | 12298228 | C4H7NO4 | 3083.34 | Calibration Amount Diff 46.707% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C9 | 424732 |  | 5324070 | C4H7NO4 | 1771.86 | Calibration Amount Diff 104.496% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C10 | 206811 |  | 2474505 | C4H7NO4 | 364.89 | Calibration Amount Diff 235.185% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C11 | 90979 |  | 1113405 | C4H7NO4 | 215.55 | Calibration Amount Diff 510.581% (absolute) > max calibration amount diff 20%
Aspartate pos | 10.26 | Target Compound | KD100520_100520_C12 | 38710 |  | 553451 | C4H7NO4 | 98.07 | Calibration Amount Diff 1070.437% (absolute) > max calibration amount diff 20%

## Aspartate U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A1 | 9694770 |  | 139365647 | [13]C4H7NO4 | 347887.89 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 25.189% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A2 | 12305381 |  | 180939922 | [13]C4H7NO4 | 115667.95 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 67.48% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A3 | 13943346 |  | 207620179 | [13]C4H7NO4 | 175109.87 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 184.195% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A4 | 14544878 |  | 234269110 | [13]C4H7NO4 | 139147.48 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 391.562% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A5 | 18778850 |  | 279183435 | [13]C4H7NO4 | 36322.25 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 978.623% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A6 | 19658241 |  | 299509680 | [13]C4H7NO4 | 157157.10 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -1211.83% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A7 | 17872368 |  | 289102108 | [13]C4H7NO4 | 97660.90 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -3095.778% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A8 | 17977613 |  | 301279313 | [13]C4H7NO4 | 77167.32 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -7773.744% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A9 | 18665476 |  | 285650813 | [13]C4H7NO4 | 34273.74 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -16245.688% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A10 | 18238213 |  | 294719737 | [13]C4H7NO4 | 54180.55 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -25762.042% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A11 | 16664694 |  | 273647607 | [13]C4H7NO4 | 23898.95 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 48237.407% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_A12 | 18828298 |  | 289213058 | [13]C4H7NO4 | 28776.39 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 125181.695% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B1 | 9875602 |  | 144541742 | [13]C4H7NO4 | 96173.63 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -21.659% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B2 | 11330705 |  | 174195129 | [13]C4H7NO4 | 97541.66 | Rsquared 0.259 < R2Threshold 0.99
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B3 | 15759106 |  | 229464360 | [13]C4H7NO4 | 109789.76 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -42.497% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B4 | 16890866 |  | 255908952 | [13]C4H7NO4 | 68505.99 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 57.822% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B5 | 18479720 |  | 295553335 | [13]C4H7NO4 | 46419.84 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -292.273% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B6 | 18687551 |  | 289023169 | [13]C4H7NO4 | 101753.20 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -488.897% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B7 | 18992437 |  | 292369008 | [13]C4H7NO4 | 20103.78 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 241.551% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B8 | 19819371 |  | 291575920 | [13]C4H7NO4 | 32139.60 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -646.348% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B9 | 19364484 |  | 301015922 | [13]C4H7NO4 | 57394.21 | Rsquared 0.259 < R2Threshold 0.99
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B10 | 19225968 |  | 303880078 | [13]C4H7NO4 | 45582.78 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -12411.485% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B11 | 17493357 |  | 281761361 | [13]C4H7NO4 | 43243.48 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 54902.481% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_B12 | 16615180 |  | 272054171 | [13]C4H7NO4 | 47187.97 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 95528.526% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C4H7NO4 | N/A
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C1 | 9558228 |  | 140146969 | [13]C4H7NO4 | 15806.76 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 45.486% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C2 | 13185600 |  | 184882070 | [13]C4H7NO4 | 187872.30 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -207.561% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C3 | 14772025 |  | 218820476 | [13]C4H7NO4 | 24162.89 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -308.815% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C4 | 16976462 |  | 265578022 | [13]C4H7NO4 | 53962.44 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -597.687% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C5 | 19161594 |  | 289379369 | [13]C4H7NO4 | 58288.16 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -1238.786% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C6 | 19074580 |  | 285861620 | [13]C4H7NO4 | 35523.57 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff -2418.805% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C7 | 19860474 |  | 294164178 | [13]C4H7NO4 | 154507.86 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 3622.162% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C8 | 19760891 |  | 298140586 | [13]C4H7NO4 | 25386.08 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 5859% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C9 | 19762001 |  | 308326995 | [13]C4H7NO4 | 27841.26 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 11949.411% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C10 | 18498461 |  | 296331052 | [13]C4H7NO4 | 32322.66 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 22581.807% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C11 | 18334051 |  | 295335400 | [13]C4H7NO4 | 48183.21 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 43509.55% (absolute) > max calibration amount diff 20%
Aspartate U-13C neg | 10.26 | Target Compound | KD100520_100520_C12 | 17839426 |  | 271417017 | [13]C4H7NO4 | 88010.80 | Rsquared 0.259 < R2Threshold 0.99. Calibration Amount Diff 80778.356% (absolute) > max calibration amount diff 20%

## Aspartate U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A1 | 2618469 |  | 33606975 | [13]C4H7NO4 | 8830.96 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 24.374% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A2 | 3800922 |  | 50715474 | [13]C4H7NO4 | INF | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 58.663% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A3 | 4763241 |  | 58818760 | [13]C4H7NO4 | 10969.88 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 246.616% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A4 | 5776226 |  | 74739836 | [13]C4H7NO4 | 18336.73 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 232.484% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A5 | 6377157 |  | 81111788 | [13]C4H7NO4 | 6880.55 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 634.641% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A6 | 7401046 |  | 97697257 | [13]C4H7NO4 | 14626.26 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -836.053% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A7 | 7982022 |  | 104888833 | [13]C4H7NO4 | 10438.10 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -1578.246% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A8 | 8963929 |  | 110764265 | [13]C4H7NO4 | 21449.68 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -2702.454% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A9 | 8371762 |  | 107786185 | [13]C4H7NO4 | 6049.95 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -1544.133% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A10 | 8565216 |  | 118875611 | [13]C4H7NO4 | 7955.41 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -12060.883% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A11 | 8605193 |  | 106956061 | [13]C4H7NO4 | 5900.87 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 23923.316% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_A12 | 8917315 |  | 114497821 | [13]C4H7NO4 | 5524.74 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 30018.378% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B1 | 2901732 |  | 37688947 | [13]C4H7NO4 | INF | Rsquared 0.577 < R2Threshold 0.99
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B2 | 4069063 |  | 54743190 | [13]C4H7NO4 | INF | Rsquared 0.577 < R2Threshold 0.99
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B3 | 5182307 |  | 64193505 | [13]C4H7NO4 | 35220.65 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 110.983% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B4 | 6075601 |  | 75719314 | [13]C4H7NO4 | 20794.06 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 150.391% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B5 | 7252771 |  | 97037486 | [13]C4H7NO4 | 25423.95 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -72.049% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B6 | 7911341 |  | 104014504 | [13]C4H7NO4 | 14022.12 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -197.377% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B7 | 7427518 |  | 100736733 | [13]C4H7NO4 | 9725.80 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -622.488% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B8 | 8151517 |  | 108848233 | [13]C4H7NO4 | 16937.15 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -1807.776% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B9 | 7896408 |  | 108368187 | [13]C4H7NO4 | 10862.38 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -2682.444% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B10 | 8521838 |  | 115078877 | [13]C4H7NO4 | 10400.23 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -6057.76% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B11 | 8173958 |  | 109757737 | [13]C4H7NO4 | 6325.11 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 25077.878% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_B12 | 8616066 |  | 107261327 | [13]C4H7NO4 | 7109.45 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 30887.248% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C4H7NO4 | N/A
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C1 | 3385281 |  | 42517754 | [13]C4H7NO4 | INF | Rsquared 0.577 < R2Threshold 0.99
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C2 | 4230722 |  | 52144533 | [13]C4H7NO4 | 11730.00 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -102.778% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C3 | 5246239 |  | 65532191 | [13]C4H7NO4 | 16235.79 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -62.978% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C4 | 6066545 |  | 80269307 | [13]C4H7NO4 | 19458.11 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -222.575% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C5 | 7377923 |  | 97590696 | [13]C4H7NO4 | 23878.19 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -533.647% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C6 | 7865116 |  | 102652670 | [13]C4H7NO4 | 10199.16 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -1563.504% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C7 | 8467331 |  | 106544708 | [13]C4H7NO4 | 9329.30 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 916.427% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C8 | 8078690 |  | 108878173 | [13]C4H7NO4 | 4034.52 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff -304.158% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C9 | 8235363 |  | 110798312 | [13]C4H7NO4 | 6977.24 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 1005.78% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C10 | 9054568 |  | 116551311 | [13]C4H7NO4 | 8985.23 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 9887.917% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C11 | 8955138 |  | 114260835 | [13]C4H7NO4 | 9722.01 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 7967.67% (absolute) > max calibration amount diff 20%
Aspartate U-13C pos | 10.26 | Target Compound | KD100520_100520_C12 | 7935901 |  | 108203601 | [13]C4H7NO4 | 16184.66 | Rsquared 0.577 < R2Threshold 0.99. Calibration Amount Diff 40388.978% (absolute) > max calibration amount diff 20%

## Aspartate U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A1 | 19037920 |  | 280643581 | [13]C4H7[15]NO4 | 1120040.52
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A2 | 23963501 |  | 356986169 | [13]C4H7[15]NO4 | 46982.12
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A3 | 27065255 |  | 407101577 | [13]C4H7[15]NO4 | 63717.67
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A4 | 28002177 |  | 457223332 | [13]C4H7[15]NO4 | 24862.12
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A5 | 36823435 |  | 546458563 | [13]C4H7[15]NO4 | 29515.60
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A6 | 36963175 |  | 558592743 | [13]C4H7[15]NO4 | 28565.04
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A7 | 33400205 |  | 536194956 | [13]C4H7[15]NO4 | 18605.01
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A8 | 33162472 |  | 555430116 | [13]C4H7[15]NO4 | 23666.94
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A9 | 34525340 |  | 525869578 | [13]C4H7[15]NO4 | 21117.19
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A10 | 33773294 |  | 545779050 | [13]C4H7[15]NO4 | 21702.40
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A11 | 32399484 |  | 530421817 | [13]C4H7[15]NO4 | 24110.67
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_A12 | 36844222 |  | 564373570 | [13]C4H7[15]NO4 | 21813.94
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B1 | 19297873 |  | 284424100 | [13]C4H7[15]NO4 | 61223.86
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B2 | 21939564 |  | 339683581 | [13]C4H7[15]NO4 | 203059.96
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B3 | 30040144 |  | 437887371 | [13]C4H7[15]NO4 | 59233.91
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B4 | 31973758 |  | 489589565 | [13]C4H7[15]NO4 | 26033.16
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B5 | 35228412 |  | 557059510 | [13]C4H7[15]NO4 | 19178.58
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B6 | 35455910 |  | 544716890 | [13]C4H7[15]NO4 | 21868.17
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B7 | 35914959 |  | 555556265 | [13]C4H7[15]NO4 | 27739.35
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B8 | 37219366 |  | 551557126 | [13]C4H7[15]NO4 | 35787.15
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B9 | 36914520 |  | 570644259 | [13]C4H7[15]NO4 | 30969.20
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B10 | 36002497 |  | 569554225 | [13]C4H7[15]NO4 | 18345.33
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B11 | 34142065 |  | 547859161 | [13]C4H7[15]NO4 | 39359.20
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_B12 | 32013758 |  | 527204029 | [13]C4H7[15]NO4 | 17484.71
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C4H7[15]NO4 | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C1 | 19238101 |  | 285101688 | [13]C4H7[15]NO4 | 191482.13
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C2 | 24254444 |  | 341805349 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C3 | 27052959 |  | 404842018 | [13]C4H7[15]NO4 | 93264.76
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C4 | 31169477 |  | 489106907 | [13]C4H7[15]NO4 | 22431.58
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C5 | 35126919 |  | 530772114 | [13]C4H7[15]NO4 | 15426.60
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C6 | 35029962 |  | 524013365 | [13]C4H7[15]NO4 | 21161.81
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C7 | 39053119 |  | 573211439 | [13]C4H7[15]NO4 | 28431.68
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C8 | 38380981 |  | 577724082 | [13]C4H7[15]NO4 | 34586.78
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C9 | 38614890 |  | 597610019 | [13]C4H7[15]NO4 | 22987.64
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C10 | 35699993 |  | 573603632 | [13]C4H7[15]NO4 | 16637.45
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C11 | 35290122 |  | 571195357 | [13]C4H7[15]NO4 | 26768.91
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD100520_100520_C12 | 34323696 |  | 524159301 | [13]C4H7[15]NO4 | 33671.91

## Aspartate U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A1 | 5566874 |  | 75499465 | [13]C4H7[15]NO4 | 46229.11 | 10.284:Peak area 75499465.131 is out of bounds (ISTD Minimum Recovery 88093468.301 and ISTD Max Recovery 264280404.904)
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A2 | 7649313 |  | 107552240 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A3 | 10168412 |  | 125883188 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A4 | 11386794 |  | 151389363 | [13]C4H7[15]NO4 | 5730.85
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A5 | 12543090 |  | 165150313 | [13]C4H7[15]NO4 | 67863.45
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A6 | 13776464 |  | 183795677 | [13]C4H7[15]NO4 | 5922.89
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A7 | 15053810 |  | 197303827 | [13]C4H7[15]NO4 | 10471.41
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A8 | 16787177 |  | 208994266 | [13]C4H7[15]NO4 | 8974.51
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A9 | 15971645 |  | 206735356 | [13]C4H7[15]NO4 | 7262.43
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A10 | 16111858 |  | 223546821 | [13]C4H7[15]NO4 | 4482.56
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A11 | 16877584 |  | 212081800 | [13]C4H7[15]NO4 | 8839.95
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_A12 | 17271734 |  | 224749781 | [13]C4H7[15]NO4 | 4162.35
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B1 | 5939742 |  | 80437719 | [13]C4H7[15]NO4 | INF | 10.297:Peak area 80437719.209 is out of bounds (ISTD Minimum Recovery 88093468.301 and ISTD Max Recovery 264280404.904)
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B2 | 7856793 |  | 110910147 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B3 | 10215325 |  | 131778086 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B4 | 11782661 |  | 151526385 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B5 | 14027324 |  | 187678720 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B6 | 15033596 |  | 200085841 | [13]C4H7[15]NO4 | 5967.13
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B7 | 14386234 |  | 192667527 | [13]C4H7[15]NO4 | 6345.63
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B8 | 15252145 |  | 206979699 | [13]C4H7[15]NO4 | 5222.96
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B9 | 15053989 |  | 206817350 | [13]C4H7[15]NO4 | 7486.51
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B10 | 15909001 |  | 219243082 | [13]C4H7[15]NO4 | 3523.74
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B11 | 16573143 |  | 217922756 | [13]C4H7[15]NO4 | 3748.05
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_B12 | 16800190 |  | 210647952 | [13]C4H7[15]NO4 | 7703.53
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C4H7[15]NO4 | N/A
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C1 | 7165102 |  | 91202039 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C2 | 7699954 |  | 100494506 | [13]C4H7[15]NO4 | INF
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C3 | 9955379 |  | 127833434 | [13]C4H7[15]NO4 | 14651.73
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C4 | 11558796 |  | 152299221 | [13]C4H7[15]NO4 | 5138.37
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C5 | 13684153 |  | 182768243 | [13]C4H7[15]NO4 | 7860.39
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C6 | 14442029 |  | 188393378 | [13]C4H7[15]NO4 | 4760.82
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C7 | 16644559 |  | 209424126 | [13]C4H7[15]NO4 | 10235.19
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C8 | 15179264 |  | 209784469 | [13]C4H7[15]NO4 | 3743.42
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C9 | 15839151 |  | 214920337 | [13]C4H7[15]NO4 | 4807.71
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C10 | 17766204 |  | 230056675 | [13]C4H7[15]NO4 | 6944.46
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C11 | 17214905 |  | 222536617 | [13]C4H7[15]NO4 | 5408.99
Aspartate U-13C, U-15N pos | 10.26 | Internal Standard | KD100520_100520_C12 | 15461923 |  | 213639315 | [13]C4H7[15]NO4 | 5577.10

## Cystine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Cystine neg | 10.6 | Target Compound | KD100520_100520_A1 | 94505460 |  | 1252513134 | C6H12N2O4S2 | INF
Cystine neg | 10.6 | Target Compound | KD100520_100520_A2 | 47005979 |  | 632309084 | C6H12N2O4S2 | 730883.05
Cystine neg | 10.6 | Target Compound | KD100520_100520_A3 | 36435674 |  | 497322659 | C6H12N2O4S2 | INF
Cystine neg | 10.6 | Target Compound | KD100520_100520_A4 | 19487145 |  | 245647400 | C6H12N2O4S2 | 31282.43
Cystine neg | 10.6 | Target Compound | KD100520_100520_A5 | 13177847 |  | 164106598 | C6H12N2O4S2 | 8492.34 | Calibration Amount Diff 38.337% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_A6 | 5794496 |  | 67820811 | C6H12N2O4S2 | 11039.43
Cystine neg | 10.6 | Target Compound | KD100520_100520_A7 | 2769950 |  | 31711915 | C6H12N2O4S2 | 4145.91
Cystine neg | 10.6 | Target Compound | KD100520_100520_A8 | 1170071 |  | 13637384 | C6H12N2O4S2 | 3028.45 | Calibration Amount Diff -63.034% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_A9 | 673190 |  | 6462602 | C6H12N2O4S2 | 3838.87 | Calibration Amount Diff -139.5% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_A10 | 243494 |  | 2439398 | C6H12N2O4S2 | 1853.96 | Calibration Amount Diff -312.71% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_A11 | 132681 |  | 1261729 | C6H12N2O4S2 | 23240.90 | Calibration Amount Diff -602.549% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_A12 | 54436 |  | 450351 | C6H12N2O4S2 | INF | Calibration Amount Diff -1210.792% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_B1 | 89814658 |  | 1137504241 | C6H12N2O4S2 | INF
Cystine neg | 10.6 | Target Compound | KD100520_100520_B2 | 58264328 |  | 764445445 | C6H12N2O4S2 | INF
Cystine neg | 10.6 | Target Compound | KD100520_100520_B3 | 35752999 |  | 472444705 | C6H12N2O4S2 | INF
Cystine neg | 10.6 | Target Compound | KD100520_100520_B4 | 21813155 |  | 268915330 | C6H12N2O4S2 | 44201.57
Cystine neg | 10.6 | Target Compound | KD100520_100520_B5 | 10528720 |  | 123563194 | C6H12N2O4S2 | 14308.54
Cystine neg | 10.6 | Target Compound | KD100520_100520_B6 | 5725701 |  | 64826112 | C6H12N2O4S2 | 9425.49
Cystine neg | 10.6 | Target Compound | KD100520_100520_B7 | 2466618 |  | 28874365 | C6H12N2O4S2 | 4620.68 | Calibration Amount Diff -22.095% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_B8 | 1268694 |  | 13720135 | C6H12N2O4S2 | 2072.54 | Calibration Amount Diff -59.961% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_B9 | 715919 |  | 6910224 | C6H12N2O4S2 | 2550.18 | Calibration Amount Diff -138.468% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_B10 | 206409 |  | 2056634 | C6H12N2O4S2 | 3978.58 | Calibration Amount Diff -323.358% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_B11 | 83271 |  | 976748 | C6H12N2O4S2 | INF | Calibration Amount Diff -619.756% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_B12 | 38525 |  | 390003 | C6H12N2O4S2 | INF | Calibration Amount Diff -1216.534% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H12N2O4S2 | N/A
Cystine neg | 10.6 | Target Compound | KD100520_100520_C1 | 79826476 |  | 1049871319 | C6H12N2O4S2 | INF
Cystine neg | 10.6 | Target Compound | KD100520_100520_C2 | 54567030 |  | 736739362 | C6H12N2O4S2 | INF
Cystine neg | 10.6 | Target Compound | KD100520_100520_C3 | 38235246 |  | 445067293 | C6H12N2O4S2 | INF
Cystine neg | 10.6 | Target Compound | KD100520_100520_C4 | 20496198 |  | 252618035 | C6H12N2O4S2 | 11133.23
Cystine neg | 10.6 | Target Compound | KD100520_100520_C5 | 10779957 |  | 125351007 | C6H12N2O4S2 | 17257.73
Cystine neg | 10.6 | Target Compound | KD100520_100520_C6 | 4781212 |  | 57555074 | C6H12N2O4S2 | 8385.99
Cystine neg | 10.6 | Target Compound | KD100520_100520_C7 | 2815535 |  | 33109630 | C6H12N2O4S2 | 12589.84
Cystine neg | 10.6 | Target Compound | KD100520_100520_C8 | 1326372 |  | 14224665 | C6H12N2O4S2 | 5224.86 | Calibration Amount Diff -63.732% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_C9 | 651801 |  | 6542208 | C6H12N2O4S2 | 3811.94 | Calibration Amount Diff -145.888% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_C10 | 140317 |  | 1613369 | C6H12N2O4S2 | 5069.58 | Calibration Amount Diff -337.376% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_C11 | 92528 |  | 1014112 | C6H12N2O4S2 | INF | Calibration Amount Diff -622.595% (absolute) > max calibration amount diff 20%
Cystine neg | 10.6 | Target Compound | KD100520_100520_C12 | 28859 |  | 351420 | C6H12N2O4S2 | INF | Calibration Amount Diff -1213.07% (absolute) > max calibration amount diff 20%

## Cystine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Cystine pos | 10.6 | Target Compound | KD100520_100520_A1 | 160359697 |  | 2014598388 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_A2 | 88407033 |  | 1214519945 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_A3 | 65017337 |  | 896068423 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_A4 | 39090842 |  | 455025580 | C6H12N2O4S2 | 149334.47
Cystine pos | 10.6 | Target Compound | KD100520_100520_A5 | 27980462 |  | 331065273 | C6H12N2O4S2 | 38525.65 | Calibration Amount Diff 36.416% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_A6 | 8883048 |  | 122549351 | C6H12N2O4S2 | 6660.30
Cystine pos | 10.6 | Target Compound | KD100520_100520_A7 | 4139223 |  | 55993344 | C6H12N2O4S2 | 9873.50
Cystine pos | 10.6 | Target Compound | KD100520_100520_A8 | 1976381 |  | 25376084 | C6H12N2O4S2 | 4101.88 | Calibration Amount Diff -45.64% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_A9 | 1185545 |  | 12407448 | C6H12N2O4S2 | 3388.08 | Calibration Amount Diff -99.631% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_A10 | 363857 |  | 4376600 | C6H12N2O4S2 | 3063.62 | Calibration Amount Diff -234.554% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_A11 | 142842 |  | 2046381 | C6H12N2O4S2 | 8043.85 | Calibration Amount Diff -443.455% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_A12 | 73234 |  | 839629 | C6H12N2O4S2 | INF | Calibration Amount Diff -873.634% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_B1 | 188541399 |  | 2244505909 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_B2 | 100346349 |  | 1305136074 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_B3 | 61404362 |  | 829313636 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_B4 | 40431330 |  | 515895193 | C6H12N2O4S2 | 1086464.69
Cystine pos | 10.6 | Target Compound | KD100520_100520_B5 | 18716964 |  | 233255475 | C6H12N2O4S2 | 18575.43
Cystine pos | 10.6 | Target Compound | KD100520_100520_B6 | 8522332 |  | 117075172 | C6H12N2O4S2 | 15823.92
Cystine pos | 10.6 | Target Compound | KD100520_100520_B7 | 5379675 |  | 63031874 | C6H12N2O4S2 | 4944.08
Cystine pos | 10.6 | Target Compound | KD100520_100520_B8 | 2170889 |  | 27222839 | C6H12N2O4S2 | 3526.66 | Calibration Amount Diff -41.251% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_B9 | 1075446 |  | 12197989 | C6H12N2O4S2 | 2664.06 | Calibration Amount Diff -101.761% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_B10 | 362244 |  | 4577153 | C6H12N2O4S2 | 2204.64 | Calibration Amount Diff -237.106% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_B11 | 170309 |  | 2120554 | C6H12N2O4S2 | 10401.79 | Calibration Amount Diff -447.142% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_B12 | 72616 |  | 840066 | C6H12N2O4S2 | INF | Calibration Amount Diff -873.256% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H12N2O4S2 | N/A
Cystine pos | 10.6 | Target Compound | KD100520_100520_C1 | 131433347 |  | 1833882898 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_C2 | 111046905 |  | 1420644342 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_C3 | 65989415 |  | 854009252 | C6H12N2O4S2 | INF
Cystine pos | 10.6 | Target Compound | KD100520_100520_C4 | 32635650 |  | 438288508 | C6H12N2O4S2 | 310302.37
Cystine pos | 10.6 | Target Compound | KD100520_100520_C5 | 16674126 |  | 229852041 | C6H12N2O4S2 | 107868.25
Cystine pos | 10.6 | Target Compound | KD100520_100520_C6 | 8570796 |  | 108545045 | C6H12N2O4S2 | 17879.54
Cystine pos | 10.6 | Target Compound | KD100520_100520_C7 | 5352161 |  | 64794191 | C6H12N2O4S2 | 9332.15
Cystine pos | 10.6 | Target Compound | KD100520_100520_C8 | 2063776 |  | 26093443 | C6H12N2O4S2 | 5632.88 | Calibration Amount Diff -46.216% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_C9 | 859697 |  | 10155595 | C6H12N2O4S2 | 3282.42 | Calibration Amount Diff -116.858% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_C10 | 342287 |  | 3646386 | C6H12N2O4S2 | 5716.40 | Calibration Amount Diff -249.955% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_C11 | 168575 |  | 1929059 | C6H12N2O4S2 | 1207.29 | Calibration Amount Diff -455.899% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD100520_100520_C12 | 57465 |  | 796026 | C6H12N2O4S2 | INF | Calibration Amount Diff -874.078% (absolute) > max calibration amount diff 20%

## Cystine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A1 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A2 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A3 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A4 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A5 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A6 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A7 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A8 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A9 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B1 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B2 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B3 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B4 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B5 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B6 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B7 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B8 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B9 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B10 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_B12 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H12N2O4S2 | N/A
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C1 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C2 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C3 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C4 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C5 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C6 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C7 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C8 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C9 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C10 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C11 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99
Cystine U-13C neg | 10.6 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0 < R2Threshold 0.99

## Cystine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A1 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A2 | 43378 |  | 186362 | [13]C6H12N2O4S2 | INF | Rsquared 0.811 < R2Threshold 0.99. Calibration Amount Diff 42.513% (absolute) > max calibration amount diff 20%
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A3 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A4 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A5 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A6 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A7 | 6607 |  | 104460 | [13]C6H12N2O4S2 | INF | Rsquared 0.811 < R2Threshold 0.99. Calibration Amount Diff 1247.694% (absolute) > max calibration amount diff 20%
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A8 | 1559 |  | 5603 | [13]C6H12N2O4S2 | INF | Rsquared 0.811 < R2Threshold 0.99. Calibration Amount Diff -3177.634% (absolute) > max calibration amount diff 20%
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A9 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B1 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B2 | 35846 |  | 125498 | [13]C6H12N2O4S2 | INF | Rsquared 0.811 < R2Threshold 0.99. Calibration Amount Diff -20.762% (absolute) > max calibration amount diff 20%
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B3 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B4 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B5 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B6 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B7 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B8 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B9 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B10 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_B12 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H12N2O4S2 | N/A
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C1 | 34153 |  | 160842 | [13]C6H12N2O4S2 | INF | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C2 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C3 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C4 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C5 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C6 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C7 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C8 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C9 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C10 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C11 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99
Cystine U-13C pos | 10.6 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | [13]C6H12N2O4S2 | N/A | Rsquared 0.811 < R2Threshold 0.99

## Cystine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A1 | 6241528 |  | 77789612 | [13]C6H12[15]N2O4S2 | 184586.37
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A2 | 6134370 |  | 80715391 | [13]C6H12[15]N2O4S2 | 85429.73
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A3 | 7861965 |  | 105320163 | [13]C6H12[15]N2O4S2 | 26112.78
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A4 | 8693709 |  | 102142096 | [13]C6H12[15]N2O4S2 | 23206.74
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A5 | 8589687 |  | 104570704 | [13]C6H12[15]N2O4S2 | 14342.74
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A6 | 8041482 |  | 94757296 | [13]C6H12[15]N2O4S2 | 13009.35
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A7 | 8471507 |  | 100844923 | [13]C6H12[15]N2O4S2 | 8267.31
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A8 | 7585169 |  | 93370297 | [13]C6H12[15]N2O4S2 | 12410.53
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A9 | 8303574 |  | 91535788 | [13]C6H12[15]N2O4S2 | 9657.79
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A10 | 7078781 |  | 93487071 | [13]C6H12[15]N2O4S2 | 9049.13
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A11 | 8401924 |  | 94541302 | [13]C6H12[15]N2O4S2 | 5755.75
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_A12 | 9180358 |  | 98721950 | [13]C6H12[15]N2O4S2 | 8484.32
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B1 | 5013400 |  | 61444564 | [13]C6H12[15]N2O4S2 | 9911.71
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B2 | 7651573 |  | 91191411 | [13]C6H12[15]N2O4S2 | 31097.80
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B3 | 8013507 |  | 100624798 | [13]C6H12[15]N2O4S2 | 17004.94
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B4 | 9153724 |  | 107576766 | [13]C6H12[15]N2O4S2 | 14188.56
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B5 | 9413634 |  | 107560087 | [13]C6H12[15]N2O4S2 | 14710.14
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B6 | 8980391 |  | 104572849 | [13]C6H12[15]N2O4S2 | 14872.02
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B7 | 8287784 |  | 94920213 | [13]C6H12[15]N2O4S2 | 8506.21
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B8 | 8044720 |  | 91378658 | [13]C6H12[15]N2O4S2 | 6457.34
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B9 | 8527701 |  | 96933844 | [13]C6H12[15]N2O4S2 | 7656.61
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B10 | 7709614 |  | 91203105 | [13]C6H12[15]N2O4S2 | 7828.82
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B11 | 7357945 |  | 93150625 | [13]C6H12[15]N2O4S2 | 11173.91
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_B12 | 8009710 |  | 95482172 | [13]C6H12[15]N2O4S2 | 7563.32
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H12[15]N2O4S2 | N/A
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C1 | 5246368 |  | 63150764 | [13]C6H12[15]N2O4S2 | 10789.75
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C2 | 6254312 |  | 80351744 | [13]C6H12[15]N2O4S2 | 44271.84
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C3 | 8434812 |  | 94645000 | [13]C6H12[15]N2O4S2 | 12766.85
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C4 | 8726813 |  | 105204003 | [13]C6H12[15]N2O4S2 | 8657.81
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C5 | 8715107 |  | 98910698 | [13]C6H12[15]N2O4S2 | 11513.06
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C6 | 8213729 |  | 99550159 | [13]C6H12[15]N2O4S2 | 9889.03
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C7 | 8766680 |  | 105388444 | [13]C6H12[15]N2O4S2 | 12800.70
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C8 | 8580281 |  | 98014024 | [13]C6H12[15]N2O4S2 | 9231.24
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C9 | 8581575 |  | 98597405 | [13]C6H12[15]N2O4S2 | 7589.62
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C10 | 7214411 |  | 90207122 | [13]C6H12[15]N2O4S2 | 4145.32
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C11 | 8423428 |  | 101270820 | [13]C6H12[15]N2O4S2 | 8789.27
Cystine U-13C, U-15N neg | 10.6 | Internal Standard | KD100520_100520_C12 | 6382881 |  | 80370261 | [13]C6H12[15]N2O4S2 | 8359.11

## Cystine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A1 | 10282158 |  | 117057415 | [13]C6H12[15]N2O4S2 | 24028.12
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A2 | 12264248 |  | 156160350 | [13]C6H12[15]N2O4S2 | 22032.06
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A3 | 14272823 |  | 191103602 | [13]C6H12[15]N2O4S2 | 263389.61
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A4 | 15904892 |  | 188762862 | [13]C6H12[15]N2O4S2 | 68579.23
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A5 | 18448870 |  | 214096661 | [13]C6H12[15]N2O4S2 | 55804.45
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A6 | 12186839 |  | 173003801 | [13]C6H12[15]N2O4S2 | 12419.55
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A7 | 13001468 |  | 179638089 | [13]C6H12[15]N2O4S2 | 13180.27
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A8 | 12991130 |  | 176633050 | [13]C6H12[15]N2O4S2 | 25715.11
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A9 | 14688866 |  | 176267519 | [13]C6H12[15]N2O4S2 | 26791.08
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A10 | 13835087 |  | 177774092 | [13]C6H12[15]N2O4S2 | 11737.48
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A11 | 11627062 |  | 169095783 | [13]C6H12[15]N2O4S2 | 18058.75
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_A12 | 13225409 |  | 175856595 | [13]C6H12[15]N2O4S2 | 25746.12
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B1 | 11335080 |  | 122679912 | [13]C6H12[15]N2O4S2 | INF
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B2 | 12304548 |  | 155368882 | [13]C6H12[15]N2O4S2 | 79799.33
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B3 | 13469402 |  | 179367737 | [13]C6H12[15]N2O4S2 | 163803.77
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B4 | 17611315 |  | 207796221 | [13]C6H12[15]N2O4S2 | 17030.78
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B5 | 15960029 |  | 203075115 | [13]C6H12[15]N2O4S2 | 55045.26
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B6 | 14151540 |  | 183744664 | [13]C6H12[15]N2O4S2 | 49526.74
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B7 | 17197383 |  | 202858893 | [13]C6H12[15]N2O4S2 | 13714.87
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B8 | 14374569 |  | 181968595 | [13]C6H12[15]N2O4S2 | 14443.04
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B9 | 13844445 |  | 176912128 | [13]C6H12[15]N2O4S2 | 18429.66
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B10 | 14380608 |  | 192677946 | [13]C6H12[15]N2O4S2 | 36073.02
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B11 | 14237588 |  | 184734284 | [13]C6H12[15]N2O4S2 | 6388.46
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_B12 | 13854605 |  | 174776577 | [13]C6H12[15]N2O4S2 | 22170.36
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H12[15]N2O4S2 | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C1 | 9501428 |  | 108961425 | [13]C6H12[15]N2O4S2 | 104650.87
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C2 | 13575045 |  | 156471157 | [13]C6H12[15]N2O4S2 | 46368.02
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C3 | 14563192 |  | 187012024 | [13]C6H12[15]N2O4S2 | 37724.66
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C4 | 13417477 |  | 181509261 | [13]C6H12[15]N2O4S2 | 23720.56
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C5 | 13388744 |  | 179141645 | [13]C6H12[15]N2O4S2 | 29917.74
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C6 | 14773869 |  | 183913425 | [13]C6H12[15]N2O4S2 | 33033.38
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C7 | 16663094 |  | 204615395 | [13]C6H12[15]N2O4S2 | 10704.78
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C8 | 13619897 |  | 182616564 | [13]C6H12[15]N2O4S2 | 17237.19
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C9 | 13504038 |  | 172893588 | [13]C6H12[15]N2O4S2 | 16076.37
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C10 | 13960578 |  | 187876996 | [13]C6H12[15]N2O4S2 | 14688.60
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C11 | 15107665 |  | 192923567 | [13]C6H12[15]N2O4S2 | 10718.02
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD100520_100520_C12 | 12435075 |  | 168044432 | [13]C6H12[15]N2O4S2 | 14342.65

## Glutamate neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A1 | 144319877 |  | 1891177579 | C5H9NO4 | 33540.73
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A2 | 81518289 |  | 999916215 | C5H9NO4 | 56027.67
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A3 | 51253552 |  | 674705812 | C5H9NO4 | 40295.22
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A4 | 29529166 |  | 383683318 | C5H9NO4 | 14845.36
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A5 | 23150363 |  | 292947592 | C5H9NO4 | 19696.27 | Calibration Amount Diff 43.08% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A6 | 9675700 |  | 134293289 | C5H9NO4 | 6948.79
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A7 | 4513569 |  | 61254237 | C5H9NO4 | 3315.66 | Calibration Amount Diff -27.045% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A8 | 2648804 |  | 32285163 | C5H9NO4 | 2983.51 | Calibration Amount Diff -82.48% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A9 | 1409513 |  | 17575356 | C5H9NO4 | 2261.61 | Calibration Amount Diff -184.069% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A10 | 651105 |  | 7668437 | C5H9NO4 | 778.12 | Calibration Amount Diff -429.016% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A11 | 368043 |  | 4511966 | C5H9NO4 | 661.30 | Calibration Amount Diff -858.174% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_A12 | 138294 |  | 1667331 | C5H9NO4 | 277.37 | Calibration Amount Diff -1786.903% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B1 | 150296666 |  | 1968337135 | C5H9NO4 | 60016.26
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B2 | 84353509 |  | 1062581292 | C5H9NO4 | 32787.30
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B3 | 56064729 |  | 711195131 | C5H9NO4 | 18687.36
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B4 | 32954820 |  | 431462227 | C5H9NO4 | 30851.33
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B5 | 16772214 |  | 212880834 | C5H9NO4 | 9506.67
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B6 | 9422645 |  | 122641376 | C5H9NO4 | 6091.64
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B7 | 4871764 |  | 62292229 | C5H9NO4 | 5282.77 | Calibration Amount Diff -29.279% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B8 | 2351183 |  | 31122481 | C5H9NO4 | 2417.99 | Calibration Amount Diff -79.577% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B9 | 1307874 |  | 17771609 | C5H9NO4 | 1874.46 | Calibration Amount Diff -183.882% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B10 | 602891 |  | 7705467 | C5H9NO4 | 728.90 | Calibration Amount Diff -430.166% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B11 | 331857 |  | 3678580 | C5H9NO4 | 456.30 | Calibration Amount Diff -878.282% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_B12 | 167013 |  | 1867602 | C5H9NO4 | 292.59 | Calibration Amount Diff -1777.924% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H9NO4 | N/A
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C1 | 137235060 |  | 1751910410 | C5H9NO4 | 78843.57
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C2 | 88127470 |  | 1133259534 | C5H9NO4 | 33429.27
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C3 | 48533557 |  | 661142136 | C5H9NO4 | 35987.14
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C4 | 30848782 |  | 399903387 | C5H9NO4 | 16796.85
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C5 | 16069843 |  | 213612772 | C5H9NO4 | 14569.86
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C6 | 7246687 |  | 97903204 | C5H9NO4 | 5638.41
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C7 | 4859840 |  | 65517504 | C5H9NO4 | 3835.95 | Calibration Amount Diff -25.184% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C8 | 2335796 |  | 31438097 | C5H9NO4 | 2605.90 | Calibration Amount Diff -86.542% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C9 | 1180599 |  | 15477089 | C5H9NO4 | 959.86 | Calibration Amount Diff -201.848% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C10 | 503942 |  | 6779077 | C5H9NO4 | 972.54 | Calibration Amount Diff -439.161% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C11 | 287447 |  | 3519020 | C5H9NO4 | 384.46 | Calibration Amount Diff -889.297% (absolute) > max calibration amount diff 20%
Glutamate neg | 10.05 | Target Compound | KD100520_100520_C12 | 119839 |  | 1525541 | C5H9NO4 | 139.66 | Calibration Amount Diff -1792.016% (absolute) > max calibration amount diff 20%

## Glu_neg_calibration
Compound | Filename | Sample ID | Area | Area_IS | Response_ratio | Conc.
Glutamate neg | KD100520_100520_A1 |  | 1891177579 | 240986600 | 7.8476 | 500
Glutamate neg | KD100520_100520_A2 |  | 999916215 | 264284568 | 3.7835 | 250
Glutamate neg | KD100520_100520_A3 |  | 674705812 | 284820366 | 2.3689 | 125
Glutamate neg | KD100520_100520_A4 |  | 383683318 | 321033146 | 1.1952 | 62.5
Glutamate neg | KD100520_100520_A5 |  | 292947592 | 368244124 | 0.79553 | 31.25
Glutamate neg | KD100520_100520_A6 |  | 134293289 | 374733452 | 0.35837 | 15.625
Glutamate neg | KD100520_100520_A7 |  | 61254237 | 374957436 | 0.16336 | 7.8125
Glutamate neg | KD100520_100520_A8 |  | 32285163 | 393294487 | 0.082089 | 3.9062
Glutamate neg | KD100520_100520_A9 |  | 17575356 | 395879776 | 0.044396 | 1.9531
Glutamate neg | KD100520_100520_A10 |  | 7668437 | 405462037 | 0.018913 | 0.97656
Glutamate neg | KD100520_100520_A11 |  | 4511966 | 408420148 | 0.011047 | 0.48828
Glutamate neg | KD100520_100520_A12 |  | 1667331 | 387392843 | 0.004304 | 0.24414
Glutamate neg | KD100520_100520_B1 |  | 1968337135 | 226296878 | 8.698 | 500
Glutamate neg | KD100520_100520_B2 |  | 1062581292 | 262005249 | 4.0556 | 250
Glutamate neg | KD100520_100520_B3 |  | 711195131 | 306639279 | 2.3193 | 125
Glutamate neg | KD100520_100520_B4 |  | 431462227 | 348213107 | 1.2391 | 62.5
Glutamate neg | KD100520_100520_B5 |  | 212880834 | 367357769 | 0.57949 | 31.25
Glutamate neg | KD100520_100520_B6 |  | 122641376 | 381096420 | 0.32181 | 15.625
Glutamate neg | KD100520_100520_B7 |  | 62292229 | 388029988 | 0.16053 | 7.8125
Glutamate neg | KD100520_100520_B8 |  | 31122481 | 370830305 | 0.083926 | 3.9062
Glutamate neg | KD100520_100520_B9 |  | 17771609 | 399769412 | 0.044455 | 1.9531
Glutamate neg | KD100520_100520_B10 |  | 7705467 | 411381272 | 0.018731 | 0.97656
Glutamate neg | KD100520_100520_B11 |  | 3678580 | 388966420 | 0.0094573 | 0.48828
Glutamate neg | KD100520_100520_B12 |  | 1867602 | 400861525 | 0.004659 | 0.24414
Glutamate neg | KD100520_100520_C1 |  | 1751910410 | 224674955 | 7.7975 | 500
Glutamate neg | KD100520_100520_C2 |  | 1133259534 | 256857834 | 4.412 | 250
Glutamate neg | KD100520_100520_C3 |  | 661142136 | 288491267 | 2.2917 | 125
Glutamate neg | KD100520_100520_C4 |  | 399903387 | 335061254 | 1.1935 | 62.5
Glutamate neg | KD100520_100520_C5 |  | 213612772 | 332299980 | 0.64283 | 31.25
Glutamate neg | KD100520_100520_C6 |  | 97903204 | 330178927 | 0.29652 | 15.625
Glutamate neg | KD100520_100520_C7 |  | 65517504 | 395354074 | 0.16572 | 7.8125
Glutamate neg | KD100520_100520_C8 |  | 31438097 | 395358079 | 0.079518 | 3.9062
Glutamate neg | KD100520_100520_C9 |  | 15477089 | 399213049 | 0.038769 | 1.9531
Glutamate neg | KD100520_100520_C10 |  | 6779077 | 391700520 | 0.017307 | 0.97656
Glutamate neg | KD100520_100520_C11 |  | 3519020 | 409844561 | 0.0085862 | 0.48828
Glutamate neg | KD100520_100520_C12 |  | 1525541 | 371917302 | 0.0041018 | 0.24414
[5 empty] | A/B | A/C | B/C
[5 empty] | 0.90223 | 1.0064 | 1.1155
[5 empty] | 0.93291 | 0.85754 | 0.91921
[5 empty] | 1.0214 | 1.0337 | 1.012
[5 empty] | 0.96455 | 1.0014 | 1.0382
[5 empty] | 1.3728 | 1.2375 | 0.90147
[5 empty] | 1.1136 | 1.2086 | 1.0853
[5 empty] | 1.0176 | 0.98579 | 0.96872
[5 empty] | 0.97811 | 1.0323 | 1.0554
[5 empty] | 0.99867 | 1.1451 | 1.1467
[5 empty] | 1.0097 | 1.0928 | 1.0823
[5 empty] | 1.1681 | 1.2866 | 1.1015
[5 empty] | 0.9238 | 1.0493 | 1.1358
|  |  |  | Mean | 1.0336 | 1.0781 | 1.0468

## Glutamate pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A1 | 149877078 |  | 2118585044 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A2 | 96091800 |  | 1296908539 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A3 | 64793982 |  | 908923903 | C5H9NO4 | 32334.31
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A4 | 44864006 |  | 576269420 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A5 | 34112605 |  | 436713477 | C5H9NO4 | 19138.17 | Calibration Amount Diff 32.538% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A6 | 16382590 |  | 231182530 | C5H9NO4 | 9981.04
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A7 | 9021126 |  | 113199053 | C5H9NO4 | 4627.57
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A8 | 5018346 |  | 65744014 | C5H9NO4 | 3560.30
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A9 | 2305664 |  | 32811601 | C5H9NO4 | 2063.51 | Calibration Amount Diff -42.486% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A10 | 1116507 |  | 14670510 | C5H9NO4 | 1287.06 | Calibration Amount Diff -114.896% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A11 | 620519 |  | 7784040 | C5H9NO4 | 549.71 | Calibration Amount Diff -217.652% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_A12 | 305755 |  | 3570231 | C5H9NO4 | 316.85 | Calibration Amount Diff -450.58% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B1 | 185046951 |  | 2369621337 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B2 | 106852110 |  | 1448765432 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B3 | 66045356 |  | 917913653 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B4 | 46877553 |  | 617328846 | C5H9NO4 | 14705.51
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B5 | 25694441 |  | 354119036 | C5H9NO4 | 14474.15
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B6 | 15054904 |  | 215139858 | C5H9NO4 | 10855.61
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B7 | 7746788 |  | 107312180 | C5H9NO4 | 6777.11
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B8 | 4748031 |  | 64403840 | C5H9NO4 | 4190.55
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B9 | 2603193 |  | 33609060 | C5H9NO4 | 1837.29 | Calibration Amount Diff -43.309% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B10 | 1144005 |  | 15094208 | C5H9NO4 | 774.16 | Calibration Amount Diff -116.083% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B11 | 580687 |  | 7348363 | C5H9NO4 | 657.29 | Calibration Amount Diff -223.145% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_B12 | 318933 |  | 3680767 | C5H9NO4 | 311.01 | Calibration Amount Diff -444.777% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H9NO4 | N/A
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C1 | 164267549 |  | 2213502602 | C5H9NO4 | 128339.65
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C2 | 110684704 |  | 1450022767 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C3 | 65851902 |  | 903545136 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C4 | 42115786 |  | 578393943 | C5H9NO4 | INF
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C5 | 23976818 |  | 371356114 | C5H9NO4 | 213570.72
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C6 | 13591248 |  | 189240241 | C5H9NO4 | 28930.60
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C7 | 9465444 |  | 122894611 | C5H9NO4 | 13468.67
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C8 | 4961320 |  | 62286239 | C5H9NO4 | 3211.40 | Calibration Amount Diff -23.413% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C9 | 2344186 |  | 31624442 | C5H9NO4 | 1799.87 | Calibration Amount Diff -50.781% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C10 | 973885 |  | 11512080 | C5H9NO4 | 1024.61 | Calibration Amount Diff -133.962% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C11 | 523333 |  | 7039288 | C5H9NO4 | 907.15 | Calibration Amount Diff -233.272% (absolute) > max calibration amount diff 20%
Glutamate pos | 10.05 | Target Compound | KD100520_100520_C12 | 225988 |  | 3150658 | C5H9NO4 | 227.55 | Calibration Amount Diff -454.595% (absolute) > max calibration amount diff 20%

## Glutamate U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A1 | 25386884 |  | 324280633 | [13]C5H9NO4 | 19345.42 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 95.534% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A2 | 29467672 |  | 358599022 | [13]C5H9NO4 | 112898.85 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 195.512% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A3 | 29927233 |  | 389360311 | [13]C5H9NO4 | 45481.56 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 317.925% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A4 | 33378091 |  | 438109566 | [13]C5H9NO4 | 93466.68 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 815.9% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A5 | 38789640 |  | 498494455 | [13]C5H9NO4 | 36802.56 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 2479.257% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A6 | 38000998 |  | 532753001 | [13]C5H9NO4 | 44370.39 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -4196.781% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A7 | 37690982 |  | 533582041 | [13]C5H9NO4 | 14615.68 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -8663.823% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A8 | 42916659 |  | 554710110 | [13]C5H9NO4 | 38546.34 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -10352.342% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A9 | 41989134 |  | 561815668 | [13]C5H9NO4 | 30781.98 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -30122.772% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A10 | 44820666 |  | 575564753 | [13]C5H9NO4 | 57359.88 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -60922.224% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A11 | 39896757 |  | 550191074 | [13]C5H9NO4 | 39701.17 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 193778.779% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_A12 | 39769844 |  | 529747449 | [13]C5H9NO4 | 44189.99 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 210250.816% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B1 | 24716232 |  | 315532755 | [13]C5H9NO4 | INF | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -111.64% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B2 | 28804610 |  | 361401787 | [13]C5H9NO4 | 54322.98 | Rsquared 0.042 < R2Threshold 0.99
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B3 | 33764898 |  | 431897585 | [13]C5H9NO4 | 32686.04 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -387.489% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B4 | 35610165 |  | 479785631 | [13]C5H9NO4 | 24314.01 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 367.824% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B5 | 38645289 |  | 508505372 | [13]C5H9NO4 | 38798.83 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 401.791% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B6 | 39642384 |  | 533196976 | [13]C5H9NO4 | 29784.21 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -1123.586% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B7 | 40434144 |  | 544705804 | [13]C5H9NO4 | 27068.48 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -3415.71% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B8 | 36909927 |  | 516505586 | [13]C5H9NO4 | 50139.57 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -775.594% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B9 | 39400601 |  | 562928834 | [13]C5H9NO4 | 51877.81 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -18115.265% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B10 | 41735929 |  | 573757137 | [13]C5H9NO4 | 31376.96 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -6880.416% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B11 | 40713820 |  | 529377366 | [13]C5H9NO4 | 37645.90 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 133339.39% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_B12 | 41755075 |  | 539956881 | [13]C5H9NO4 | 36705.88 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 388784.601% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H9NO4 | N/A
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C1 | 23987310 |  | 298031673 | [13]C5H9NO4 | 64328.55 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 176.952% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C2 | 28895526 |  | 369740751 | [13]C5H9NO4 | 22221.92 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -507.439% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C3 | 29920115 |  | 412518854 | [13]C5H9NO4 | 24033.57 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -752.209% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C4 | 36967069 |  | 486571839 | [13]C5H9NO4 | 27795.74 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -2162.424% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C5 | 35268281 |  | 482000908 | [13]C5H9NO4 | 39895.94 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -4109.913% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C6 | 34279479 |  | 476196156 | [13]C5H9NO4 | 18960.17 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff -6994.862% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C7 | 38603769 |  | 538882713 | [13]C5H9NO4 | 25231.36 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 7675.501% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C8 | 38013282 |  | 538529452 | [13]C5H9NO4 | 64358.11 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 15947.156% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C9 | 39766047 |  | 543905236 | [13]C5H9NO4 | 37583.67 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 31653.762% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C10 | 38601386 |  | 533255176 | [13]C5H9NO4 | 41954.92 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 65680.092% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C11 | 41229059 |  | 558031206 | [13]C5H9NO4 | 34696.43 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 130797.187% (absolute) > max calibration amount diff 20%
Glutamate U-13C neg | 10.05 | Target Compound | KD100520_100520_C12 | 36454706 |  | 505021087 | [13]C5H9NO4 | 34255.97 | Rsquared 0.042 < R2Threshold 0.99. Calibration Amount Diff 293800.528% (absolute) > max calibration amount diff 20%

## Glutamate U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A1 | 25452640 |  | 349697045 | [13]C5H9NO4 | INF | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -49.249% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A2 | 33819675 |  | 461021984 | [13]C5H9NO4 | 9649.94 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -22.24% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A3 | 39160422 |  | 543494581 | [13]C5H9NO4 | 5312.69 | Rsquared 0.09 < R2Threshold 0.99
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A4 | 54521419 |  | 703675969 | [13]C5H9NO4 | 9812.82 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -451.114% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A5 | 62252198 |  | 803758611 | [13]C5H9NO4 | 13581.18 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -858.792% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A6 | 69568717 |  | 995055890 | [13]C5H9NO4 | 17574.82 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 2558.579% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A7 | 83104388 |  | 1081571069 | [13]C5H9NO4 | 32974.17 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 4731.944% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A8 | 91265512 |  | 1251573154 | [13]C5H9NO4 | 32655.64 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 5230.696% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A9 | 80840056 |  | 1200091115 | [13]C5H9NO4 | 33223.66 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 8787.408% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A10 | 89806137 |  | 1272974508 | [13]C5H9NO4 | 27141.95 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 15525.521% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A11 | 85151854 |  | 1212707751 | [13]C5H9NO4 | 52758.63 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -102486.942% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_A12 | 98215057 |  | 1288070800 | [13]C5H9NO4 | 23229.40 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -225110.868% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B1 | 29274806 |  | 357163916 | [13]C5H9NO4 | INF | Rsquared 0.09 < R2Threshold 0.99
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B2 | 36656233 |  | 490947676 | [13]C5H9NO4 | INF | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 29.715% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B3 | 41032576 |  | 569018801 | [13]C5H9NO4 | 67965.22 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -48.416% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B4 | 55314803 |  | 731495212 | [13]C5H9NO4 | 7805.45 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 362.821% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B5 | 64938817 |  | 915803331 | [13]C5H9NO4 | 10741.76 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 638.988% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B6 | 69263453 |  | 1022937451 | [13]C5H9NO4 | 21752.78 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 527.523% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B7 | 70504034 |  | 1029399395 | [13]C5H9NO4 | 16216.11 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 1642.604% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B8 | 83532562 |  | 1178046967 | [13]C5H9NO4 | 17826.74 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 2258.706% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B9 | 87636302 |  | 1238046049 | [13]C5H9NO4 | 30120.41 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 8583.736% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B10 | 91953176 |  | 1311736261 | [13]C5H9NO4 | 22871.69 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -3565.463% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B11 | 82670022 |  | 1211150366 | [13]C5H9NO4 | 26654.92 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -100800.964% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_B12 | 95383696 |  | 1241579396 | [13]C5H9NO4 | 41623.18 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -243525.211% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H9NO4 | N/A
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C1 | 27989727 |  | 362715237 | [13]C5H9NO4 | 9983.32 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -119.508% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C2 | 37217658 |  | 465211061 | [13]C5H9NO4 | 7842.04 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 309.826% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C3 | 42709410 |  | 580892182 | [13]C5H9NO4 | 7253.14 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 969.274% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C4 | 53096522 |  | 739494615 | [13]C5H9NO4 | 6772.90 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 1272.264% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C5 | 59466705 |  | 923654019 | [13]C5H9NO4 | 9572.37 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 2443.151% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C6 | 69881256 |  | 1000768000 | [13]C5H9NO4 | 8015.92 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff 4887.309% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C7 | 81107544 |  | 1105048097 | [13]C5H9NO4 | 23000.86 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -6799.977% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C8 | 90192050 |  | 1190264858 | [13]C5H9NO4 | 19169.18 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -9507.314% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C9 | 81494565 |  | 1203860472 | [13]C5H9NO4 | 36881.82 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -22429.931% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C10 | 87156927 |  | 1201948653 | [13]C5H9NO4 | 51869.96 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -52937.365% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C11 | 93279207 |  | 1297956931 | [13]C5H9NO4 | 30225.69 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -99611.815% (absolute) > max calibration amount diff 20%
Glutamate U-13C pos | 10.05 | Target Compound | KD100520_100520_C12 | 84585337 |  | 1191463170 | [13]C5H9NO4 | 28136.17 | Rsquared 0.09 < R2Threshold 0.99. Calibration Amount Diff -225839.502% (absolute) > max calibration amount diff 20%

## Glutamate U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A1 | 18914442 |  | 240986600 | [13]C5H9[15]NO4 | 321787.68
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A2 | 21968466 |  | 264284568 | [13]C5H9[15]NO4 | INF
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A3 | 21789763 |  | 284820366 | [13]C5H9[15]NO4 | INF
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A4 | 24571541 |  | 321033146 | [13]C5H9[15]NO4 | 120135.22
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A5 | 28772134 |  | 368244124 | [13]C5H9[15]NO4 | 46438.78
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A6 | 26549730 |  | 374733452 | [13]C5H9[15]NO4 | 24045.13
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A7 | 26757200 |  | 374957436 | [13]C5H9[15]NO4 | 22998.94
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A8 | 30619070 |  | 393294487 | [13]C5H9[15]NO4 | 19186.92
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A9 | 29740371 |  | 395879776 | [13]C5H9[15]NO4 | 12685.78
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A10 | 31593086 |  | 405462037 | [13]C5H9[15]NO4 | 26467.17
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A11 | 29795965 |  | 408420148 | [13]C5H9[15]NO4 | 39274.68
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_A12 | 29139009 |  | 387392843 | [13]C5H9[15]NO4 | 28715.42
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B1 | 17842215 |  | 226296878 | [13]C5H9[15]NO4 | INF
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B2 | 21224782 |  | 262005249 | [13]C5H9[15]NO4 | INF
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B3 | 24150681 |  | 306639279 | [13]C5H9[15]NO4 | 258335.41
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B4 | 25896322 |  | 348213107 | [13]C5H9[15]NO4 | 79203.81
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B5 | 28046070 |  | 367357769 | [13]C5H9[15]NO4 | 63288.62
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B6 | 28593102 |  | 381096420 | [13]C5H9[15]NO4 | 33989.67
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B7 | 28856180 |  | 388029988 | [13]C5H9[15]NO4 | 71077.69
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B8 | 26727533 |  | 370830305 | [13]C5H9[15]NO4 | 30106.12
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B9 | 27862504 |  | 399769412 | [13]C5H9[15]NO4 | 41561.28
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B10 | 29985528 |  | 411381272 | [13]C5H9[15]NO4 | 20867.41
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B11 | 30105960 |  | 388966420 | [13]C5H9[15]NO4 | 20852.24
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_B12 | 30984304 |  | 400861525 | [13]C5H9[15]NO4 | 17754.01
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C5H9[15]NO4 | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C1 | 18061761 |  | 224674955 | [13]C5H9[15]NO4 | INF
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C2 | 20084988 |  | 256857834 | [13]C5H9[15]NO4 | 160162.68
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C3 | 21002594 |  | 288491267 | [13]C5H9[15]NO4 | 276647.49
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C4 | 25495348 |  | 335061254 | [13]C5H9[15]NO4 | 146828.75
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C5 | 24438293 |  | 332299980 | [13]C5H9[15]NO4 | 19973.82
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C6 | 23710595 |  | 330178927 | [13]C5H9[15]NO4 | 38907.56
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C7 | 28470000 |  | 395354074 | [13]C5H9[15]NO4 | 17720.16
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C8 | 27898843 |  | 395358079 | [13]C5H9[15]NO4 | 28838.44
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C9 | 29419417 |  | 399213049 | [13]C5H9[15]NO4 | 13560.30
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C10 | 28319173 |  | 391700520 | [13]C5H9[15]NO4 | 20759.55
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C11 | 30411041 |  | 409844561 | [13]C5H9[15]NO4 | 20290.80
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD100520_100520_C12 | 26735384 |  | 371917302 | [13]C5H9[15]NO4 | 20188.87

## Glutamate U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A1 | 19014866 |  | 251365100 | [13]C5H9[15]NO4 | 18521.67 | 10.090:Peak area 251365099.613 is out of bounds (ISTD Minimum Recovery 342079114.772 and ISTD Max Recovery 1026237344.317)
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A2 | 25652329 |  | 332551222 | [13]C5H9[15]NO4 | 60299.23 | 10.074:Peak area 332551221.761 is out of bounds (ISTD Minimum Recovery 342079114.772 and ISTD Max Recovery 1026237344.317)
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A3 | 28248749 |  | 393235324 | [13]C5H9[15]NO4 | 6236.17
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A4 | 40314020 |  | 520338587 | [13]C5H9[15]NO4 | 13852.52
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A5 | 45864946 |  | 594983934 | [13]C5H9[15]NO4 | 15098.29
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A6 | 49748920 |  | 708495660 | [13]C5H9[15]NO4 | 17789.07
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A7 | 59879835 |  | 771805003 | [13]C5H9[15]NO4 | 6873.35
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A8 | 65701035 |  | 902065143 | [13]C5H9[15]NO4 | 36273.81
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A9 | 58014723 |  | 866736598 | [13]C5H9[15]NO4 | 6818.98
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A10 | 64719520 |  | 920516130 | [13]C5H9[15]NO4 | 7560.74
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A11 | 64367678 |  | 912264261 | [13]C5H9[15]NO4 | 11200.75
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_A12 | 74392759 |  | 971920348 | [13]C5H9[15]NO4 | 40262.31
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B1 | 21345901 |  | 253282229 | [13]C5H9[15]NO4 | INF | 10.103:Peak area 253282228.810 is out of bounds (ISTD Minimum Recovery 342079114.772 and ISTD Max Recovery 1026237344.317)
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B2 | 26381297 |  | 351434348 | [13]C5H9[15]NO4 | 197680.97
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B3 | 30329213 |  | 413635584 | [13]C5H9[15]NO4 | 6734.21
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B4 | 39386178 |  | 524705215 | [13]C5H9[15]NO4 | 8981.15
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B5 | 46592249 |  | 659174638 | [13]C5H9[15]NO4 | 10425.09
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B6 | 50227897 |  | 742114800 | [13]C5H9[15]NO4 | 15339.18
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B7 | 51227481 |  | 745111959 | [13]C5H9[15]NO4 | 7267.25
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B8 | 60381151 |  | 854944188 | [13]C5H9[15]NO4 | 8393.80
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B9 | 63843056 |  | 894359564 | [13]C5H9[15]NO4 | 11467.94
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B10 | 68034883 |  | 959166628 | [13]C5H9[15]NO4 | 12027.51
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B11 | 62775695 |  | 910629995 | [13]C5H9[15]NO4 | 13138.05
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_B12 | 72925056 |  | 939454894 | [13]C5H9[15]NO4 | 13615.48
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C5H9[15]NO4 | N/A
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C1 | 20611282 |  | 266242200 | [13]C5H9[15]NO4 | INF | 10.049:Peak area 266242199.749 is out of bounds (ISTD Minimum Recovery 342079114.772 and ISTD Max Recovery 1026237344.317)
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C2 | 25212534 |  | 319848211 | [13]C5H9[15]NO4 | 8661.71 | 10.088:Peak area 319848210.887 is out of bounds (ISTD Minimum Recovery 342079114.772 and ISTD Max Recovery 1026237344.317)
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C3 | 29227341 |  | 392470408 | [13]C5H9[15]NO4 | 21956.60
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C4 | 37142132 |  | 513262702 | [13]C5H9[15]NO4 | 35957.83
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C5 | 42030430 |  | 643389792 | [13]C5H9[15]NO4 | 14623.08
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C6 | 48740774 |  | 697722265 | [13]C5H9[15]NO4 | 9188.31
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C7 | 62898769 |  | 832501600 | [13]C5H9[15]NO4 | 21473.95
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C8 | 67528643 |  | 888130381 | [13]C5H9[15]NO4 | 15929.29
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C9 | 60765494 |  | 902068717 | [13]C5H9[15]NO4 | 10329.83
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C10 | 66021852 |  | 905097074 | [13]C5H9[15]NO4 | 11374.66
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C11 | 70085199 |  | 975547966 | [13]C5H9[15]NO4 | 9801.32
Glutamate U-13C, U-15N pos | 10.05 | Internal Standard | KD100520_100520_C12 | 64032811 |  | 899123596 | [13]C5H9[15]NO4 | 5541.51

## Glutamine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A1 | 63897771 |  | 928454969 | C5H10N2O3 | INF | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -24.924% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A2 | 38796218 |  | 536965753 | C5H10N2O3 | 72795.93 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -29.644% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A3 | 25529922 |  | 383328999 | C5H10N2O3 | 104584.09 | Rsquared 0.84 < R2Threshold 0.99
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A4 | 14966577 |  | 222616256 | C5H10N2O3 | 45691.97 | Rsquared 0.84 < R2Threshold 0.99
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A5 | 11043301 |  | 158575379 | C5H10N2O3 | 11883.81 | Rsquared 0.84 < R2Threshold 0.99
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A6 | 6329095 |  | 85879610 | C5H10N2O3 | 8520.15 | Rsquared 0.84 < R2Threshold 0.99
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A7 | 2996653 |  | 42935280 | C5H10N2O3 | 5682.85 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -32.551% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A8 | 1499080 |  | 23780993 | C5H10N2O3 | 3830.59 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -59.684% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A9 | 957240 |  | 12988656 | C5H10N2O3 | 2211.77 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -109.601% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A10 | 388380 |  | 5170754 | C5H10N2O3 | 1985.05 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -229.856% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A11 | 241087 |  | 3034836 | C5H10N2O3 | 624.94 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -445.243% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_A12 | 91754 |  | 1118903 | C5H10N2O3 | 240.95 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -902.373% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B1 | 118113954 |  | 1695940012 | C5H10N2O3 | INF | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff 55.857% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B2 | 65824934 |  | 984308747 | C5H10N2O3 | INF | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff 40.492% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B3 | 44545797 |  | 673098761 | C5H10N2O3 | 132062.67 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff 56.919% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B4 | 30053478 |  | 421610498 | C5H10N2O3 | 37725.34 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff 66.268% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B5 | 16213714 |  | 242101005 | C5H10N2O3 | 22903.63 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff 50.581% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B6 | 8981668 |  | 135093203 | C5H10N2O3 | 9947.15 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff 59.377% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B7 | 4938302 |  | 76544526 | C5H10N2O3 | 8505.80 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff 46.004% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B8 | 2909317 |  | 43648524 | C5H10N2O3 | 5204.17 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff 30.063% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B9 | 1732269 |  | 23926254 | C5H10N2O3 | 3856.16 | Rsquared 0.84 < R2Threshold 0.99
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B10 | 771344 |  | 11013457 | C5H10N2O3 | 1219.26 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -145.37% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B11 | 448266 |  | 5951996 | C5H10N2O3 | 1267.76 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -345.427% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_B12 | 183404 |  | 2356188 | C5H10N2O3 | 443.48 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -819.387% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H10N2O3 | N/A
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C1 | 50992954 |  | 755177283 | C5H10N2O3 | INF | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -29.474% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C2 | 33773855 |  | 498563893 | C5H10N2O3 | 76110.27 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -23.754% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C3 | 20845945 |  | 307551396 | C5H10N2O3 | 50346.22 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -23.25% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C4 | 11941994 |  | 179999537 | C5H10N2O3 | 16852.35 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -23.148% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C5 | 7616876 |  | 114397271 | C5H10N2O3 | 10107.87 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -22.579% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C6 | 3944562 |  | 58576534 | C5H10N2O3 | 4344.31 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -34.537% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C7 | 2398734 |  | 37026703 | C5H10N2O3 | 2428.17 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -40.551% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C8 | 1298615 |  | 18495334 | C5H10N2O3 | 1867.25 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -72.417% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C9 | 787588 |  | 10487211 | C5H10N2O3 | 1479.58 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -122.514% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C10 | 286660 |  | 3770610 | C5H10N2O3 | 793.83 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -253.342% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C11 | 206732 |  | 2293729 | C5H10N2O3 | 852.42 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -463.942% (absolute) > max calibration amount diff 20%
Glutamine neg | 9.64 | Target Compound | KD100520_100520_C12 | 53555 |  | 875228 | C5H10N2O3 | 230.71 | Rsquared 0.84 < R2Threshold 0.99. Calibration Amount Diff -906.74% (absolute) > max calibration amount diff 20%

## Glutamine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A1 | 664081342 |  | 9759949987 | C5H10N2O3 | INF | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -24.019% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A2 | 462075111 |  | 6503979990 | C5H10N2O3 | INF | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -28.274% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A3 | 315626125 |  | 4756187457 | C5H10N2O3 | 120506.53 | Rsquared 0.846 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A4 | 197777803 |  | 2815090907 | C5H10N2O3 | 48500.49 | Rsquared 0.846 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A5 | 129232244 |  | 1861707230 | C5H10N2O3 | 30045.25 | Rsquared 0.846 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A6 | 67677925 |  | 966156966 | C5H10N2O3 | 28671.84 | Rsquared 0.846 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A7 | 29893883 |  | 444498438 | C5H10N2O3 | 15615.91 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -41.489% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A8 | 15029277 |  | 234824534 | C5H10N2O3 | 10151.13 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -84.246% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A9 | 8960998 |  | 131396337 | C5H10N2O3 | 3739.31 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -157.226% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A10 | 3760929 |  | 56792269 | C5H10N2O3 | 1826.33 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -335.486% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A11 | 2257027 |  | 31921390 | C5H10N2O3 | 2502.36 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -659.528% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_A12 | 981426 |  | 13851501 | C5H10N2O3 | 625.27 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -1345.578% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B1 | 1158167898 |  | 16666991916 | C5H10N2O3 | 202294.83 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff 54.008% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B2 | 800430649 |  | 11748563972 | C5H10N2O3 | 6866959.04 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff 40.083% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B3 | 536675750 |  | 8045181041 | C5H10N2O3 | 104202.72 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff 57.552% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B4 | 318418973 |  | 4992039526 | C5H10N2O3 | 88759.53 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff 69.095% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B5 | 172336515 |  | 2617488049 | C5H10N2O3 | 32229.95 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff 54.016% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B6 | 114023926 |  | 1603258248 | C5H10N2O3 | 42290.35 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff 61.434% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B7 | 54729157 |  | 838356142 | C5H10N2O3 | 18717.36 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff 44.66% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B8 | 27885417 |  | 457587333 | C5H10N2O3 | 14551.42 | Rsquared 0.846 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B9 | 16395002 |  | 239950472 | C5H10N2O3 | 7967.54 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -58.1% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B10 | 7634294 |  | 112450276 | C5H10N2O3 | 3156.33 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -242.667% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B11 | 3923172 |  | 58526679 | C5H10N2O3 | 1723.62 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -558.896% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_B12 | 1667120 |  | 24847765 | C5H10N2O3 | 899.44 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -1254.83% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H10N2O3 | N/A
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C1 | 704414871 |  | 9577020619 | C5H10N2O3 | 80545.00 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -29.743% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C2 | 474781852 |  | 6730875502 | C5H10N2O3 | INF | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -21.969% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C3 | 292210669 |  | 4270419392 | C5H10N2O3 | 83675.59 | Rsquared 0.846 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C4 | 167557279 |  | 2430994815 | C5H10N2O3 | 25734.22 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -20.209% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C5 | 94247337 |  | 1415283401 | C5H10N2O3 | 10389.14 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -20.013% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C6 | 47469161 |  | 695794575 | C5H10N2O3 | 14351.98 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -36.355% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C7 | 26000679 |  | 415738035 | C5H10N2O3 | 12551.64 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -48.227% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C8 | 13136671 |  | 195242014 | C5H10N2O3 | 4332.28 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -95.16% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C9 | 7061142 |  | 108318924 | C5H10N2O3 | 3119.30 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -175.712% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C10 | 2890883 |  | 44407295 | C5H10N2O3 | 1159.48 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -355.493% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C11 | 1689508 |  | 25451955 | C5H10N2O3 | 1035.30 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -684.365% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD100520_100520_C12 | 685806 |  | 9438004 | C5H10N2O3 | 283.48 | Rsquared 0.846 < R2Threshold 0.99. Calibration Amount Diff -1366.792% (absolute) > max calibration amount diff 20%

## Gln_pos_calibration
Compound | Filename | Sample ID | Area | Area_IS | Response ratio | Conc. (Asn/Gln in cali B is double, hence correction) | Conc. Lin reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss | Conc. power reg. | Lin. Reg. l1 loss | Lin. Reg. c1 loss | Conc. power reg. no norm | Lin. Reg. l1 loss | Lin. Reg. c1 loss
Glutamine pos | KD100520_100520_A1 |  | 9759949987 | 1540320659 | 6.3363 | 434.3497 | 437.2498 | 2.9001 | 0.66769 | 401.8169 | 32.5328 | 7.49 | 309.8912 | 124.4585 | 28.654
Glutamine pos | KD100520_100520_A2 |  | 6503979990 | 2154226812 | 3.0192 | 217.1748 | 208.344 | 8.8309 | 4.0663 | 188.3774 | 28.7974 | 13.26 | 196.6769 | 20.498 | 9.4385
Glutamine pos | KD100520_100520_A3 |  | 4756187457 | 2552027562 | 1.8637 | 108.5874 | 128.6076 | 20.0202 | 18.437 | 115.0605 | 6.4731 | 5.9612 | 138.5146 | 29.9272 | 27.5604
Glutamine pos | KD100520_100520_A4 |  | 2815090907 | 2990411855 | 0.94137 | 54.2937 | 64.9613 | 10.6676 | 19.6479 | 57.2556 | 2.9619 | 5.4554 | 76.9753 | 22.6816 | 41.7757
Glutamine pos | KD100520_100520_A5 |  | 1861707230 | 2995894522 | 0.62142 | 27.1469 | 42.8823 | 15.7354 | 57.9641 | 37.4534 | 10.3066 | 37.966 | 48.4378 | 21.291 | 78.4289
Glutamine pos | KD100520_100520_A6 |  | 966156966 | 3369178088 | 0.28676 | 13.5734 | 19.7887 | 6.2153 | 45.7899 | 16.9932 | 3.4198 | 25.1946 | 23.2317 | 9.6582 | 71.1555
Glutamine pos | KD100520_100520_A7 |  | 444498438 | 3432009308 | 0.12952 | 6.7867 | 8.9375 | 2.1508 | 31.6908 | 7.5425 | 0.75575 | 11.1357 | 9.7359 | 2.9491 | 43.4547
Glutamine pos | KD100520_100520_A8 |  | 234824534 | 3663757007 | 0.064094 | 3.3934 | 4.4229 | 1.0296 | 30.3408 | 3.6755 | 0.28215 | 8.3149 | 4.7636 | 1.3703 | 40.3807
Glutamine pos | KD100520_100520_A9 |  | 131396337 | 3707968263 | 0.035436 | 1.6967 | 2.4453 | 0.74867 | 44.1255 | 2.0059 | 0.30923 | 18.2258 | 2.4858 | 0.78913 | 46.51
Glutamine pos | KD100520_100520_A10 |  | 56792269 | 3578094435 | 0.015872 | 0.84834 | 1.0953 | 0.24695 | 29.1103 | 0.8828 | 0.034462 | 4.0623 | 0.97137 | 0.12303 | 14.5026
Glutamine pos | KD100520_100520_A11 |  | 31921390 | 3642009208 | 0.0087648 | 0.42417 | 0.60483 | 0.18066 | 42.5917 | 0.48119 | 0.057022 | 13.4433 | 0.50945 | 0.085281 | 20.1055
Glutamine pos | KD100520_100520_A12 |  | 13851501 | 3784889414 | 0.0036597 | 0.21208 | 0.25254 | 0.040459 | 19.0768 | 0.19711 | 0.014972 | 7.0595 | 0.19996 | 0.012129 | 5.7188
Glutamine pos | KD100520_100520_B1 |  | 16666991916 | 1303340185 | 12.7879 | 868.6994 | 882.455 | 13.7557 | 1.5835 | 823.5119 | 45.1874 | 5.2017 | 564.3577 | 304.3417 | 35.0342
Glutamine pos | KD100520_100520_B2 |  | 11748563972 | 2009943876 | 5.8452 | 434.3497 | 403.3611 | 30.9886 | 7.1345 | 370.0201 | 64.3296 | 14.8105 | 381.4406 | 52.9091 | 12.1812
Glutamine pos | KD100520_100520_B3 |  | 8045181041 | 2430064596 | 3.3107 | 217.1748 | 228.4605 | 11.2857 | 5.1966 | 206.9835 | 10.1913 | 4.6927 | 249.5808 | 32.406 | 14.9216
Glutamine pos | KD100520_100520_B4 |  | 4992039526 | 2770891557 | 1.8016 | 108.5874 | 124.323 | 15.7356 | 14.4912 | 111.1448 | 2.5573 | 2.3551 | 146.2315 | 37.6441 | 34.6671
Glutamine pos | KD100520_100520_B5 |  | 2617488049 | 3079997681 | 0.84983 | 54.2937 | 58.6445 | 4.3508 | 8.0135 | 51.5725 | 2.7212 | 5.012 | 70.9487 | 16.655 | 30.6757
Glutamine pos | KD100520_100520_B6 |  | 1603258248 | 3403627578 | 0.47104 | 27.1469 | 32.5053 | 5.3585 | 19.7388 | 28.2184 | 1.0716 | 3.9474 | 40.9708 | 13.824 | 50.923
Glutamine pos | KD100520_100520_B7 |  | 838356142 | 3481223643 | 0.24082 | 13.5734 | 16.6184 | 3.045 | 22.4335 | 14.2163 | 0.6429 | 4.7364 | 19.8178 | 6.2443 | 46.0041
Glutamine pos | KD100520_100520_B8 |  | 457587333 | 3534442295 | 0.12947 | 6.7867 | 8.934 | 2.1473 | 31.6396 | 7.5395 | 0.75276 | 11.0916 | 10.0576 | 3.2709 | 48.195
Glutamine pos | KD100520_100520_B9 |  | 239950472 | 3557439817 | 0.06745 | 3.3934 | 4.6545 | 1.2612 | 37.1664 | 3.8723 | 0.47896 | 14.1145 | 4.8803 | 1.4869 | 43.8178
Glutamine pos | KD100520_100520_B10 |  | 112450276 | 3642888299 | 0.030868 | 1.6967 | 2.1301 | 0.43346 | 25.5476 | 1.7421 | 0.045396 | 2.6756 | 2.0879 | 0.39125 | 23.0599
Glutamine pos | KD100520_100520_B11 |  | 58526679 | 3466052470 | 0.016886 | 0.84834 | 1.1652 | 0.31689 | 37.3543 | 0.94044 | 0.092105 | 10.8571 | 1.0047 | 0.15632 | 18.4269
Glutamine pos | KD100520_100520_B12 |  | 24847765 | 3393888373 | 0.0073213 | 0.42417 | 0.50522 | 0.081053 | 19.1087 | 0.40036 | 0.023805 | 5.6121 | 0.3848 | 0.039373 | 9.2825
Glutamine pos | KD100520_100520_C1 |  | 9577020619 | 1633466978 | 5.863 | 434.3497 | 404.5882 | 29.7615 | 6.852 | 371.1705 | 63.1792 | 14.5457 | 303.3921 | 130.9576 | 30.1503
Glutamine pos | KD100520_100520_C2 |  | 6730875502 | 2052188458 | 3.2799 | 217.1748 | 226.3328 | 9.1579 | 4.2168 | 205.0138 | 12.1611 | 5.5997 | 204.3788 | 12.7961 | 5.8921
Glutamine pos | KD100520_100520_C3 |  | 4270419392 | 2486439676 | 1.7175 | 108.5874 | 118.5184 | 9.931 | 9.1456 | 105.8445 | 2.7429 | 2.526 | 122.7674 | 14.18 | 13.0586
Glutamine pos | KD100520_100520_C4 |  | 2430994815 | 2766917688 | 0.87859 | 54.2937 | 60.6291 | 6.3354 | 11.6687 | 53.3566 | 0.93709 | 1.726 | 65.3108 | 11.0171 | 20.2917
Glutamine pos | KD100520_100520_C5 |  | 1415283401 | 3028841195 | 0.46727 | 27.1469 | 32.2448 | 5.098 | 18.7792 | 27.9874 | 0.84051 | 3.0961 | 35.6291 | 8.4822 | 31.2457
Glutamine pos | KD100520_100520_C6 |  | 695794575 | 3186345452 | 0.21837 | 13.5734 | 15.0689 | 1.4955 | 11.0176 | 12.8632 | 0.71026 | 5.2327 | 16.0834 | 2.5099 | 18.4916
Glutamine pos | KD100520_100520_C7 |  | 415738035 | 3441277804 | 0.12081 | 6.7867 | 8.3367 | 1.55 | 22.8383 | 7.0247 | 0.23801 | 3.5071 | 9.033 | 2.2463 | 33.0982
Glutamine pos | KD100520_100520_C8 |  | 195242014 | 3422630562 | 0.057044 | 3.3934 | 3.9365 | 0.54311 | 16.005 | 3.2629 | 0.13044 | 3.844 | 3.8737 | 0.48038 | 14.1565
Glutamine pos | KD100520_100520_C9 |  | 108318924 | 3676046453 | 0.029466 | 1.6967 | 2.0334 | 0.33669 | 19.8442 | 1.6612 | 0.035436 | 2.0885 | 2.0022 | 0.30551 | 18.0066
Glutamine pos | KD100520_100520_C10 |  | 44407295 | 3513285586 | 0.01264 | 0.84834 | 0.87224 | 0.023897 | 2.8169 | 0.69952 | 0.14882 | 17.5423 | 0.73741 | 0.11093 | 13.0762
Glutamine pos | KD100520_100520_C11 |  | 25451955 | 3764856051 | 0.0067604 | 0.42417 | 0.46652 | 0.042346 | 9.9832 | 0.36905 | 0.055123 | 12.9956 | 0.39529 | 0.028877 | 6.8079
Glutamine pos | KD100520_100520_C12 |  | 9438004 | 3366291039 | 0.0028037 | 0.21208 | 0.19347 | 0.018611 | 8.7754 | 0.15013 | 0.061956 | 29.2126 | 0.1301 | 0.081981 | 38.6547
[5 empty] | 1 | 50 | 69.007 |  |  | 60.902
[6 empty] | Mean |  |  | 19.8572 |  |  | 9.572
[6 empty] | Sum |  | 221.8201 |  |  | 295.2803
Conclusion: Power regression is best over a wide range
|  |  | Conc. | Resi_before | Resi_after
|  |  | 434.3497 | 28.654 | 7.49
|  |  | 217.1748 | 9.4385 | 13.26
|  |  | 108.5874 | 27.5604 | 5.9612
|  |  | 54.2937 | 41.7757 | 5.4554
|  |  | 27.1469 | 78.4289 | 37.966
|  |  | 13.5734 | 71.1555 | 25.1946
|  |  | 6.7867 | 43.4547 | 11.1357
|  |  | 3.3934 | 40.3807 | 8.3149
|  |  | 1.6967 | 46.51 | 18.2258
|  |  | 0.84834 | 14.5026 | 4.0623
|  |  | 0.42417 | 20.1055 | 13.4433
|  |  | 0.21208 | 5.7188 | 7.0595
|  |  | 868.6994 | 35.0342 | 5.2017
|  |  | 434.3497 | 12.1812 | 14.8105
|  |  | 217.1748 | 14.9216 | 4.6927
|  |  | 108.5874 | 34.6671 | 2.3551
|  |  | 54.2937 | 30.6757 | 5.012
|  |  | 27.1469 | 50.923 | 3.9474
|  |  | 13.5734 | 46.0041 | 4.7364
|  |  | 6.7867 | 48.195 | 11.0916
|  |  | 3.3934 | 43.8178 | 14.1145
|  |  | 1.6967 | 23.0599 | 2.6756
|  |  | 0.84834 | 18.4269 | 10.8571
|  |  | 0.42417 | 9.2825 | 5.6121
|  |  | 434.3497 | 30.1503 | 14.5457
|  |  | 217.1748 | 5.8921 | 5.5997
|  |  | 108.5874 | 13.0586 | 2.526
|  |  | 54.2937 | 20.2917 | 1.726
|  |  | 27.1469 | 31.2457 | 3.0961
|  |  | 13.5734 | 18.4916 | 5.2327
|  |  | 6.7867 | 33.0982 | 3.5071
|  |  | 3.3934 | 14.1565 | 3.844
|  |  | 1.6967 | 18.0066 | 2.0885
|  |  | 0.84834 | 13.0762 | 17.5423
|  |  | 0.42417 | 6.8079 | 12.9956
|  |  | 0.21208 | 38.6547 | 29.2126

## Glutamine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A1 | 3591541 |  | 49012932 | [13]C5H10N2O3 | 6519.06 | Rsquared 0.39 < R2Threshold 0.99
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A2 | 4451713 |  | 60351303 | [13]C5H10N2O3 | 9691.36 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 79.511% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A3 | 4863237 |  | 70790808 | [13]C5H10N2O3 | 9284.12 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 197.444% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A4 | 5615543 |  | 83745363 | [13]C5H10N2O3 | 9528.76 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 390.815% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A5 | 6330650 |  | 90793142 | [13]C5H10N2O3 | 7297.91 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 799.547% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A6 | 8483349 |  | 114127877 | [13]C5H10N2O3 | 11598.33 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -1839.454% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A7 | 8631082 |  | 126003488 | [13]C5H10N2O3 | 18116.19 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -2450.982% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A8 | 8534165 |  | 140606191 | [13]C5H10N2O3 | 15062.36 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -6325.604% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A9 | 10195007 |  | 146206959 | [13]C5H10N2O3 | 13677.09 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -10760.528% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A10 | 8740392 |  | 133023170 | [13]C5H10N2O3 | 10603.32 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -27037.468% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A11 | 10004403 |  | 144030919 | [13]C5H10N2O3 | 21172.80 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 42325.916% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_A12 | 9543804 |  | 142190950 | [13]C5H10N2O3 | 14103.46 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 92614.39% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B1 | 3275603 |  | 43956665 | [13]C5H10N2O3 | INF | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -28.776% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B2 | 4054012 |  | 56718569 | [13]C5H10N2O3 | 12726.38 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 26.2% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B3 | 4738380 |  | 70275585 | [13]C5H10N2O3 | 13196.20 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 50.123% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B4 | 5979659 |  | 83055799 | [13]C5H10N2O3 | 13155.73 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 86.586% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B5 | 6984887 |  | 104016356 | [13]C5H10N2O3 | 16749.36 | Rsquared 0.39 < R2Threshold 0.99
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B6 | 6898992 |  | 105129562 | [13]C5H10N2O3 | 6792.44 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 309.059% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B7 | 7758285 |  | 120741861 | [13]C5H10N2O3 | 14774.04 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -770.573% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B8 | 8559533 |  | 129012251 | [13]C5H10N2O3 | 11063.97 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -898.36% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B9 | 9646548 |  | 138193164 | [13]C5H10N2O3 | 17239.88 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -3307.111% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B10 | 9335975 |  | 140354503 | [13]C5H10N2O3 | 24834.41 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -6566.628% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B11 | 9251997 |  | 134189416 | [13]C5H10N2O3 | 20684.89 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 31365.501% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_B12 | 9012703 |  | 133896182 | [13]C5H10N2O3 | 15735.21 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 74142.362% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_blank | 1473 |  | 7265 | [13]C5H10N2O3 | INF
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C1 | 3096796 |  | 41377694 | [13]C5H10N2O3 | INF | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 32.67% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C2 | 3909588 |  | 54964238 | [13]C5H10N2O3 | 23906.39 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -107.386% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C3 | 4627619 |  | 67518688 | [13]C5H10N2O3 | 176759.60 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -198.857% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C4 | 5229927 |  | 77656887 | [13]C5H10N2O3 | 9029.29 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -364.046% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C5 | 6257869 |  | 94874760 | [13]C5H10N2O3 | 11204.74 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -852.569% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C6 | 7001363 |  | 103697831 | [13]C5H10N2O3 | 12509.18 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff -1694.514% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C7 | 6956616 |  | 113709551 | [13]C5H10N2O3 | 11926.38 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 1946.742% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C8 | 8087076 |  | 120053468 | [13]C5H10N2O3 | 11696.44 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 4193.661% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C9 | 9168123 |  | 130646822 | [13]C5H10N2O3 | 15790.80 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 5059.443% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C10 | 8617223 |  | 126883174 | [13]C5H10N2O3 | 14448.11 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 10601.642% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C11 | 9865924 |  | 139993425 | [13]C5H10N2O3 | 8199.33 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 16357.543% (absolute) > max calibration amount diff 20%
Glutamine U-13C neg | 9.64 | Target Compound | KD100520_100520_C12 | 7596808 |  | 119880015 | [13]C5H10N2O3 | 12485.06 | Rsquared 0.39 < R2Threshold 0.99. Calibration Amount Diff 70453.602% (absolute) > max calibration amount diff 20%

## Glutamine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A1 | 39390421 |  | 559943878 | [13]C5H10N2O3 | INF | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 78.943% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A2 | 57781862 |  | 789849530 | [13]C5H10N2O3 | INF | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 181.118% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A3 | 63706266 |  | 946331165 | [13]C5H10N2O3 | 293643.51 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 257.752% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A4 | 77759237 |  | 1102686995 | [13]C5H10N2O3 | 132535.29 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 819.242% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A5 | 76424073 |  | 1101906123 | [13]C5H10N2O3 | 1498797.36 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 1922.224% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A6 | 90538191 |  | 1305568504 | [13]C5H10N2O3 | 53012.72 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -3793.553% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A7 | 86489848 |  | 1334102205 | [13]C5H10N2O3 | 51314.17 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -8444.92% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A8 | 88473434 |  | 1415347797 | [13]C5H10N2O3 | 82569.90 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -13000.367% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A9 | 96076845 |  | 1437023371 | [13]C5H10N2O3 | 72979.56 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -29796.501% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A10 | 85934109 |  | 1382815157 | [13]C5H10N2O3 | 82847.62 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -52657.804% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A11 | 85584404 |  | 1334182578 | [13]C5H10N2O3 | 101260.51 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 147935.436% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_A12 | 87769987 |  | 1395609922 | [13]C5H10N2O3 | 73321.85 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 235583.759% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B1 | 35414107 |  | 483820811 | [13]C5H10N2O3 | INF | Rsquared 0.057 < R2Threshold 0.99
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B2 | 52771464 |  | 747899859 | [13]C5H10N2O3 | INF | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 47.34% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B3 | 62800245 |  | 912368079 | [13]C5H10N2O3 | INF | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 30.171% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B4 | 68468698 |  | 1053823007 | [13]C5H10N2O3 | 145849.82 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -317.824% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B5 | 77540831 |  | 1160799984 | [13]C5H10N2O3 | 27796.84 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 139.173% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B6 | 91714615 |  | 1291034547 | [13]C5H10N2O3 | 37757.80 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -575.419% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B7 | 84322179 |  | 1318179713 | [13]C5H10N2O3 | 37366.50 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -534.467% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B8 | 80986454 |  | 1348342044 | [13]C5H10N2O3 | 92258.52 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -5419.879% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B9 | 90370697 |  | 1358831510 | [13]C5H10N2O3 | 63001.90 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -12255.959% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B10 | 90745912 |  | 1376731472 | [13]C5H10N2O3 | 54162.34 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 1018.537% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B11 | 81214296 |  | 1270745575 | [13]C5H10N2O3 | 88414.45 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 144228.621% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_B12 | 77457757 |  | 1261688968 | [13]C5H10N2O3 | 75546.24 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 159583.713% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H10N2O3 | N/A
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C1 | 45870639 |  | 596115416 | [13]C5H10N2O3 | INF | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 61.579% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C2 | 58539439 |  | 804569986 | [13]C5H10N2O3 | INF | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -442.583% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C3 | 67893902 |  | 978592519 | [13]C5H10N2O3 | 24339.67 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -859.665% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C4 | 75948098 |  | 1095847973 | [13]C5H10N2O3 | 24545.29 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -1863.071% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C5 | 78546106 |  | 1195950987 | [13]C5H10N2O3 | 23478.59 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -3390.567% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C6 | 83407313 |  | 1245262630 | [13]C5H10N2O3 | 36365.92 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff -5093.275% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C7 | 78581749 |  | 1267044673 | [13]C5H10N2O3 | 27053.67 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 7686.082% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C8 | 81939852 |  | 1262904062 | [13]C5H10N2O3 | 40426.76 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 14222.988% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C9 | 86423382 |  | 1347833412 | [13]C5H10N2O3 | 98529.89 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 35879.429% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C10 | 85137279 |  | 1291660041 | [13]C5H10N2O3 | 99704.37 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 65557.294% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C11 | 85936736 |  | 1391515526 | [13]C5H10N2O3 | 66050.42 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 106742.037% (absolute) > max calibration amount diff 20%
Glutamine U-13C pos | 9.64 | Target Compound | KD100520_100520_C12 | 77253939 |  | 1233936301 | [13]C5H10N2O3 | 74077.34 | Rsquared 0.057 < R2Threshold 0.99. Calibration Amount Diff 290306.781% (absolute) > max calibration amount diff 20%

## Glutamine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A1 | 9710717 |  | 136946395 | [13]C5H10[15]N2O3 | INF | 9.686:Peak area 136946395.327 is out of bounds (ISTD Minimum Recovery 137753067.652 and ISTD Max Recovery 413259202.956)
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A2 | 12405609 |  | 167961214 | [13]C5H10[15]N2O3 | 291663.46
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A3 | 13607689 |  | 195022566 | [13]C5H10[15]N2O3 | 62107.10
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A4 | 15415085 |  | 228756358 | [13]C5H10[15]N2O3 | 153171.06
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A5 | 17156398 |  | 247181913 | [13]C5H10[15]N2O3 | 99800.90
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A6 | 21752670 |  | 289899721 | [13]C5H10[15]N2O3 | 35923.86
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A7 | 21312301 |  | 323518280 | [13]C5H10[15]N2O3 | 49997.79
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A8 | 22168965 |  | 358400557 | [13]C5H10[15]N2O3 | 21941.31
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A9 | 26287026 |  | 374267805 | [13]C5H10[15]N2O3 | 51331.14
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A10 | 21907250 |  | 338249055 | [13]C5H10[15]N2O3 | 45888.08
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A11 | 27184941 |  | 388385318 | [13]C5H10[15]N2O3 | 18754.68
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_A12 | 25474727 |  | 384372044 | [13]C5H10[15]N2O3 | 94176.20
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B1 | 8645001 |  | 120847728 | [13]C5H10[15]N2O3 | 427161.89 | 9.671:Peak area 120847728.301 is out of bounds (ISTD Minimum Recovery 137753067.652 and ISTD Max Recovery 413259202.956)
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B2 | 10734980 |  | 155107378 | [13]C5H10[15]N2O3 | 149281.73
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B3 | 12751997 |  | 189029808 | [13]C5H10[15]N2O3 | 50728.01
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B4 | 15719928 |  | 221389083 | [13]C5H10[15]N2O3 | 58812.70
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B5 | 17892521 |  | 274122798 | [13]C5H10[15]N2O3 | 82896.16
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B6 | 18060782 |  | 278362888 | [13]C5H10[15]N2O3 | 173920.01
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B7 | 19551312 |  | 315073938 | [13]C5H10[15]N2O3 | 37718.31
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B8 | 21997640 |  | 337546246 | [13]C5H10[15]N2O3 | 43384.74
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B9 | 25315981 |  | 360153746 | [13]C5H10[15]N2O3 | 41417.88
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B10 | 23462269 |  | 365761811 | [13]C5H10[15]N2O3 | 53949.12
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B11 | 24773678 |  | 359375429 | [13]C5H10[15]N2O3 | 34293.05
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_B12 | 23762416 |  | 359858612 | [13]C5H10[15]N2O3 | 51975.62
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C5H10[15]N2O3 | N/A
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C1 | 8327068 |  | 118532178 | [13]C5H10[15]N2O3 | 175850.37 | 9.673:Peak area 118532177.868 is out of bounds (ISTD Minimum Recovery 137753067.652 and ISTD Max Recovery 413259202.956)
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C2 | 9869036 |  | 144036388 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C3 | 11984332 |  | 174641935 | [13]C5H10[15]N2O3 | 91828.46
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C4 | 13495499 |  | 199843162 | [13]C5H10[15]N2O3 | 195954.48
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C5 | 16290363 |  | 242067879 | [13]C5H10[15]N2O3 | 104159.16
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C6 | 18306232 |  | 264130655 | [13]C5H10[15]N2O3 | 53354.65
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C7 | 19397472 |  | 304775462 | [13]C5H10[15]N2O3 | 85614.89
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C8 | 21887625 |  | 322100123 | [13]C5H10[15]N2O3 | 45614.46
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C9 | 23763267 |  | 347544802 | [13]C5H10[15]N2O3 | 99134.64
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C10 | 23426753 |  | 337695214 | [13]C5H10[15]N2O3 | 44779.72
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C11 | 25867416 |  | 371444790 | [13]C5H10[15]N2O3 | 19273.07
Glutamine U-13C, U-15N neg | 9.64 | Internal Standard | KD100520_100520_C12 | 19854368 |  | 321817592 | [13]C5H10[15]N2O3 | 23400.48

## Glutamine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A1 | 108892178 |  | 1540320659 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A2 | 155752913 |  | 2154226812 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A3 | 171420607 |  | 2552027562 | [13]C5H10[15]N2O3 | 162493.04
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A4 | 203193610 |  | 2990411855 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A5 | 205921953 |  | 2995894522 | [13]C5H10[15]N2O3 | 635039.82
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A6 | 223481653 |  | 3369178088 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A7 | 224808384 |  | 3432009308 | [13]C5H10[15]N2O3 | 389368.10
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A8 | 229829227 |  | 3663757007 | [13]C5H10[15]N2O3 | 321919.33
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A9 | 242174355 |  | 3707968263 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A10 | 215083124 |  | 3578094435 | [13]C5H10[15]N2O3 | 6555683.72
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A11 | 232465845 |  | 3642009208 | [13]C5H10[15]N2O3 | 762821.65
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_A12 | 243687976 |  | 3784889414 | [13]C5H10[15]N2O3 | 6235815.23
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B1 | 94930121 |  | 1303340185 | [13]C5H10[15]N2O3 | INF | 9.685:Peak area 1303340185.461 is out of bounds (ISTD Minimum Recovery 1525321867.274 and ISTD Max Recovery 4575965601.821)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B2 | 142742590 |  | 2009943876 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B3 | 164214851 |  | 2430064596 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B4 | 186975167 |  | 2770891557 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B5 | 212799817 |  | 3079997681 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B6 | 240317876 |  | 3403627578 | [13]C5H10[15]N2O3 | 985701.81
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B7 | 219051451 |  | 3481223643 | [13]C5H10[15]N2O3 | 685818.03
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B8 | 208683943 |  | 3534442295 | [13]C5H10[15]N2O3 | 1054108.32
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B9 | 237038553 |  | 3557439817 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B10 | 237917220 |  | 3642888299 | [13]C5H10[15]N2O3 | 884161.59
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B11 | 219619657 |  | 3466052470 | [13]C5H10[15]N2O3 | 200429.75
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_B12 | 213614378 |  | 3393888373 | [13]C5H10[15]N2O3 | 208866.72
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_blank | 3989 |  | 14926 | [13]C5H10[15]N2O3 | 181.13 | 9.851:Peak area 14926.450 is out of bounds (ISTD Minimum Recovery 1525321867.274 and ISTD Max Recovery 4575965601.821)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C1 | 121784317 |  | 1633466978 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C2 | 151553083 |  | 2052188458 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C3 | 178769387 |  | 2486439676 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C4 | 192031580 |  | 2766917688 | [13]C5H10[15]N2O3 | 189430.40
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C5 | 199493888 |  | 3028841195 | [13]C5H10[15]N2O3 | 6902601.56
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C6 | 213095241 |  | 3186345452 | [13]C5H10[15]N2O3 | 13446949.65
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C7 | 210232013 |  | 3441277804 | [13]C5H10[15]N2O3 | 541126.36
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C8 | 224509491 |  | 3422630562 | [13]C5H10[15]N2O3 | INF
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C9 | 236272002 |  | 3676046453 | [13]C5H10[15]N2O3 | 2161109.83
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C10 | 232264819 |  | 3513285586 | [13]C5H10[15]N2O3 | 595934.23
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C11 | 229646474 |  | 3764856051 | [13]C5H10[15]N2O3 | 7831200.35
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD100520_100520_C12 | 210841574 |  | 3366291039 | [13]C5H10[15]N2O3 | 479435.35

## Glycine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glycine neg | 10.11 | Target Compound | KD100520_100520_A1 | 6829006 |  | 106352914 | C2H5NO2 | 1013.75
Glycine neg | 10.11 | Target Compound | KD100520_100520_A2 | 3079047 |  | 47073346 | C2H5NO2 | 1126.77
Glycine neg | 10.11 | Target Compound | KD100520_100520_A3 | 1962885 |  | 27694143 | C2H5NO2 | 733.14
Glycine neg | 10.11 | Target Compound | KD100520_100520_A4 | 845346 |  | 12338295 | C2H5NO2 | 334.64
Glycine neg | 10.11 | Target Compound | KD100520_100520_A5 | 535139 |  | 7527109 | C2H5NO2 | 237.01
Glycine neg | 10.11 | Target Compound | KD100520_100520_A6 | 177650 |  | 2413880 | C2H5NO2 | 156.76
Glycine neg | 10.11 | Target Compound | KD100520_100520_A7 | 59251 |  | 834275 | C2H5NO2 | 53.12 | Calibration Amount Diff 42.869% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_A8 | 28998 |  | 480784 | C2H5NO2 | 39.83 | Calibration Amount Diff 146.19% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_A9 | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_B1 | 7646468 |  | 114771318 | C2H5NO2 | 2180.11
Glycine neg | 10.11 | Target Compound | KD100520_100520_B2 | 3315956 |  | 50856837 | C2H5NO2 | 1254.39
Glycine neg | 10.11 | Target Compound | KD100520_100520_B3 | 1889645 |  | 27870398 | C2H5NO2 | 906.56
Glycine neg | 10.11 | Target Compound | KD100520_100520_B4 | 1016385 |  | 13607042 | C2H5NO2 | 428.38
Glycine neg | 10.11 | Target Compound | KD100520_100520_B5 | 336932 |  | 5026486 | C2H5NO2 | 166.00
Glycine neg | 10.11 | Target Compound | KD100520_100520_B6 | 147236 |  | 1959164 | C2H5NO2 | 130.56
Glycine neg | 10.11 | Target Compound | KD100520_100520_B7 | 72815 |  | 827928 | C2H5NO2 | 61.18 | Calibration Amount Diff 41.188% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_B8 | 28309 |  | 357373 | C2H5NO2 | 40.93 | Calibration Amount Diff 135.703% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_B9 | 18438 |  | 288757 | C2H5NO2 | 35.05 | Calibration Amount Diff 355.108% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_B10 | 16779 |  | 296585 | C2H5NO2 | 34.69 | Calibration Amount Diff 811.44% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_B12 | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_C1 | 6550414 |  | 102926361 | C2H5NO2 | 1584.04
Glycine neg | 10.11 | Target Compound | KD100520_100520_C2 | 3812771 |  | 54925308 | C2H5NO2 | 1645.72
Glycine neg | 10.11 | Target Compound | KD100520_100520_C3 | 1834312 |  | 26447973 | C2H5NO2 | 766.19
Glycine neg | 10.11 | Target Compound | KD100520_100520_C4 | 859469 |  | 12581863 | C2H5NO2 | 523.39
Glycine neg | 10.11 | Target Compound | KD100520_100520_C5 | 376976 |  | 5397268 | C2H5NO2 | 234.04
Glycine neg | 10.11 | Target Compound | KD100520_100520_C6 | 125483 |  | 1612633 | C2H5NO2 | 93.21
Glycine neg | 10.11 | Target Compound | KD100520_100520_C7 | 54123 |  | 807584 | C2H5NO2 | 64.24 | Calibration Amount Diff 39.172% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_C8 | 32983 |  | 486910 | C2H5NO2 | 60.94 | Calibration Amount Diff 149.135% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_C9 | 20418 |  | 304031 | C2H5NO2 | 42.23 | Calibration Amount Diff 357.298% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD100520_100520_C10 | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_C11 | N/F |  | N/F | C2H5NO2 | N/A
Glycine neg | 10.11 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | C2H5NO2 | N/A

## Glycine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glycine pos | 10.11 | Target Compound | KD100520_100520_A1 | 7510000 |  | 99992985 | C2H5NO2 | INF
Glycine pos | 10.11 | Target Compound | KD100520_100520_A2 | 4103939 |  | 52784289 | C2H5NO2 | INF
Glycine pos | 10.11 | Target Compound | KD100520_100520_A3 | 2252718 |  | 31043289 | C2H5NO2 | 7383.40
Glycine pos | 10.11 | Target Compound | KD100520_100520_A4 | 1298890 |  | 17535859 | C2H5NO2 | 1904.65 | Calibration Amount Diff -22.895% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_A5 | 874316 |  | 12452031 | C2H5NO2 | 1132.39
Glycine pos | 10.11 | Target Compound | KD100520_100520_A6 | 324819 |  | 4694584 | C2H5NO2 | 765.04
Glycine pos | 10.11 | Target Compound | KD100520_100520_A7 | 152693 |  | 2241617 | C2H5NO2 | 534.47 | Calibration Amount Diff 38.595% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_A8 | 74042 |  | 1133825 | C2H5NO2 | INF | Calibration Amount Diff 124.044% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_A9 | 42876 |  | 585384 | C2H5NO2 | 323.61 | Calibration Amount Diff 305.945% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_A10 | 21796 |  | 269103 | C2H5NO2 | 2760.17 | Calibration Amount Diff 655.256% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_A11 | 19001 |  | 125422 | C2H5NO2 | INF | Calibration Amount Diff 1371.017% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_A12 | 13825 |  | 91291 | C2H5NO2 | INF | Calibration Amount Diff 2817.715% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_B1 | 7853106 |  | 103881127 | C2H5NO2 | INF
Glycine pos | 10.11 | Target Compound | KD100520_100520_B2 | 4353200 |  | 58178921 | C2H5NO2 | INF
Glycine pos | 10.11 | Target Compound | KD100520_100520_B3 | 2377818 |  | 32340121 | C2H5NO2 | 6169.04
Glycine pos | 10.11 | Target Compound | KD100520_100520_B4 | 1568971 |  | 20577927 | C2H5NO2 | 1369.59
Glycine pos | 10.11 | Target Compound | KD100520_100520_B5 | 590866 |  | 8794163 | C2H5NO2 | 1476.39 | Calibration Amount Diff -22.971% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_B6 | 351971 |  | 4737523 | C2H5NO2 | 533.85
Glycine pos | 10.11 | Target Compound | KD100520_100520_B7 | 167283 |  | 2487201 | C2H5NO2 | 423.90 | Calibration Amount Diff 40.879% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_B8 | 101690 |  | 1221396 | C2H5NO2 | 375.31 | Calibration Amount Diff 125.615% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_B9 | 44774 |  | 636399 | C2H5NO2 | 13319.04 | Calibration Amount Diff 306.278% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_B10 | 21643 |  | 280506 | C2H5NO2 | 5065.27 | Calibration Amount Diff 654.647% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_B11 | 12938 |  | 158457 | C2H5NO2 | 508.19 | Calibration Amount Diff 1379.659% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_B12 | 8679 |  | 52523 | C2H5NO2 | 206.63 | Calibration Amount Diff 2796.86% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_blank | 1876 |  | 15144 | C2H5NO2 | 99.82
Glycine pos | 10.11 | Target Compound | KD100520_100520_C1 | 7385377 |  | 103051944 | C2H5NO2 | INF
Glycine pos | 10.11 | Target Compound | KD100520_100520_C2 | 4099321 |  | 60116451 | C2H5NO2 | INF
Glycine pos | 10.11 | Target Compound | KD100520_100520_C3 | 2542568 |  | 32977341 | C2H5NO2 | 23425.90
Glycine pos | 10.11 | Target Compound | KD100520_100520_C4 | 1478708 |  | 20286569 | C2H5NO2 | 3538.71
Glycine pos | 10.11 | Target Compound | KD100520_100520_C5 | 772288 |  | 9712325 | C2H5NO2 | 1321.41
Glycine pos | 10.11 | Target Compound | KD100520_100520_C6 | 297930 |  | 3961748 | C2H5NO2 | 416.50
Glycine pos | 10.11 | Target Compound | KD100520_100520_C7 | 144081 |  | 2229526 | C2H5NO2 | 358.70 | Calibration Amount Diff 31.819% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_C8 | 84110 |  | 1134078 | C2H5NO2 | 390.68 | Calibration Amount Diff 120.561% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_C9 | 54253 |  | 682108 | C2H5NO2 | 1778.53 | Calibration Amount Diff 307.856% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_C10 | 29798 |  | 260653 | C2H5NO2 | 1039.87 | Calibration Amount Diff 653.688% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_C11 | 16648 |  | 158877 | C2H5NO2 | 69.43 | Calibration Amount Diff 1377.791% (absolute) > max calibration amount diff 20%
Glycine pos | 10.11 | Target Compound | KD100520_100520_C12 | 8087 |  | 81320 | C2H5NO2 | INF | Calibration Amount Diff 2816.525% (absolute) > max calibration amount diff 20%

## Glycine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A1 | 21689 |  | 211127 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A2 | 24602 |  | 256473 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -28.469% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A3 | 28283 |  | 256617 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 108.277% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A4 | 23500 |  | 215127 | [13]C2H5NO2 | 992.15 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 1547.124% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A5 | 21794 |  | 329089 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -1326.967% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A6 | 18180 |  | 271966 | [13]C2H5NO2 | 4647.52 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 576.762% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A7 | 19039 |  | 264944 | [13]C2H5NO2 | 779.86 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 1957.57% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A8 | 21377 |  | 274511 | [13]C2H5NO2 | 712.03 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 4176.925% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A9 | 23626 |  | 326535 | [13]C2H5NO2 | 281.56 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -39493.18% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A10 | 20121 |  | 287631 | [13]C2H5NO2 | 4391.65 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -8210.625% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A11 | 19853 |  | 291822 | [13]C2H5NO2 | 443.47 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -28753.476% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_A12 | 22399 |  | 309110 | [13]C2H5NO2 | 1869.73 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -235589.951% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B1 | 24228 |  | 191409 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 98.616% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B2 | 17969 |  | 225678 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 151.555% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B3 | 22459 |  | 204289 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 790.054% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B4 | 16401 |  | 166008 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 2651.271% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B5 | 25443 |  | 323792 | [13]C2H5NO2 | 320.92 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -857.584% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B6 | 19582 |  | 258063 | [13]C2H5NO2 | 318.93 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 2085.766% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B7 | 19674 |  | 299568 | [13]C2H5NO2 | 1333.93 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -2649.379% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B8 | 26067 |  | 270085 | [13]C2H5NO2 | 251.49 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 91.45% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B9 | 22476 |  | 309435 | [13]C2H5NO2 | 272.90 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -23850.88% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B10 | 20689 |  | 283873 | [13]C2H5NO2 | 144.66 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -5361.254% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B11 | 18044 |  | 228347 | [13]C2H5NO2 | 5018.51 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 113422.074% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_B12 | 18642 |  | 220034 | [13]C2H5NO2 | 1205.23 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 292761.013% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C2H5NO2 | N/A
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C1 | 25717 |  | 274335 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -205.619% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C2 | 32298 |  | 232900 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 90.204% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C3 | 23790 |  | 279562 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -308.216% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C4 | 28280 |  | 330050 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -1181.831% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C5 | 33836 |  | 370866 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -5328.359% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C6 | 18504 |  | 253782 | [13]C2H5NO2 | 1775.41 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -410.624% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C7 | 15561 |  | 208660 | [13]C2H5NO2 | 214.21 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 14106.687% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C8 | 23420 |  | 323417 | [13]C2H5NO2 | INF | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -18045.005% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C9 | 17489 |  | 269542 | [13]C2H5NO2 | 3418.12 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 8144.984% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C10 | 20788 |  | 272358 | [13]C2H5NO2 | 174.97 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -812.315% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C11 | 20885 |  | 270853 | [13]C2H5NO2 | 513.65 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff 5521.63% (absolute) > max calibration amount diff 20%
Glycine U-13C neg | 10.11 | Target Compound | KD100520_100520_C12 | 24911 |  | 294604 | [13]C2H5NO2 | 4849.68 | Rsquared 0.051 < R2Threshold 0.99. Calibration Amount Diff -193621.11% (absolute) > max calibration amount diff 20%

## Glycine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A1 | 27001 |  | 91911 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 24.536% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A2 | 17062 |  | 58276 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 229.884% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A3 | 24846 |  | 85141 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 521.948% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A4 | 23941 |  | 172889 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 947.426% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A5 | 44057 |  | 460939 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 174.089% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A6 | 49406 |  | 536019 | [13]C2H5NO2 | 8626.77 | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 375.975% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A7 | 60601 |  | 755315 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -3249.326% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A8 | 62733 |  | 736812 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -2510.181% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A9 | 49605 |  | 658469 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -225.33% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A10 | 66103 |  | 949801 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -32823.829% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A11 | 46966 |  | 641740 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 19940.356% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_A12 | 75081 |  | 847268 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -54098.645% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B1 | 52507 |  | 240632 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -128.035% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B2 | 22964 |  | 78410 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 206.357% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B3 | 42009 |  | 220198 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 181.319% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B4 | 45942 |  | 407911 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 89.659% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B5 | 43004 |  | 534138 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -57.294% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B6 | 49172 |  | 682396 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -596.915% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B7 | 56295 |  | 669215 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -486.78% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B8 | 68079 |  | 658866 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 1680.988% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B9 | 45896 |  | 676939 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 2392.245% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B10 | 66296 |  | 852255 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -11184.711% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B11 | 49324 |  | 744912 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -5202.646% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_B12 | 57839 |  | 793573 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -42557.31% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C2H5NO2 | N/A
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C1 | N/F |  | N/F | [13]C2H5NO2 | N/A | Rsquared 0.175 < R2Threshold 0.99
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C2 | 32461 |  | 370751 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -261.552% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C3 | 57345 |  | 305694 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -20.35% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C4 | 43019 |  | 341628 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 320.028% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C5 | 61418 |  | 589075 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -588.559% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C6 | 50164 |  | 658330 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -829.568% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C7 | 53229 |  | 662577 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff 780.615% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C8 | 64575 |  | 746158 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -552.08% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C9 | 76038 |  | 877310 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -10578.952% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C10 | 76472 |  | 845413 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -16149.243% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C11 | 58073 |  | 818615 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -16917.22% (absolute) > max calibration amount diff 20%
Glycine U-13C pos | 10.11 | Target Compound | KD100520_100520_C12 | 54635 |  | 704161 | [13]C2H5NO2 | INF | Rsquared 0.175 < R2Threshold 0.99. Calibration Amount Diff -10710.039% (absolute) > max calibration amount diff 20%

## Glycine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A1 | 768023 |  | 9816572 | [13]C2H5[15]NO2 | 1142.19
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A2 | 869277 |  | 10971779 | [13]C2H5[15]NO2 | 1691.17
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A3 | 835380 |  | 11223954 | [13]C2H5[15]NO2 | 1661.66
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A4 | 849879 |  | 11932169 | [13]C2H5[15]NO2 | 1056.70
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A5 | 832526 |  | 12229384 | [13]C2H5[15]NO2 | 898.50
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A6 | 818027 |  | 11410327 | [13]C2H5[15]NO2 | 1105.24
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A7 | 812708 |  | 11279367 | [13]C2H5[15]NO2 | 1011.37
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A8 | 868240 |  | 11706403 | [13]C2H5[15]NO2 | 1331.78
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A9 | 772418 |  | 11131108 | [13]C2H5[15]NO2 | 854.51
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A10 | 780784 |  | 11506277 | [13]C2H5[15]NO2 | 961.52
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A11 | 915215 |  | 11498744 | [13]C2H5[15]NO2 | 804.67
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_A12 | 818981 |  | 10995957 | [13]C2H5[15]NO2 | 775.11
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B1 | 760275 |  | 10483702 | [13]C2H5[15]NO2 | 1878.89
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B2 | 828511 |  | 10983561 | [13]C2H5[15]NO2 | 1018.90
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B3 | 877864 |  | 11669135 | [13]C2H5[15]NO2 | 1347.35
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B4 | 873404 |  | 12124766 | [13]C2H5[15]NO2 | 805.30
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B5 | 860128 |  | 12459423 | [13]C2H5[15]NO2 | 849.58
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B6 | 824823 |  | 11544877 | [13]C2H5[15]NO2 | 1206.99
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B7 | 877328 |  | 11632988 | [13]C2H5[15]NO2 | 955.42
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B8 | 825861 |  | 11045201 | [13]C2H5[15]NO2 | 858.27
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B9 | 837056 |  | 11287002 | [13]C2H5[15]NO2 | 744.56
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B10 | 832769 |  | 11435987 | [13]C2H5[15]NO2 | 959.14
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B11 | 776525 |  | 10864536 | [13]C2H5[15]NO2 | 1042.88
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_B12 | 794349 |  | 10997140 | [13]C2H5[15]NO2 | 866.88
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C2H5[15]NO2 | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C1 | 705540 |  | 9861421 | [13]C2H5[15]NO2 | 1003.55
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C2 | 848122 |  | 10827078 | [13]C2H5[15]NO2 | 762.32
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C3 | 802021 |  | 10696593 | [13]C2H5[15]NO2 | 679.55
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C4 | 835768 |  | 11480178 | [13]C2H5[15]NO2 | 801.61
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C5 | 733540 |  | 10666539 | [13]C2H5[15]NO2 | 603.29
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C6 | 697808 |  | 10231189 | [13]C2H5[15]NO2 | 714.10
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C7 | 819465 |  | 11907724 | [13]C2H5[15]NO2 | 876.83
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C8 | 800485 |  | 11188857 | [13]C2H5[15]NO2 | 1081.64
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C9 | 815211 |  | 11475937 | [13]C2H5[15]NO2 | 786.61
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C10 | 822385 |  | 11096978 | [13]C2H5[15]NO2 | 896.51
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C11 | 842170 |  | 11133743 | [13]C2H5[15]NO2 | 815.92
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD100520_100520_C12 | 742270 |  | 10725509 | [13]C2H5[15]NO2 | 862.99

## Glycine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A1 | 818262 |  | 10137262 | [13]C2H5[15]NO2 | 3613.66 | 10.118:Peak area 10137261.502 is out of bounds (ISTD Minimum Recovery 12721421.141 and ISTD Max Recovery 38164263.422)
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A2 | 887414 |  | 11779688 | [13]C2H5[15]NO2 | 1242.13 | 10.101:Peak area 11779687.919 is out of bounds (ISTD Minimum Recovery 12721421.141 and ISTD Max Recovery 38164263.422)
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A3 | 1143961 |  | 14403518 | [13]C2H5[15]NO2 | 1276.20
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A4 | 1690323 |  | 20545417 | [13]C2H5[15]NO2 | 3622.10
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A5 | 1563467 |  | 23027365 | [13]C2H5[15]NO2 | 6113.94
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A6 | 1980912 |  | 26473928 | [13]C2H5[15]NO2 | 4994.00
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A7 | 2035712 |  | 28203630 | [13]C2H5[15]NO2 | 11314.09
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A8 | 2089829 |  | 31111258 | [13]C2H5[15]NO2 | 5264.93
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A9 | 2075676 |  | 30186417 | [13]C2H5[15]NO2 | 5084.67
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A10 | 2292783 |  | 33582490 | [13]C2H5[15]NO2 | 3409.40
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A11 | 2235194 |  | 32461330 | [13]C2H5[15]NO2 | 4255.24
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_A12 | 2315932 |  | 34652933 | [13]C2H5[15]NO2 | 2977.17
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B1 | 707795 |  | 9773103 | [13]C2H5[15]NO2 | 16475.40 | 10.103:Peak area 9773102.803 is out of bounds (ISTD Minimum Recovery 12721421.141 and ISTD Max Recovery 38164263.422)
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B2 | 1108897 |  | 12756903 | [13]C2H5[15]NO2 | 3911.94
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B3 | 1096227 |  | 15089099 | [13]C2H5[15]NO2 | 1786.90
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B4 | 1533308 |  | 21084531 | [13]C2H5[15]NO2 | 2861.97
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B5 | 1783769 |  | 24853583 | [13]C2H5[15]NO2 | 4530.37
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B6 | 2056594 |  | 29228817 | [13]C2H5[15]NO2 | 4977.05
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B7 | 2016534 |  | 29902786 | [13]C2H5[15]NO2 | 3759.43
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B8 | 2319265 |  | 32384609 | [13]C2H5[15]NO2 | 7233.62
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B9 | 2250716 |  | 32590566 | [13]C2H5[15]NO2 | 12080.26
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B10 | 2355648 |  | 35552451 | [13]C2H5[15]NO2 | 2016.20
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B11 | 2349434 |  | 33447238 | [13]C2H5[15]NO2 | 2364.95
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_B12 | 2285908 |  | 33237400 | [13]C2H5[15]NO2 | 8358.68
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C2H5[15]NO2 | N/A
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C1 | 730735 |  | 9317053 | [13]C2H5[15]NO2 | 1096.06 | 10.105:Peak area 9317052.994 is out of bounds (ISTD Minimum Recovery 12721421.141 and ISTD Max Recovery 38164263.422)
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C2 | 916606 |  | 12358821 | [13]C2H5[15]NO2 | 2714.59 | 10.116:Peak area 12358820.871 is out of bounds (ISTD Minimum Recovery 12721421.141 and ISTD Max Recovery 38164263.422)
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C3 | 1200871 |  | 15491237 | [13]C2H5[15]NO2 | 2717.39
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C4 | 1569764 |  | 20817487 | [13]C2H5[15]NO2 | 2444.95
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C5 | 1831609 |  | 23679839 | [13]C2H5[15]NO2 | 3223.56
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C6 | 1921911 |  | 27330295 | [13]C2H5[15]NO2 | 6455.74
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C7 | 2384902 |  | 32542158 | [13]C2H5[15]NO2 | 2783.30
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C8 | 2286904 |  | 33726836 | [13]C2H5[15]NO2 | 3244.99
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C9 | 2343841 |  | 33825261 | [13]C2H5[15]NO2 | 6106.99
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C10 | 2404935 |  | 33869667 | [13]C2H5[15]NO2 | 4726.05
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C11 | 2410384 |  | 34928430 | [13]C2H5[15]NO2 | 4535.14
Glycine U-13C, U-15N pos | 10.11 | Internal Standard | KD100520_100520_C12 | 2125960 |  | 31588919 | [13]C2H5[15]NO2 | 2767.91

## Histidine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Histidine neg | 11.94 | Target Compound | KD100520_100520_A1 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A2 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A3 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A4 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A5 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A6 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A7 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A8 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A9 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B1 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B2 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B3 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B4 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B5 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B6 | 85039 |  | 691881 | C6H9N3O2 | 42.21
Histidine neg | 11.94 | Target Compound | KD100520_100520_B7 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B8 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B9 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B10 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_B12 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C1 | 102087 |  | 338342 | C6H9N3O2 | 42.89
Histidine neg | 11.94 | Target Compound | KD100520_100520_C2 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C3 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C4 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C5 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C6 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C7 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C8 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C9 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C10 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C11 | N/F |  | N/F | C6H9N3O2 | N/A
Histidine neg | 11.94 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | C6H9N3O2 | N/A

## Histidine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Histidine pos | 11.94 | Target Compound | KD100520_100520_A1 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_A2 | 1754809 |  | 9555802 | C6H9N3O2 | 31.31 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -107.17% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_A3 | 3919636 |  | 58232133 | C6H9N3O2 | 135.31 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff 42.358% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_A4 | 3872062 |  | 49164520 | C6H9N3O2 | 84.35 | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_A5 | 2654572 |  | 40312456 | C6H9N3O2 | 64.69 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff 44.666% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_A6 | 2229907 |  | 27743250 | C6H9N3O2 | 97.32 | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_A7 | 1169517 |  | 13492673 | C6H9N3O2 | 61.73 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -138.415% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_A8 | 946543 |  | 16961037 | C6H9N3O2 | 45.84 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -388.859% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_A9 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_B1 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_B2 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_B3 | 4281058 |  | 52625567 | C6H9N3O2 | 106.81 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff 27.681% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_B4 | 2606068 |  | 44714318 | C6H9N3O2 | 68.63 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff 30.07% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_B5 | 2320149 |  | 22257979 | C6H9N3O2 | 68.11 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -23.432% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_B6 | 1603320 |  | 21981886 | C6H9N3O2 | 179.33 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -42.017% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_B7 | 1136765 |  | 15411420 | C6H9N3O2 | 50.75 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -147.965% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_B8 | 770857 |  | 9588819 | C6H9N3O2 | 65.80 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -423.142% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_B9 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_B10 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_B12 | 108059 |  | 1566251 | C6H9N3O2 | 30.28 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -7125.554% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H9N3O2 | N/A
Histidine pos | 11.94 | Target Compound | KD100520_100520_C1 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_C2 | 1198745 |  | 4952450 | C6H9N3O2 | 31.21 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff 76.83% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_C3 | 4148952 |  | 32074567 | C6H9N3O2 | 118.73 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff 34.259% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_C4 | 2925701 |  | 39804015 | C6H9N3O2 | 126.89 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff 27.439% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_C5 | 2289209 |  | 22055231 | C6H9N3O2 | 129.05 | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_C6 | 1189311 |  | 15969744 | C6H9N3O2 | 56.50 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -43.885% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_C7 | 893958 |  | 17219069 | C6H9N3O2 | 49.93 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -153.476% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_C8 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_C9 | 360686 |  | 4737480 | C6H9N3O2 | 33.91 | Rsquared 0.495 < R2Threshold 0.99. Calibration Amount Diff -786.478% (absolute) > max calibration amount diff 20%
Histidine pos | 11.94 | Target Compound | KD100520_100520_C10 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_C11 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99
Histidine pos | 11.94 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | C6H9N3O2 | N/A | Rsquared 0.495 < R2Threshold 0.99

## Histidine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A1 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A2 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A3 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A4 | 7910 |  | 80104 | [13]C6H9N3O2 | INF | Rsquared 0.202 < R2Threshold 0.99. Calibration Amount Diff 112.825% (absolute) > max calibration amount diff 20%
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A5 | 5350 |  | 21084 | [13]C6H9N3O2 | 875.19 | Rsquared 0.202 < R2Threshold 0.99. Calibration Amount Diff -96.986% (absolute) > max calibration amount diff 20%
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A6 | 12000 |  | 56750 | [13]C6H9N3O2 | INF | Rsquared 0.202 < R2Threshold 0.99. Calibration Amount Diff -93.972% (absolute) > max calibration amount diff 20%
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A7 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A8 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A9 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B1 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B2 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B3 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B4 | 51475 |  | 177500 | [13]C6H9N3O2 | 16423.92 | Rsquared 0.202 < R2Threshold 0.99. Calibration Amount Diff -98.493% (absolute) > max calibration amount diff 20%
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B5 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B6 | 9878 |  | 92420 | [13]C6H9N3O2 | 242.33 | Rsquared 0.202 < R2Threshold 0.99. Calibration Amount Diff 274.592% (absolute) > max calibration amount diff 20%
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B7 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B8 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B9 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B10 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_B12 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H9N3O2 | N/A
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C1 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C2 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C3 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C4 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C5 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C6 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C7 | 15655 |  | 108411 | [13]C6H9N3O2 | 66.88 | Rsquared 0.202 < R2Threshold 0.99. Calibration Amount Diff -87.945% (absolute) > max calibration amount diff 20%
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C8 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C9 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C10 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C11 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99
Histidine U-13C neg | 11.94 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.202 < R2Threshold 0.99

## Histidine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A1 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A2 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A3 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A4 | 204046 |  | 2460707 | [13]C6H9N3O2 | 88.63 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff -90.187% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A5 | 207911 |  | 2756894 | [13]C6H9N3O2 | 69.31 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff -193.863% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A6 | 412820 |  | 5453732 | [13]C6H9N3O2 | 91.94 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 1397.447% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A7 | 433794 |  | 5108214 | [13]C6H9N3O2 | 48.21 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 1379.358% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A8 | 667466 |  | 11283081 | [13]C6H9N3O2 | 44.97 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff -567.224% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A9 | 695184 |  | 10559540 | [13]C6H9N3O2 | 37.15 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff -25638.635% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A10 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A11 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_A12 | 742892 |  | 9115194 | [13]C6H9N3O2 | 30.05 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff -204513.752% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B1 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B2 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B3 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B4 | 137582 |  | 2057709 | [13]C6H9N3O2 | 65.65 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff -103.727% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B5 | 247953 |  | 2929229 | [13]C6H9N3O2 | 119.34 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 220.315% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B6 | 294516 |  | 4222149 | [13]C6H9N3O2 | 52.80 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 314.666% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B7 | 476312 |  | 6589490 | [13]C6H9N3O2 | 65.81 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 1952.677% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B8 | 530809 |  | 9729852 | [13]C6H9N3O2 | 52.59 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 2113.27% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B9 | 530606 |  | 8986887 | [13]C6H9N3O2 | 33.08 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 3566.334% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B10 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B11 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_B12 | 795137 |  | 11873767 | [13]C6H9N3O2 | 117.58 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 60379.389% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H9N3O2 | N/A
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C1 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C2 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C3 | 120978 |  | 920396 | [13]C6H9N3O2 | 87.63 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff -35.393% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C4 | 167180 |  | 2235332 | [13]C6H9N3O2 | 87.84 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 55.222% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C5 | 254458 |  | 2498857 | [13]C6H9N3O2 | 95.36 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 323.563% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C6 | 237942 |  | 3370985 | [13]C6H9N3O2 | 48.72 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 628.789% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C7 | 377290 |  | 6312194 | [13]C6H9N3O2 | 44.88 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 529.636% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C8 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C9 | 445565 |  | 5564384 | [13]C6H9N3O2 | 34.72 | Rsquared 0.028 < R2Threshold 0.99. Calibration Amount Diff 3913.867% (absolute) > max calibration amount diff 20%
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C10 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C11 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99
Histidine U-13C pos | 11.94 | Target Compound | KD100520_100520_C12 | N/F |  | N/F | [13]C6H9N3O2 | N/A | Rsquared 0.028 < R2Threshold 0.99

## Histidine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A1 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A2 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A3 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A4 | 190853 |  | 1163893 | [13]C6H9[15]N3O2 | 43.46 | 11.741:Peak area 1163893.007 is out of bounds (ISTD Minimum Recovery 325103.395 and ISTD Max Recovery 975310.186). 11.741:Apex Retention Time 11.741 is out of bounds (ISTD Min RT 1.377 and ISTD Max RT 1.877)
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A5 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A6 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A7 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A8 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A9 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A10 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A11 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_A12 | 955524 |  | 12630841 | [13]C6H9[15]N3O2 | 48.31 | 11.694:Peak area 12630840.796 is out of bounds (ISTD Minimum Recovery 325103.395 and ISTD Max Recovery 975310.186). 11.694:Apex Retention Time 11.694 is out of bounds (ISTD Min RT 1.377 and ISTD Max RT 1.877)
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B1 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B2 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B3 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B4 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B5 | 227958 |  | 3329503 | [13]C6H9[15]N3O2 | 63.13 | 11.693:Peak area 3329502.863 is out of bounds (ISTD Minimum Recovery 325103.395 and ISTD Max Recovery 975310.186). 11.693:Apex Retention Time 11.693 is out of bounds (ISTD Min RT 1.377 and ISTD Max RT 1.877)
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B6 | 298370 |  | 3079721 | [13]C6H9[15]N3O2 | 85.65 | 11.760:Peak area 3079721.462 is out of bounds (ISTD Minimum Recovery 325103.395 and ISTD Max Recovery 975310.186). 11.760:Apex Retention Time 11.760 is out of bounds (ISTD Min RT 1.377 and ISTD Max RT 1.877)
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B7 | 376950 |  | 3203486 | [13]C6H9[15]N3O2 | 48.25 | 11.695:Peak area 3203486.332 is out of bounds (ISTD Minimum Recovery 325103.395 and ISTD Max Recovery 975310.186). 11.695:Apex Retention Time 11.695 is out of bounds (ISTD Min RT 1.377 and ISTD Max RT 1.877)
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B8 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B9 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B10 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B11 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_B12 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C1 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C2 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C3 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C4 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C5 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C6 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C7 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C8 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C9 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C10 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C11 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N neg | 11.94 | Internal Standard | KD100520_100520_C12 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.

## Histidine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A1 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A2 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A3 | 1789776 |  | 25344160 | [13]C6H9[15]N3O2 | 192.66 | 11.726:Peak area 25344159.758 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.726:Apex Retention Time 11.726 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A4 | 3453326 |  | 45172562 | [13]C6H9[15]N3O2 | 86.20 | 11.725:Apex Retention Time 11.725 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A5 | 3422668 |  | 54433274 | [13]C6H9[15]N3O2 | 72.25 | 11.810:Apex Retention Time 11.810 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A6 | 6035010 |  | 68986247 | [13]C6H9[15]N3O2 | 125.93 | 11.797:Apex Retention Time 11.797 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A7 | 6429724 |  | 77067978 | [13]C6H9[15]N3O2 | 38.33 | 11.815:Apex Retention Time 11.815 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A8 | 10714909 |  | 217639893 | [13]C6H9[15]N3O2 | 87.55 | 11.813:Peak area 217639892.899 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.813:Apex Retention Time 11.813 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A9 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A10 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A11 | 8683221 |  | 149516901 | [13]C6H9[15]N3O2 | 36.85 | 11.791:Peak area 149516900.808 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.791:Apex Retention Time 11.791 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_A12 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B1 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B2 | 350641 |  | 1457039 | [13]C6H9[15]N3O2 | 30.51 | 11.761:Peak area 1457038.756 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.761:Apex Retention Time 11.761 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B3 | 2044442 |  | 25271096 | [13]C6H9[15]N3O2 | 124.74 | 11.701:Peak area 25271095.751 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.701:Apex Retention Time 11.701 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B4 | 2228979 |  | 38418420 | [13]C6H9[15]N3O2 | 63.58 | 11.779:Apex Retention Time 11.779 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B5 | 4005766 |  | 45336110 | [13]C6H9[15]N3O2 | 78.52 | 11.801:Apex Retention Time 11.801 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B6 | 4986861 |  | 69440770 | [13]C6H9[15]N3O2 | 110.54 | 11.776:Apex Retention Time 11.776 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B7 | 6794215 |  | 92660093 | [13]C6H9[15]N3O2 | 55.90 | 11.802:Apex Retention Time 11.802 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B8 | 8685245 |  | 154102326 | [13]C6H9[15]N3O2 | 74.10 | 11.829:Peak area 154102325.595 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.829:Apex Retention Time 11.829 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B9 | 8695310 |  | 146040162 | [13]C6H9[15]N3O2 | 32.79 | 11.870:Peak area 146040161.844 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.870:Apex Retention Time 11.870 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B10 | 9318823 |  | 157425240 | [13]C6H9[15]N3O2 | 36.21 | 11.892:Peak area 157425240.111 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.892:Apex Retention Time 11.892 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B11 | 9513959 |  | 161429256 | [13]C6H9[15]N3O2 | 40.26 | 11.716:Peak area 161429255.520 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.716:Apex Retention Time 11.716 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_B12 | 11369015 |  | 170274661 | [13]C6H9[15]N3O2 | 364.08 | 11.839:Peak area 170274660.609 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.839:Apex Retention Time 11.839 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C1 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C2 | 267880 |  | 917809 | [13]C6H9[15]N3O2 | 31.16 | 11.751:Peak area 917809.343 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.751:Apex Retention Time 11.751 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C3 | 1896871 |  | 14720506 | [13]C6H9[15]N3O2 | 139.73 | 11.697:Peak area 14720506.399 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.697:Apex Retention Time 11.697 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C4 | 2520335 |  | 34775659 | [13]C6H9[15]N3O2 | 146.03 | 11.767:Apex Retention Time 11.767 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C5 | 3699697 |  | 36698003 | [13]C6H9[15]N3O2 | 99.42 | 11.801:Apex Retention Time 11.801 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C6 | 3805695 |  | 51000154 | [13]C6H9[15]N3O2 | 122.87 | 11.779:Apex Retention Time 11.779 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C7 | 5328651 |  | 106770453 | [13]C6H9[15]N3O2 | 40.60 | 11.876:Peak area 106770452.829 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.876:Apex Retention Time 11.876 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C8 | N/F |  | N/F | [13]C6H9[15]N3O2 | N/A | Internal Standard Compound has peaks that are not found.
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C9 | 7103761 |  | 89359890 | [13]C6H9[15]N3O2 | 30.81 | 11.879:Apex Retention Time 11.879 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C10 | 6959695 |  | 61291736 | [13]C6H9[15]N3O2 | 65.13 | 11.823:Apex Retention Time 11.823 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C11 | 6706982 |  | 91622467 | [13]C6H9[15]N3O2 | 45.78 | 11.791:Apex Retention Time 11.791 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)
Histidine U-13C, U-15N pos | 11.94 | Internal Standard | KD100520_100520_C12 | 8835732 |  | 152628781 | [13]C6H9[15]N3O2 | 106.35 | 11.892:Peak area 152628780.642 is out of bounds (ISTD Minimum Recovery 32497245.073 and ISTD Max Recovery 97491735.218). 11.892:Apex Retention Time 11.892 is out of bounds (ISTD Min RT 8.925 and ISTD Max RT 9.425)

## Isoleucine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A1 | 167761120 |  | 2235736675 | C6H13NO2 | 15194.91
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A2 | 82870489 |  | 1042381364 | C6H13NO2 | 6926.52
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A3 | 50588756 |  | 665691044 | C6H13NO2 | 16421.98
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A4 | 23609689 |  | 312173865 | C6H13NO2 | 14654.85
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A5 | 16557625 |  | 208681754 | C6H13NO2 | 9264.34 | Calibration Amount Diff 42.482% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A6 | 7189322 |  | 98293149 | C6H13NO2 | 3410.96
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A7 | 3192940 |  | 43890911 | C6H13NO2 | 1856.05
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A8 | 1508186 |  | 20683092 | C6H13NO2 | 1826.79 | Calibration Amount Diff -79.604% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A9 | 926815 |  | 11806607 | C6H13NO2 | 1630.99 | Calibration Amount Diff -175.689% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A10 | 391075 |  | 4624426 | C6H13NO2 | 407.62 | Calibration Amount Diff -407.507% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A11 | 231448 |  | 2631107 | C6H13NO2 | 403.21 | Calibration Amount Diff -824.169% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_A12 | 82278 |  | 1048967 | C6H13NO2 | 211.07 | Calibration Amount Diff -1711.478% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B1 | 179305351 |  | 2452521308 | C6H13NO2 | 12616.27
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B2 | 85840987 |  | 1124540917 | C6H13NO2 | 7967.58
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B3 | 45717440 |  | 635610340 | C6H13NO2 | 8735.51
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B4 | 23515700 |  | 321516545 | C6H13NO2 | 13271.94
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B5 | 11685546 |  | 151872530 | C6H13NO2 | 7915.35
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B6 | 6038801 |  | 81257146 | C6H13NO2 | 5467.83
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B7 | 3134431 |  | 41767842 | C6H13NO2 | 2261.53 | Calibration Amount Diff -26.166% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B8 | 1680804 |  | 22479088 | C6H13NO2 | 1738.18 | Calibration Amount Diff -62.899% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B9 | 831343 |  | 10978921 | C6H13NO2 | 926.23 | Calibration Amount Diff -177.938% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B10 | 369141 |  | 4438419 | C6H13NO2 | 614.56 | Calibration Amount Diff -411.475% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B11 | 184440 |  | 2090354 | C6H13NO2 | 363.60 | Calibration Amount Diff -839.58% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_B12 | 81508 |  | 906593 | C6H13NO2 | 264.35 | Calibration Amount Diff -1716.8% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H13NO2 | N/A
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C1 | 143126502 |  | 2161235650 | C6H13NO2 | 6968.75
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C2 | 78580875 |  | 1136323470 | C6H13NO2 | 8193.56
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C3 | 42405775 |  | 600471580 | C6H13NO2 | 3135.39
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C4 | 21729224 |  | 272697692 | C6H13NO2 | 6981.35
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C5 | 11234814 |  | 146487885 | C6H13NO2 | 5841.35
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C6 | 5424391 |  | 75397792 | C6H13NO2 | 3229.30
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C7 | 3075715 |  | 40919731 | C6H13NO2 | 2914.19 | Calibration Amount Diff -24.1% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C8 | 1512514 |  | 20403611 | C6H13NO2 | 1940.73 | Calibration Amount Diff -81.961% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C9 | 789629 |  | 10103153 | C6H13NO2 | 1031.48 | Calibration Amount Diff -188.845% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C10 | 334700 |  | 3691139 | C6H13NO2 | 397.12 | Calibration Amount Diff -431.183% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C11 | 160647 |  | 1994913 | C6H13NO2 | 290.08 | Calibration Amount Diff -855.17% (absolute) > max calibration amount diff 20%
Isoleucine neg | 7.05 | Target Compound | KD100520_100520_C12 | 64049 |  | 864529 | C6H13NO2 | 165.81 | Calibration Amount Diff -1719.173% (absolute) > max calibration amount diff 20%

## Isoleucine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A1 | 1201955953 |  | 17130485587 | C6H13NO2 | 9962.20
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A2 | 658694089 |  | 9405032982 | C6H13NO2 | 11233.32
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A3 | 523392778 |  | 7610355785 | C6H13NO2 | 31341.27
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A4 | 266820443 |  | 3874480117 | C6H13NO2 | 14358.23
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A5 | 189146353 |  | 2607654180 | C6H13NO2 | 23366.29 | Calibration Amount Diff 39.264% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A6 | 91801618 |  | 1271647514 | C6H13NO2 | 8872.42
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A7 | 42514622 |  | 559628232 | C6H13NO2 | 4923.28 | Calibration Amount Diff -21.068% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A8 | 21705134 |  | 281955596 | C6H13NO2 | 5292.86 | Calibration Amount Diff -72.523% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A9 | 12248058 |  | 161999422 | C6H13NO2 | 2589.03 | Calibration Amount Diff -161.619% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A10 | 6109276 |  | 74756411 | C6H13NO2 | 1642.58 | Calibration Amount Diff -399.512% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A11 | 3072978 |  | 39805454 | C6H13NO2 | 1050.01 | Calibration Amount Diff -820.843% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_A12 | 1296742 |  | 18158666 | C6H13NO2 | 471.92 | Calibration Amount Diff -1710.294% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B1 | 1251260269 |  | 18529623887 | C6H13NO2 | 24694.31
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B2 | 771504912 |  | 10461849194 | C6H13NO2 | 10573.61
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B3 | 540664862 |  | 7389443337 | C6H13NO2 | 27779.27
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B4 | 267055818 |  | 3867996355 | C6H13NO2 | 26762.01
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B5 | 136323703 |  | 1928755896 | C6H13NO2 | 16613.05
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B6 | 76092247 |  | 1081646390 | C6H13NO2 | 17055.49
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B7 | 41167461 |  | 549088927 | C6H13NO2 | 12245.97 | Calibration Amount Diff -21.995% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B8 | 21639048 |  | 281559375 | C6H13NO2 | 3159.17 | Calibration Amount Diff -67.083% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B9 | 11370948 |  | 151714013 | C6H13NO2 | 2097.59 | Calibration Amount Diff -170.816% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B10 | 6019047 |  | 71203647 | C6H13NO2 | 2413.95 | Calibration Amount Diff -398.821% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B11 | 2849366 |  | 36827808 | C6H13NO2 | 944.05 | Calibration Amount Diff -827.936% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_B12 | 1299772 |  | 16222367 | C6H13NO2 | 477.61 | Calibration Amount Diff -1715.743% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H13NO2 | N/A
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C1 | 1183051448 |  | 16669244299 | C6H13NO2 | 13417.66
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C2 | 746893933 |  | 10734465388 | C6H13NO2 | 8131.11
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C3 | 462473891 |  | 6514631903 | C6H13NO2 | 12685.46
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C4 | 243581928 |  | 3395247825 | C6H13NO2 | 7145.77
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C5 | 142415861 |  | 1911201539 | C6H13NO2 | 14219.93
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C6 | 71790524 |  | 947040217 | C6H13NO2 | 22047.57
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C7 | 44255012 |  | 608908828 | C6H13NO2 | 6667.84
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C8 | 21067849 |  | 277471937 | C6H13NO2 | 4125.63 | Calibration Amount Diff -69.59% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C9 | 10717735 |  | 146617758 | C6H13NO2 | 2925.12 | Calibration Amount Diff -181.889% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C10 | 4613981 |  | 62047735 | C6H13NO2 | 1102.64 | Calibration Amount Diff -417.524% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C11 | 2621595 |  | 33559811 | C6H13NO2 | 892.65 | Calibration Amount Diff -842.335% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD100520_100520_C12 | 1286191 |  | 17133755 | C6H13NO2 | 456.17 | Calibration Amount Diff -1706.469% (absolute) > max calibration amount diff 20%

## Isoleucine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A1 | 260761 |  | 3024818 | [13]C6H13NO2 | 4744.19 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -54.877% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A2 | 262398 |  | 3081800 | [13]C6H13NO2 | 450.40 | Rsquared 0.64 < R2Threshold 0.99
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A3 | 290793 |  | 2993562 | [13]C6H13NO2 | 536.97 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 66.602% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A4 | 250674 |  | 2857977 | [13]C6H13NO2 | 377.37 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 314.035% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A5 | 261130 |  | 2977817 | [13]C6H13NO2 | 347.43 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 453.899% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A6 | 276101 |  | 3475459 | [13]C6H13NO2 | 477.55 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -439.383% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A7 | 263445 |  | 3403081 | [13]C6H13NO2 | 549.67 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -2388.19% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A8 | 243985 |  | 3374666 | [13]C6H13NO2 | 300.06 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -3217.403% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A9 | 285558 |  | 3346746 | [13]C6H13NO2 | 562.96 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -260.274% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A10 | 267553 |  | 3150903 | [13]C6H13NO2 | 466.81 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -9023.092% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A11 | 263643 |  | 3108793 | [13]C6H13NO2 | 622.31 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 2984.852% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_A12 | 258411 |  | 3222533 | [13]C6H13NO2 | 392.67 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 17809.704% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B1 | 197131 |  | 2609159 | [13]C6H13NO2 | 204.61 | Rsquared 0.64 < R2Threshold 0.99
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B2 | 260528 |  | 2884699 | [13]C6H13NO2 | 373.77 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 36.959% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B3 | 259090 |  | 2918027 | [13]C6H13NO2 | 394.64 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 73.395% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B4 | 230337 |  | 2810952 | [13]C6H13NO2 | 429.86 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 287.724% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B5 | 253239 |  | 3241779 | [13]C6H13NO2 | 536.29 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -213.811% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B6 | 240919 |  | 3238245 | [13]C6H13NO2 | 365.86 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -503.251% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B7 | 277104 |  | 3456018 | [13]C6H13NO2 | 654.56 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -2035.473% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B8 | 261514 |  | 3133172 | [13]C6H13NO2 | 497.54 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -1446.11% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B9 | 235842 |  | 3229129 | [13]C6H13NO2 | 403.04 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -2314.988% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B10 | 270482 |  | 3039750 | [13]C6H13NO2 | 493.58 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -3193.433% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B11 | 231429 |  | 2710002 | [13]C6H13NO2 | 542.82 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 15399.576% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_B12 | 257737 |  | 2928564 | [13]C6H13NO2 | 579.12 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 21483.683% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H13NO2 | N/A
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C1 | 217266 |  | 2509811 | [13]C6H13NO2 | 1531.29 | Rsquared 0.64 < R2Threshold 0.99
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C2 | 204042 |  | 2585490 | [13]C6H13NO2 | 415.43 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 59.915% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C3 | 214841 |  | 2830744 | [13]C6H13NO2 | 274.02 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 116.247% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C4 | 257006 |  | 2692920 | [13]C6H13NO2 | 592.33 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 137.396% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C5 | 231669 |  | 2703484 | [13]C6H13NO2 | 291.67 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 180.602% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C6 | 256061 |  | 3354787 | [13]C6H13NO2 | 498.23 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -1003.096% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C7 | 242443 |  | 2984317 | [13]C6H13NO2 | 465.53 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 750.926% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C8 | 245131 |  | 3211720 | [13]C6H13NO2 | 526.23 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -418.637% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C9 | 260538 |  | 3311034 | [13]C6H13NO2 | 1059.28 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -4826.086% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C10 | 265029 |  | 3078831 | [13]C6H13NO2 | 698.05 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -5347.9% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C11 | 259793 |  | 3226983 | [13]C6H13NO2 | 607.08 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff -15073.051% (absolute) > max calibration amount diff 20%
Isoleucine U-13C neg | 7.05 | Target Compound | KD100520_100520_C12 | 205814 |  | 2798226 | [13]C6H13NO2 | 976.65 | Rsquared 0.64 < R2Threshold 0.99. Calibration Amount Diff 39443.048% (absolute) > max calibration amount diff 20%

## Isoleucine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A1 | 2167072 |  | 39659525 | [13]C6H13NO2 | 2511.49 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -353.972% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A2 | 2289753 |  | 32858248 | [13]C6H13NO2 | 1045.56 | Rsquared 0.125 < R2Threshold 0.99
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A3 | 3159172 |  | 42823119 | [13]C6H13NO2 | 1062.15 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 106.368% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A4 | 3234581 |  | 46243022 | [13]C6H13NO2 | 2511.20 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 57.717% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A5 | 3316466 |  | 45219430 | [13]C6H13NO2 | 2121.28 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 619.171% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A6 | 3660307 |  | 49170516 | [13]C6H13NO2 | 1873.73 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -509.952% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A7 | 3716641 |  | 47811583 | [13]C6H13NO2 | 2482.25 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -1295.961% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A8 | 3867363 |  | 48073636 | [13]C6H13NO2 | 2642.73 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -5021.54% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A9 | 3752725 |  | 49639098 | [13]C6H13NO2 | 2925.00 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -17429.536% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A10 | 4026618 |  | 50931074 | [13]C6H13NO2 | 4596.14 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -22846.756% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A11 | 3379420 |  | 43370009 | [13]C6H13NO2 | 3041.56 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 52106.368% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_A12 | 3124951 |  | 46412193 | [13]C6H13NO2 | 1998.66 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -11105.097% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B1 | 1678211 |  | 22727428 | [13]C6H13NO2 | 5384.97 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 118.806% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B2 | 2618696 |  | 31915460 | [13]C6H13NO2 | 1170.07 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 200.512% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B3 | 3467813 |  | 42031510 | [13]C6H13NO2 | 3209.16 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 124.282% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B4 | 2937749 |  | 38944503 | [13]C6H13NO2 | 1410.60 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 870.347% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B5 | 3257934 |  | 44470859 | [13]C6H13NO2 | 3130.76 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 589.082% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B6 | 3276330 |  | 46500130 | [13]C6H13NO2 | 3038.13 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 56.261% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B7 | 3574057 |  | 47834661 | [13]C6H13NO2 | 3178.65 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -2021.64% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B8 | 3393615 |  | 45397788 | [13]C6H13NO2 | 1986.91 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -3051.475% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B9 | 3322257 |  | 46663805 | [13]C6H13NO2 | 1535.51 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -4093.719% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B10 | 3899263 |  | 48501325 | [13]C6H13NO2 | 2154.32 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -25128.808% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B11 | 3263801 |  | 44751985 | [13]C6H13NO2 | 3569.54 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -2423.744% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_B12 | 3259598 |  | 42788203 | [13]C6H13NO2 | 1657.42 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 3321.878% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H13NO2 | N/A
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C1 | 1821226 |  | 22079723 | [13]C6H13NO2 | 716.00 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 130.141% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C2 | 2581919 |  | 32416559 | [13]C6H13NO2 | 1508.47 | Rsquared 0.125 < R2Threshold 0.99
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C3 | 2910097 |  | 35090237 | [13]C6H13NO2 | 1276.95 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 260.302% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C4 | 2994346 |  | 39014194 | [13]C6H13NO2 | 1494.63 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 262.162% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C5 | 3378459 |  | 43462269 | [13]C6H13NO2 | 3393.87 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -573.419% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C6 | 3572339 |  | 45884235 | [13]C6H13NO2 | 2457.97 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -1378.275% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C7 | 3769460 |  | 48042302 | [13]C6H13NO2 | 2209.22 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 1529.738% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C8 | 3480000 |  | 45520328 | [13]C6H13NO2 | 2501.26 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -2999.329% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C9 | 3693747 |  | 47786786 | [13]C6H13NO2 | 1765.43 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 410.081% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C10 | 3389446 |  | 47274972 | [13]C6H13NO2 | 3560.69 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -9690.955% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C11 | 3457768 |  | 46074787 | [13]C6H13NO2 | 3135.63 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff -20185.664% (absolute) > max calibration amount diff 20%
Isoleucine U-13C pos | 7.05 | Target Compound | KD100520_100520_C12 | 3249584 |  | 41299285 | [13]C6H13NO2 | 2119.93 | Rsquared 0.125 < R2Threshold 0.99. Calibration Amount Diff 45763.115% (absolute) > max calibration amount diff 20%

## Isoleucine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A1 | 24676295 |  | 301466296 | [13]C6H13[15]NO2 | 77284.46
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A2 | 24712163 |  | 308742320 | [13]C6H13[15]NO2 | 4357.93
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A3 | 23812877 |  | 295354439 | [13]C6H13[15]NO2 | 5883.47
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A4 | 21556118 |  | 290468059 | [13]C6H13[15]NO2 | 7531.11
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A5 | 22305898 |  | 287942813 | [13]C6H13[15]NO2 | 7240.74
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A6 | 21556745 |  | 297866518 | [13]C6H13[15]NO2 | 20629.35
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A7 | 20244856 |  | 274324337 | [13]C6H13[15]NO2 | 3613.74
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A8 | 18896757 |  | 279568840 | [13]C6H13[15]NO2 | 9816.49
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A9 | 21844468 |  | 294214292 | [13]C6H13[15]NO2 | 7015.42
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A10 | 20055967 |  | 265492987 | [13]C6H13[15]NO2 | 3920.03
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A11 | 20889239 |  | 275882204 | [13]C6H13[15]NO2 | 3613.31
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_A12 | 19365339 |  | 290303811 | [13]C6H13[15]NO2 | 10002.27
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B1 | 23870897 |  | 300590135 | [13]C6H13[15]NO2 | 25630.70
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B2 | 23481173 |  | 308564639 | [13]C6H13[15]NO2 | 5380.42
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B3 | 22162628 |  | 289323419 | [13]C6H13[15]NO2 | 23862.84
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B4 | 21004497 |  | 282915659 | [13]C6H13[15]NO2 | 11697.87
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B5 | 20117709 |  | 280299078 | [13]C6H13[15]NO2 | 11320.24
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B6 | 20191909 |  | 276150741 | [13]C6H13[15]NO2 | 15211.98
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B7 | 20278376 |  | 282269100 | [13]C6H13[15]NO2 | 3553.68
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B8 | 20966265 |  | 268591240 | [13]C6H13[15]NO2 | 5045.91
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B9 | 19564969 |  | 278120428 | [13]C6H13[15]NO2 | 4632.40
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B10 | 19725486 |  | 263544802 | [13]C6H13[15]NO2 | 10704.27
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B11 | 18296581 |  | 248322285 | [13]C6H13[15]NO2 | 7222.92
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_B12 | 20197381 |  | 265076833 | [13]C6H13[15]NO2 | 6670.97
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H13[15]NO2 | N/A
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C1 | 20906611 |  | 314316568 | [13]C6H13[15]NO2 | 7338.06
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C2 | 19880382 |  | 286888313 | [13]C6H13[15]NO2 | 6967.98
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C3 | 20496199 |  | 289692399 | [13]C6H13[15]NO2 | 10467.39
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C4 | 19741597 |  | 256794506 | [13]C6H13[15]NO2 | 10101.59
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C5 | 18216086 |  | 249339682 | [13]C6H13[15]NO2 | 30457.39
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C6 | 18950626 |  | 275334391 | [13]C6H13[15]NO2 | 5633.28
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C7 | 19925152 |  | 272119481 | [13]C6H13[15]NO2 | 5846.96
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C8 | 19721554 |  | 280994435 | [13]C6H13[15]NO2 | 6246.08
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C9 | 20271783 |  | 278281189 | [13]C6H13[15]NO2 | 6799.85
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C10 | 20368039 |  | 264106426 | [13]C6H13[15]NO2 | 2631.83
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C11 | 20753602 |  | 273812617 | [13]C6H13[15]NO2 | 6957.88
Isoleucine U-13C, U-15N neg | 7.05 | Internal Standard | KD100520_100520_C12 | 16684403 |  | 259311540 | [13]C6H13[15]NO2 | 1890.82

## Isoleucine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A1 | 182889672 |  | 2424721725 | [13]C6H13[15]NO2 | 19621.74
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A2 | 204751845 |  | 2771183287 | [13]C6H13[15]NO2 | 86736.46
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A3 | 258470964 |  | 3583573554 | [13]C6H13[15]NO2 | 35418.49
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A4 | 256859274 |  | 3726410625 | [13]C6H13[15]NO2 | 10829.43
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A5 | 266063370 |  | 3754005753 | [13]C6H13[15]NO2 | 36635.97
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A6 | 275584815 |  | 3818007432 | [13]C6H13[15]NO2 | 7242.87
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A7 | 271733678 |  | 3688216892 | [13]C6H13[15]NO2 | 10012.99
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A8 | 278147797 |  | 3628671968 | [13]C6H13[15]NO2 | 17197.15
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A9 | 273794470 |  | 3631252341 | [13]C6H13[15]NO2 | 11463.18
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A10 | 284536178 |  | 3819410824 | [13]C6H13[15]NO2 | 26228.75
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A11 | 259373387 |  | 3626543702 | [13]C6H13[15]NO2 | 77667.19
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_A12 | 251192224 |  | 3634105904 | [13]C6H13[15]NO2 | 8136.74
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B1 | 173460161 |  | 2382709894 | [13]C6H13[15]NO2 | 13856.52
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B2 | 223769437 |  | 3031765295 | [13]C6H13[15]NO2 | 70614.24
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B3 | 274752243 |  | 3536441848 | [13]C6H13[15]NO2 | 17236.60
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B4 | 236607077 |  | 3558301078 | [13]C6H13[15]NO2 | 18992.08
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B5 | 247989123 |  | 3683568160 | [13]C6H13[15]NO2 | 10438.03
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B6 | 249391549 |  | 3683640589 | [13]C6H13[15]NO2 | 17536.21
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B7 | 258477458 |  | 3644030187 | [13]C6H13[15]NO2 | 20155.24
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B8 | 256487014 |  | 3485077700 | [13]C6H13[15]NO2 | 22411.90
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B9 | 248100062 |  | 3612087264 | [13]C6H13[15]NO2 | 11703.60
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B10 | 280671589 |  | 3619744245 | [13]C6H13[15]NO2 | 23002.14
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B11 | 245483021 |  | 3516429398 | [13]C6H13[15]NO2 | 14683.64
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_B12 | 244302375 |  | 3377192645 | [13]C6H13[15]NO2 | 51036.72
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H13[15]NO2 | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C1 | 183744830 |  | 2355165974 | [13]C6H13[15]NO2 | 21588.86
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C2 | 207349359 |  | 2724125828 | [13]C6H13[15]NO2 | 25010.52
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C3 | 233803488 |  | 3079475749 | [13]C6H13[15]NO2 | 23526.16
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C4 | 227227324 |  | 3240111713 | [13]C6H13[15]NO2 | 7676.54
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C5 | 234075556 |  | 3312539645 | [13]C6H13[15]NO2 | 82351.42
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C6 | 251349557 |  | 3457770545 | [13]C6H13[15]NO2 | 13604.32
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C7 | 279798630 |  | 3897452128 | [13]C6H13[15]NO2 | 14162.83
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C8 | 259742547 |  | 3496061381 | [13]C6H13[15]NO2 | 43322.69
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C9 | 271945328 |  | 3773106478 | [13]C6H13[15]NO2 | 14746.57
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C10 | 253261413 |  | 3646608911 | [13]C6H13[15]NO2 | 34049.74
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C11 | 255400780 |  | 3550623224 | [13]C6H13[15]NO2 | 13719.14
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD100520_100520_C12 | 247379398 |  | 3338373070 | [13]C6H13[15]NO2 | 14086.86

## Leucine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Leucine neg | 6.74 | Target Compound | KD100520_100520_A1 | 176718001 |  | 2196741929 | C6H13NO2 | 1370087.86
Leucine neg | 6.74 | Target Compound | KD100520_100520_A2 | 91227984 |  | 1217616026 | C6H13NO2 | 177454.31
Leucine neg | 6.74 | Target Compound | KD100520_100520_A3 | 60991946 |  | 749280194 | C6H13NO2 | 131797.69
Leucine neg | 6.74 | Target Compound | KD100520_100520_A4 | 27014397 |  | 336442663 | C6H13NO2 | 66579.68
Leucine neg | 6.74 | Target Compound | KD100520_100520_A5 | 20028868 |  | 235251477 | C6H13NO2 | 61331.15 | Calibration Amount Diff 51.574% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_A6 | 8619033 |  | 105087351 | C6H13NO2 | 19971.10
Leucine neg | 6.74 | Target Compound | KD100520_100520_A7 | 4139655 |  | 48702745 | C6H13NO2 | 11933.66 | Calibration Amount Diff -53.845% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_A8 | 2106047 |  | 24987833 | C6H13NO2 | 1799.53 | Calibration Amount Diff -139.993% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_A9 | 996368 |  | 12702862 | C6H13NO2 | 2033.65 | Calibration Amount Diff -319.164% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_A10 | 549364 |  | 6246299 | C6H13NO2 | 628.39 | Calibration Amount Diff -716.568% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_A11 | 283456 |  | 3074038 | C6H13NO2 | 503.59 | Calibration Amount Diff -1466.115% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_A12 | 111017 |  | 1300269 | C6H13NO2 | 323.49 | Calibration Amount Diff -3019.591% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_B1 | 196342232 |  | 2463811490 | C6H13NO2 | 784590.45
Leucine neg | 6.74 | Target Compound | KD100520_100520_B2 | 87878367 |  | 1226530648 | C6H13NO2 | 202835.67
Leucine neg | 6.74 | Target Compound | KD100520_100520_B3 | 50797992 |  | 673480753 | C6H13NO2 | 101563.60
Leucine neg | 6.74 | Target Compound | KD100520_100520_B4 | 29246691 |  | 359211695 | C6H13NO2 | 91675.16 | Calibration Amount Diff 24.494% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_B5 | 12834239 |  | 168229770 | C6H13NO2 | 24303.19
Leucine neg | 6.74 | Target Compound | KD100520_100520_B6 | 7797908 |  | 96812427 | C6H13NO2 | 14572.30
Leucine neg | 6.74 | Target Compound | KD100520_100520_B7 | 3792040 |  | 46680081 | C6H13NO2 | 22452.20 | Calibration Amount Diff -47.581% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_B8 | 1914173 |  | 23190326 | C6H13NO2 | 2092.47 | Calibration Amount Diff -143.833% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_B9 | 1017426 |  | 12506565 | C6H13NO2 | 1359.19 | Calibration Amount Diff -317.633% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_B10 | 493808 |  | 5461498 | C6H13NO2 | 1346.41 | Calibration Amount Diff -716.968% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_B11 | 239521 |  | 2690852 | C6H13NO2 | 549.54 | Calibration Amount Diff -1480.288% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_B12 | 91272 |  | 1132490 | C6H13NO2 | 224.94 | Calibration Amount Diff -3019.556% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H13NO2 | N/A
Leucine neg | 6.74 | Target Compound | KD100520_100520_C1 | 185168795 |  | 2312007163 | C6H13NO2 | 543994.77
Leucine neg | 6.74 | Target Compound | KD100520_100520_C2 | 98910796 |  | 1283718571 | C6H13NO2 | 55162.00
Leucine neg | 6.74 | Target Compound | KD100520_100520_C3 | 48547345 |  | 622307681 | C6H13NO2 | 759586.01
Leucine neg | 6.74 | Target Compound | KD100520_100520_C4 | 26878090 |  | 325238474 | C6H13NO2 | 208031.50 | Calibration Amount Diff 20.652% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_C5 | 14378824 |  | 181155114 | C6H13NO2 | 2361516.36
Leucine neg | 6.74 | Target Compound | KD100520_100520_C6 | 5885602 |  | 78907485 | C6H13NO2 | 17181.59
Leucine neg | 6.74 | Target Compound | KD100520_100520_C7 | 3944343 |  | 46704548 | C6H13NO2 | 7792.14 | Calibration Amount Diff -43.257% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_C8 | 1828680 |  | 22068252 | C6H13NO2 | 3338.32 | Calibration Amount Diff -144.535% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_C9 | 951572 |  | 11319467 | C6H13NO2 | 2623.25 | Calibration Amount Diff -329.437% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_C10 | 398070 |  | 4870789 | C6H13NO2 | 532.75 | Calibration Amount Diff -735.353% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_C11 | 225938 |  | 2426758 | C6H13NO2 | 567.73 | Calibration Amount Diff -1499.446% (absolute) > max calibration amount diff 20%
Leucine neg | 6.74 | Target Compound | KD100520_100520_C12 | 78140 |  | 931124 | C6H13NO2 | 159.99 | Calibration Amount Diff -3035.13% (absolute) > max calibration amount diff 20%

## Leucine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Leucine pos | 6.74 | Target Compound | KD100520_100520_A1 | 1108324252 |  | 15169810769 | C6H13NO2 | 541170.01
Leucine pos | 6.74 | Target Compound | KD100520_100520_A2 | 712291437 |  | 9525945794 | C6H13NO2 | 105439.40
Leucine pos | 6.74 | Target Compound | KD100520_100520_A3 | 476360574 |  | 6120794045 | C6H13NO2 | 55809.57
Leucine pos | 6.74 | Target Compound | KD100520_100520_A4 | 273105706 |  | 3357440667 | C6H13NO2 | 30927.74 | Calibration Amount Diff 23.979% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_A5 | 183333385 |  | 2303546831 | C6H13NO2 | 25020.53 | Calibration Amount Diff 64.088% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_A6 | 89461644 |  | 1116802160 | C6H13NO2 | 20777.49 | Calibration Amount Diff 30.132% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_A7 | 42475515 |  | 560693122 | C6H13NO2 | 10146.10 | Calibration Amount Diff -40.601% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_A8 | 22233351 |  | 272035334 | C6H13NO2 | 5645.77 | Calibration Amount Diff -149.273% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_A9 | 11833182 |  | 145458832 | C6H13NO2 | 2292.92 | Calibration Amount Diff -373.758% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_A10 | 5492561 |  | 67286354 | C6H13NO2 | 1930.62 | Calibration Amount Diff -836.632% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_A11 | 2997999 |  | 38386782 | C6H13NO2 | 1078.71 | Calibration Amount Diff -1724.535% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_A12 | 1495295 |  | 17579557 | C6H13NO2 | 363.68 | Calibration Amount Diff -3588.13% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_B1 | 1207736458 |  | 16386405560 | C6H13NO2 | 925959.67
Leucine pos | 6.74 | Target Compound | KD100520_100520_B2 | 714723868 |  | 10109461642 | C6H13NO2 | 63055.14
Leucine pos | 6.74 | Target Compound | KD100520_100520_B3 | 475929902 |  | 6293286090 | C6H13NO2 | 48034.15
Leucine pos | 6.74 | Target Compound | KD100520_100520_B4 | 270458460 |  | 3524574230 | C6H13NO2 | 27169.87 | Calibration Amount Diff 31.743% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_B5 | 143985414 |  | 1777382374 | C6H13NO2 | 38813.14
Leucine pos | 6.74 | Target Compound | KD100520_100520_B6 | 74439131 |  | 950392404 | C6H13NO2 | 20054.88
Leucine pos | 6.74 | Target Compound | KD100520_100520_B7 | 37149923 |  | 500895843 | C6H13NO2 | 7669.19 | Calibration Amount Diff -44.753% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_B8 | 19099036 |  | 260572731 | C6H13NO2 | 7018.73 | Calibration Amount Diff -141.476% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_B9 | 10981029 |  | 134923898 | C6H13NO2 | 2593.43 | Calibration Amount Diff -358.123% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_B10 | 5060427 |  | 68020303 | C6H13NO2 | 2299.60 | Calibration Amount Diff -834.037% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_B11 | 2712882 |  | 33046898 | C6H13NO2 | 582.12 | Calibration Amount Diff -1740.344% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_B12 | 1361783 |  | 16959853 | C6H13NO2 | 291.05 | Calibration Amount Diff -3579.816% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H13NO2 | N/A
Leucine pos | 6.74 | Target Compound | KD100520_100520_C1 | 1117838223 |  | 15278038306 | C6H13NO2 | 547975.62
Leucine pos | 6.74 | Target Compound | KD100520_100520_C2 | 722637494 |  | 9689597625 | C6H13NO2 | 53275.11
Leucine pos | 6.74 | Target Compound | KD100520_100520_C3 | 452145519 |  | 5677235630 | C6H13NO2 | 79605.88
Leucine pos | 6.74 | Target Compound | KD100520_100520_C4 | 264480826 |  | 3315407104 | C6H13NO2 | 49289.87 | Calibration Amount Diff 25.865% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_C5 | 141728345 |  | 1895693054 | C6H13NO2 | 35233.81 | Calibration Amount Diff 20.267% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_C6 | 71373340 |  | 898015297 | C6H13NO2 | 12345.72
Leucine pos | 6.74 | Target Compound | KD100520_100520_C7 | 36643581 |  | 489769798 | C6H13NO2 | 10426.77 | Calibration Amount Diff -37.88% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_C8 | 21021232 |  | 276174639 | C6H13NO2 | 5705.99 | Calibration Amount Diff -143.92% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_C9 | 9583307 |  | 122516308 | C6H13NO2 | 3098.98 | Calibration Amount Diff -375.572% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_C10 | 4437198 |  | 58978638 | C6H13NO2 | 1848.80 | Calibration Amount Diff -857.431% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_C11 | 2540078 |  | 32004742 | C6H13NO2 | 697.22 | Calibration Amount Diff -1758.015% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD100520_100520_C12 | 1306566 |  | 15030123 | C6H13NO2 | 411.89 | Calibration Amount Diff -3596.803% (absolute) > max calibration amount diff 20%

## Leucine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A1 | 185893 |  | 1876138 | [13]C6H13NO2 | INF | Rsquared 0.56 < R2Threshold 0.99
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A2 | 203492 |  | 2246982 | [13]C6H13NO2 | 1124.70 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 44.777% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A3 | 231195 |  | 2509208 | [13]C6H13NO2 | 1708.92 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 21.452% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A4 | 185597 |  | 2062064 | [13]C6H13NO2 | 951.13 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 157.991% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A5 | 199823 |  | 1994636 | [13]C6H13NO2 | 355.30 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 909.154% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A6 | 199760 |  | 2055819 | [13]C6H13NO2 | 407.51 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 679.117% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A7 | 227884 |  | 2389578 | [13]C6H13NO2 | 613.78 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -832.881% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A8 | 205744 |  | 2255629 | [13]C6H13NO2 | 323.17 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -108.36% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A9 | 198268 |  | 2329210 | [13]C6H13NO2 | 518.29 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -12009.513% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A10 | 212265 |  | 2512871 | [13]C6H13NO2 | 636.25 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -18137.968% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A11 | 191951 |  | 2148657 | [13]C6H13NO2 | 1234.13 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -4617.772% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_A12 | 218902 |  | 2477052 | [13]C6H13NO2 | 407.43 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -110323.123% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B1 | 193011 |  | 2122241 | [13]C6H13NO2 | 197.75 | Rsquared 0.56 < R2Threshold 0.99
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B2 | 166301 |  | 2056192 | [13]C6H13NO2 | 713.30 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 61.295% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B3 | 163253 |  | 2141908 | [13]C6H13NO2 | 721.49 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 130.191% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B4 | 200444 |  | 2061906 | [13]C6H13NO2 | 3356.50 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 213.36% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B5 | 175496 |  | 2020132 | [13]C6H13NO2 | 497.84 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 519.273% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B6 | 197790 |  | 2260397 | [13]C6H13NO2 | 435.12 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -22.683% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B7 | 205647 |  | 2235570 | [13]C6H13NO2 | 737.38 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -1483.305% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B8 | 182790 |  | 2152534 | [13]C6H13NO2 | 500.78 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -238.558% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B9 | 181068 |  | 2012243 | [13]C6H13NO2 | 965.71 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 4356.139% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B10 | 189900 |  | 2180168 | [13]C6H13NO2 | 580.24 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -15195.107% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B11 | 180392 |  | 2169501 | [13]C6H13NO2 | 898.03 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -30668.595% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_B12 | 148117 |  | 1902488 | [13]C6H13NO2 | 654.26 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 26706.367% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H13NO2 | N/A
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C1 | 185341 |  | 2092413 | [13]C6H13NO2 | 2727.46 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -20.46% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C2 | 222490 |  | 2326059 | [13]C6H13NO2 | 1336.35 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -83.316% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C3 | 190606 |  | 2131443 | [13]C6H13NO2 | 336.91 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -105.023% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C4 | 185382 |  | 1968963 | [13]C6H13NO2 | 887.94 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 118.702% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C5 | 198624 |  | 2300795 | [13]C6H13NO2 | 367.93 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -488.063% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C6 | 166953 |  | 1942977 | [13]C6H13NO2 | 574.05 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 122.673% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C7 | 197047 |  | 2005710 | [13]C6H13NO2 | 511.96 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 1152.659% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C8 | 172094 |  | 2016031 | [13]C6H13NO2 | 912.14 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 1112.243% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C9 | 185204 |  | 1996738 | [13]C6H13NO2 | 632.86 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 2212.134% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C10 | 171864 |  | 2040694 | [13]C6H13NO2 | 904.03 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 8362.241% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C11 | 190013 |  | 2168676 | [13]C6H13NO2 | 971.72 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff -8045.167% (absolute) > max calibration amount diff 20%
Leucine U-13C neg | 6.74 | Target Compound | KD100520_100520_C12 | 147571 |  | 1766653 | [13]C6H13NO2 | 1557.20 | Rsquared 0.56 < R2Threshold 0.99. Calibration Amount Diff 37216.716% (absolute) > max calibration amount diff 20%

## Leucine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A1 | N/F |  | N/F | [13]C6H13NO2 | N/A | Rsquared 0.233 < R2Threshold 0.99
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A2 | 1976085 |  | 19406679 | [13]C6H13NO2 | 15460.01 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 101.603% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A3 | 1954047 |  | 23419774 | [13]C6H13NO2 | INF | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 110.589% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A4 | 2189636 |  | 23697168 | [13]C6H13NO2 | 84316.19 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 215.718% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A5 | 2152744 |  | 24736440 | [13]C6H13NO2 | 10805.11 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 245.795% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A6 | 2118997 |  | 25613671 | [13]C6H13NO2 | 13679.50 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 197.763% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A7 | 2347203 |  | 28906517 | [13]C6H13NO2 | 3881.27 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff -444.952% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A8 | 2249530 |  | 26942473 | [13]C6H13NO2 | 16920.02 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff -616.059% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A9 | 2218370 |  | 27664296 | [13]C6H13NO2 | 2784.50 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff -327.756% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A10 | 2187216 |  | 25455805 | [13]C6H13NO2 | 9048.24 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 2890.6% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A11 | 1813827 |  | 24684197 | [13]C6H13NO2 | 24202.53 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 3066.309% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_A12 | 1978680 |  | 23725956 | [13]C6H13NO2 | 6675.63 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 40077.817% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B1 | 1174860 |  | 15320088 | [13]C6H13NO2 | INF | Rsquared 0.233 < R2Threshold 0.99
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B2 | 1635596 |  | 20640897 | [13]C6H13NO2 | 54915.97 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 22.789% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B3 | 2010874 |  | 25793160 | [13]C6H13NO2 | INF | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 51.599% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B4 | 2068893 |  | 26556364 | [13]C6H13NO2 | 5837.13 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff -85.887% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B5 | 2231441 |  | 25588249 | [13]C6H13NO2 | 31125.04 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 204.351% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B6 | 1897508 |  | 23168974 | [13]C6H13NO2 | 10735.77 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 834.475% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B7 | 1971753 |  | 25294087 | [13]C6H13NO2 | 32843.21 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 510.402% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B8 | 1942954 |  | 24603225 | [13]C6H13NO2 | 5562.71 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff -373.542% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B9 | 2028960 |  | 22496594 | [13]C6H13NO2 | 2154.94 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 3968.285% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B10 | 2023850 |  | 25004478 | [13]C6H13NO2 | 5396.63 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 5355.326% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B11 | 1910901 |  | 21479762 | [13]C6H13NO2 | 34369.28 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 24296.975% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_B12 | 1772137 |  | 21798754 | [13]C6H13NO2 | 8863.39 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 45011.202% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H13NO2 | N/A
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C1 | 1272209 |  | 16124074 | [13]C6H13NO2 | INF | Rsquared 0.233 < R2Threshold 0.99
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C2 | 1581532 |  | 19957100 | [13]C6H13NO2 | 66929.15 | Rsquared 0.233 < R2Threshold 0.99
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C3 | 1998318 |  | 24232606 | [13]C6H13NO2 | 13820.17 | Rsquared 0.233 < R2Threshold 0.99
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C4 | 2989449 |  | 38886213 | [13]C6H13NO2 | 9989.50 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff -1404.377% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C5 | 3373998 |  | 43332683 | [13]C6H13NO2 | 32863.13 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff -2992.074% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C6 | 2110176 |  | 24446036 | [13]C6H13NO2 | 15448.40 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 242.705% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C7 | 1863249 |  | 24247531 | [13]C6H13NO2 | 2809.64 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 94.246% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C8 | 2050232 |  | 24647447 | [13]C6H13NO2 | 14975.12 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 2522.347% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C9 | 1851430 |  | 22004756 | [13]C6H13NO2 | 1812.24 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 5250.374% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C10 | 1873465 |  | 24050900 | [13]C6H13NO2 | 100042.68 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 10311.904% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C11 | 1967696 |  | 23726338 | [13]C6H13NO2 | 18289.25 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 11183.172% (absolute) > max calibration amount diff 20%
Leucine U-13C pos | 6.74 | Target Compound | KD100520_100520_C12 | 1727664 |  | 21289421 | [13]C6H13NO2 | 10383.59 | Rsquared 0.233 < R2Threshold 0.99. Calibration Amount Diff 37822.57% (absolute) > max calibration amount diff 20%

## Leucine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A1 | 26374594 |  | 331184034 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A2 | 27894880 |  | 354463843 | [13]C6H13[15]NO2 | 2235417.95
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A3 | 28743271 |  | 361383367 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A4 | 24585562 |  | 298148580 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A5 | 26488146 |  | 308144477 | [13]C6H13[15]NO2 | 4672586.90
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A6 | 23737716 |  | 292441955 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A7 | 25430352 |  | 316728639 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A8 | 25084843 |  | 305605875 | [13]C6H13[15]NO2 | 6867516.78
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A9 | 22506273 |  | 289387884 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A10 | 25256457 |  | 318622486 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A11 | 23246382 |  | 288669893 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_A12 | 24878556 |  | 303810575 | [13]C6H13[15]NO2 | 5972245.87
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B1 | 27128636 |  | 341421345 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B2 | 24730170 |  | 330558656 | [13]C6H13[15]NO2 | 2698188.28
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B3 | 23842795 |  | 326885029 | [13]C6H13[15]NO2 | 2688748.57
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B4 | 25565295 |  | 302477780 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B5 | 22948853 |  | 296059039 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B6 | 24641818 |  | 307739502 | [13]C6H13[15]NO2 | 1342089.14
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B7 | 22980816 |  | 290685234 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B8 | 22670119 |  | 291061313 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B9 | 22484959 |  | 282236072 | [13]C6H13[15]NO2 | 5491684.90
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B10 | 22681150 |  | 279365696 | [13]C6H13[15]NO2 | 5444492.61
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B11 | 22529788 |  | 277825227 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_B12 | 19241034 |  | 264535058 | [13]C6H13[15]NO2 | 5764505.62
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H13[15]NO2 | N/A
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C1 | 27630507 |  | 335522284 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C2 | 25180398 |  | 320395130 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C3 | 23664106 |  | 288112204 | [13]C6H13[15]NO2 | 705519.83
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C4 | 23541161 |  | 281810103 | [13]C6H13[15]NO2 | 125299.67
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C5 | 24240425 |  | 297697797 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C6 | 19648162 |  | 266897804 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C7 | 23299178 |  | 282555325 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C8 | 22096971 |  | 278313428 | [13]C6H13[15]NO2 | 3801292.49
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C9 | 21673885 |  | 275410067 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C10 | 21997745 |  | 285724996 | [13]C6H13[15]NO2 | 6842751.10
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C11 | 22473236 |  | 289488831 | [13]C6H13[15]NO2 | 680912.62
Leucine U-13C, U-15N neg | 6.74 | Internal Standard | KD100520_100520_C12 | 19667354 |  | 248192595 | [13]C6H13[15]NO2 | INF

## Leucine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A1 | 180304787 |  | 2530678550 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A2 | 243713020 |  | 3237638269 | [13]C6H13[15]NO2 | 6713986.51
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A3 | 237138884 |  | 3241628882 | [13]C6H13[15]NO2 | 2637510.76
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A4 | 260104068 |  | 3133891493 | [13]C6H13[15]NO2 | 1749974.39
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A5 | 249635038 |  | 3085186797 | [13]C6H13[15]NO2 | 2663330.07
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A6 | 248895852 |  | 3074013506 | [13]C6H13[15]NO2 | 1147006.44
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A7 | 248941434 |  | 3319585979 | [13]C6H13[15]NO2 | 1901414.30
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A8 | 253144737 |  | 3106425093 | [13]C6H13[15]NO2 | 13935962.61
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A9 | 242713328 |  | 3219454088 | [13]C6H13[15]NO2 | 1677404.84
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A10 | 243323590 |  | 3022973982 | [13]C6H13[15]NO2 | 2749566.58
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A11 | 214885580 |  | 2907051949 | [13]C6H13[15]NO2 | 14111941.72
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_A12 | 227344278 |  | 2940362249 | [13]C6H13[15]NO2 | 18341291.74
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B1 | 178663710 |  | 2486455803 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B2 | 212891889 |  | 2948218181 | [13]C6H13[15]NO2 | 11276840.63
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B3 | 242396593 |  | 3393135614 | [13]C6H13[15]NO2 | 12631505.88
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B4 | 245140119 |  | 3115102420 | [13]C6H13[15]NO2 | 9307609.07
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B5 | 258568788 |  | 3165293896 | [13]C6H13[15]NO2 | 1322689.22
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B6 | 233847604 |  | 2961330518 | [13]C6H13[15]NO2 | 20251125.73
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B7 | 220828665 |  | 3037829960 | [13]C6H13[15]NO2 | 7793852.67
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B8 | 212708792 |  | 2852637732 | [13]C6H13[15]NO2 | 6609729.52
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B9 | 226578856 |  | 2755576642 | [13]C6H13[15]NO2 | 3865988.67
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B10 | 221981669 |  | 3013420986 | [13]C6H13[15]NO2 | 8196432.24
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B11 | 216301647 |  | 2697918081 | [13]C6H13[15]NO2 | 1966333.36
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_B12 | 207650108 |  | 2722280485 | [13]C6H13[15]NO2 | 14591242.59
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H13[15]NO2 | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C1 | 189314196 |  | 2551788029 | [13]C6H13[15]NO2 | 6679764.83
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C2 | 199304616 |  | 2818075510 | [13]C6H13[15]NO2 | INF
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C3 | 240637838 |  | 3074671399 | [13]C6H13[15]NO2 | 7538153.18
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C4 | 250278999 |  | 3053045694 | [13]C6H13[15]NO2 | 9902284.45
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C5 | 232875274 |  | 3285421846 | [13]C6H13[15]NO2 | 19658150.03 | 6.988:Apex Retention Time 6.988 is out of bounds (ISTD Min RT 6.486 and ISTD Max RT 6.986)
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C6 | 236901426 |  | 2946572487 | [13]C6H13[15]NO2 | 23403262.49
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C7 | 215997941 |  | 2855153212 | [13]C6H13[15]NO2 | 1653171.59
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C8 | 238835222 |  | 3063090056 | [13]C6H13[15]NO2 | 12833653.13
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C9 | 214406453 |  | 2738256418 | [13]C6H13[15]NO2 | 1701602.95
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C10 | 214774200 |  | 2987616427 | [13]C6H13[15]NO2 | 7849686.08
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C11 | 215052070 |  | 2862494708 | [13]C6H13[15]NO2 | 2641571.67
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD100520_100520_C12 | 200357132 |  | 2629237611 | [13]C6H13[15]NO2 | 7459170.04

## Lysine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Lysine neg | 13.78 | Target Compound | KD100520_100520_A1 | 19718152 |  | 312437727 | C6H14N2O2 | 664252.78
Lysine neg | 13.78 | Target Compound | KD100520_100520_A2 | 10966018 |  | 176062776 | C6H14N2O2 | 273086.37
Lysine neg | 13.78 | Target Compound | KD100520_100520_A3 | 7715061 |  | 129795271 | C6H14N2O2 | 7111.68
Lysine neg | 13.78 | Target Compound | KD100520_100520_A4 | 3992345 |  | 66935207 | C6H14N2O2 | 2818.29
Lysine neg | 13.78 | Target Compound | KD100520_100520_A5 | 2875119 |  | 49504568 | C6H14N2O2 | 1177.95 | Calibration Amount Diff 35.362% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_A6 | 1323552 |  | 22661192 | C6H14N2O2 | 1304.28
Lysine neg | 13.78 | Target Compound | KD100520_100520_A7 | 559657 |  | 9068781 | C6H14N2O2 | 764.17
Lysine neg | 13.78 | Target Compound | KD100520_100520_A8 | 274065 |  | 4235747 | C6H14N2O2 | 781.08 | Calibration Amount Diff -42.22% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_A9 | 124739 |  | 1976219 | C6H14N2O2 | 390.23 | Calibration Amount Diff -88.142% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_A10 | 44794 |  | 754411 | C6H14N2O2 | 216.36 | Calibration Amount Diff -188.061% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_A11 | 24790 |  | 402730 | C6H14N2O2 | 544.99 | Calibration Amount Diff -331.231% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_A12 | 10783 |  | 170840 | C6H14N2O2 | INF | Calibration Amount Diff -648.663% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_B1 | 20468488 |  | 338111565 | C6H14N2O2 | 29613.10
Lysine neg | 13.78 | Target Compound | KD100520_100520_B2 | 12509048 |  | 194244812 | C6H14N2O2 | 68198.45
Lysine neg | 13.78 | Target Compound | KD100520_100520_B3 | 8363067 |  | 127954321 | C6H14N2O2 | 68767.63
Lysine neg | 13.78 | Target Compound | KD100520_100520_B4 | 4470387 |  | 73951502 | C6H14N2O2 | 2066.52
Lysine neg | 13.78 | Target Compound | KD100520_100520_B5 | 2140246 |  | 35091248 | C6H14N2O2 | 1077.81
Lysine neg | 13.78 | Target Compound | KD100520_100520_B6 | 1042503 |  | 18291576 | C6H14N2O2 | 799.70
Lysine neg | 13.78 | Target Compound | KD100520_100520_B7 | 472003 |  | 8201093 | C6H14N2O2 | 729.68
Lysine neg | 13.78 | Target Compound | KD100520_100520_B8 | 246743 |  | 3889805 | C6H14N2O2 | 826.51 | Calibration Amount Diff -43.992% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_B9 | 117044 |  | 1771864 | C6H14N2O2 | 668.30 | Calibration Amount Diff -90.713% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_B10 | 46317 |  | 738553 | C6H14N2O2 | 363.43 | Calibration Amount Diff -182.664% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_B11 | 25661 |  | 411422 | C6H14N2O2 | 1137.08 | Calibration Amount Diff -329.156% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_B12 | 9113 |  | 165781 | C6H14N2O2 | INF | Calibration Amount Diff -645.059% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H14N2O2 | N/A
Lysine neg | 13.78 | Target Compound | KD100520_100520_C1 | 19732190 |  | 317399009 | C6H14N2O2 | 415609.40
Lysine neg | 13.78 | Target Compound | KD100520_100520_C2 | 11368895 |  | 184683842 | C6H14N2O2 | 35485.09
Lysine neg | 13.78 | Target Compound | KD100520_100520_C3 | 6879361 |  | 107877882 | C6H14N2O2 | 75893.10
Lysine neg | 13.78 | Target Compound | KD100520_100520_C4 | 3142889 |  | 58476609 | C6H14N2O2 | 1689.96
Lysine neg | 13.78 | Target Compound | KD100520_100520_C5 | 1795357 |  | 31508147 | C6H14N2O2 | 1402.74
Lysine neg | 13.78 | Target Compound | KD100520_100520_C6 | 868575 |  | 14593956 | C6H14N2O2 | 818.57
Lysine neg | 13.78 | Target Compound | KD100520_100520_C7 | 518839 |  | 8052750 | C6H14N2O2 | 1525.90
Lysine neg | 13.78 | Target Compound | KD100520_100520_C8 | 241294 |  | 3837286 | C6H14N2O2 | 601.35 | Calibration Amount Diff -46.299% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_C9 | 89872 |  | 1454740 | C6H14N2O2 | 345.59 | Calibration Amount Diff -101.466% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_C10 | 28543 |  | 453359 | C6H14N2O2 | 1857.53 | Calibration Amount Diff -202.934% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_C11 | 19628 |  | 269551 | C6H14N2O2 | INF | Calibration Amount Diff -350.483% (absolute) > max calibration amount diff 20%
Lysine neg | 13.78 | Target Compound | KD100520_100520_C12 | 8520 |  | 116946 | C6H14N2O2 | INF | Calibration Amount Diff -658.74% (absolute) > max calibration amount diff 20%

## Lysine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Lysine pos | 13.78 | Target Compound | KD100520_100520_A1 | 246293155 |  | 4855026739 | C6H14N2O2 | 93127.14
Lysine pos | 13.78 | Target Compound | KD100520_100520_A2 | 146584553 |  | 2856354363 | C6H14N2O2 | 119016.43
Lysine pos | 13.78 | Target Compound | KD100520_100520_A3 | 102655102 |  | 1838400858 | C6H14N2O2 | 92190.66
Lysine pos | 13.78 | Target Compound | KD100520_100520_A4 | 60751928 |  | 1068603027 | C6H14N2O2 | 21962.89
Lysine pos | 13.78 | Target Compound | KD100520_100520_A5 | 40408154 |  | 710868071 | C6H14N2O2 | 17483.63 | Calibration Amount Diff 47.428% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_A6 | 17706400 |  | 315262601 | C6H14N2O2 | 6846.42
Lysine pos | 13.78 | Target Compound | KD100520_100520_A7 | 8070217 |  | 136135945 | C6H14N2O2 | 4231.37 | Calibration Amount Diff -33.003% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_A8 | 4278603 |  | 70111862 | C6H14N2O2 | 4040.02 | Calibration Amount Diff -107.385% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_A9 | 2016687 |  | 36950227 | C6H14N2O2 | 1854.16 | Calibration Amount Diff -245.366% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_A10 | 999963 |  | 17134588 | C6H14N2O2 | 809.13 | Calibration Amount Diff -559.724% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_A11 | 584011 |  | 9437678 | C6H14N2O2 | 560.64 | Calibration Amount Diff -1139.779% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_A12 | 248503 |  | 3926582 | C6H14N2O2 | 247.89 | Calibration Amount Diff -2370.345% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_B1 | 270880008 |  | 5328264135 | C6H14N2O2 | 114911.60
Lysine pos | 13.78 | Target Compound | KD100520_100520_B2 | 175500462 |  | 3228573344 | C6H14N2O2 | 115679.45
Lysine pos | 13.78 | Target Compound | KD100520_100520_B3 | 101030971 |  | 1878924167 | C6H14N2O2 | 115393.11
Lysine pos | 13.78 | Target Compound | KD100520_100520_B4 | 61376608 |  | 1082701673 | C6H14N2O2 | 32014.25
Lysine pos | 13.78 | Target Compound | KD100520_100520_B5 | 29308727 |  | 548852811 | C6H14N2O2 | 14889.28
Lysine pos | 13.78 | Target Compound | KD100520_100520_B6 | 16604583 |  | 293676578 | C6H14N2O2 | 5727.19
Lysine pos | 13.78 | Target Compound | KD100520_100520_B7 | 7991202 |  | 143754091 | C6H14N2O2 | 4054.52 | Calibration Amount Diff -34.192% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_B8 | 4387542 |  | 77876965 | C6H14N2O2 | 2770.98 | Calibration Amount Diff -100.784% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_B9 | 2038856 |  | 38068491 | C6H14N2O2 | 955.39 | Calibration Amount Diff -245.486% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_B10 | 1100923 |  | 18273629 | C6H14N2O2 | 1071.96 | Calibration Amount Diff -558.193% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_B11 | 494901 |  | 9050018 | C6H14N2O2 | 631.01 | Calibration Amount Diff -1151.118% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_B12 | 303721 |  | 3995167 | C6H14N2O2 | 355.39 | Calibration Amount Diff -2373.025% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C6H14N2O2 | N/A
Lysine pos | 13.78 | Target Compound | KD100520_100520_C1 | 294160259 |  | 5323255168 | C6H14N2O2 | 293908.73
Lysine pos | 13.78 | Target Compound | KD100520_100520_C2 | 174132757 |  | 3009002714 | C6H14N2O2 | 89998.21
Lysine pos | 13.78 | Target Compound | KD100520_100520_C3 | 100819609 |  | 1900365163 | C6H14N2O2 | 80361.51
Lysine pos | 13.78 | Target Compound | KD100520_100520_C4 | 53902671 |  | 1018535652 | C6H14N2O2 | 40613.61
Lysine pos | 13.78 | Target Compound | KD100520_100520_C5 | 29619836 |  | 554579515 | C6H14N2O2 | 15354.94
Lysine pos | 13.78 | Target Compound | KD100520_100520_C6 | 14152127 |  | 253940090 | C6H14N2O2 | 3677.78
Lysine pos | 13.78 | Target Compound | KD100520_100520_C7 | 8964441 |  | 153177163 | C6H14N2O2 | 3949.53 | Calibration Amount Diff -31.666% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_C8 | 4016137 |  | 73158384 | C6H14N2O2 | 2858.67 | Calibration Amount Diff -113.432% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_C9 | 2301403 |  | 39798634 | C6H14N2O2 | 1075.81 | Calibration Amount Diff -255.062% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_C10 | 889355 |  | 15364085 | C6H14N2O2 | 906.25 | Calibration Amount Diff -579.957% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_C11 | 537191 |  | 8342456 | C6H14N2O2 | 593.68 | Calibration Amount Diff -1168.938% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD100520_100520_C12 | 259073 |  | 4106539 | C6H14N2O2 | 341.79 | Calibration Amount Diff -2368.174% (absolute) > max calibration amount diff 20%

## Lysine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A1 | 123121 |  | 1595410 | [13]C6H14N2O2 | 1450.07 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -27.412% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A2 | 123541 |  | 1699277 | [13]C6H14N2O2 | 493.32 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 126.485% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A3 | 175142 |  | 2647706 | [13]C6H14N2O2 | 523.08 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 88.039% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A4 | 219214 |  | 2896343 | [13]C6H14N2O2 | 585.26 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 134.637% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A5 | 216322 |  | 3377254 | [13]C6H14N2O2 | 437.02 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 201.55% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A6 | 258167 |  | 4200113 | [13]C6H14N2O2 | 662.39 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -1463.283% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A7 | 238254 |  | 3584514 | [13]C6H14N2O2 | 868.93 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -1158.846% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A8 | 256906 |  | 3900546 | [13]C6H14N2O2 | 681.02 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -4312.973% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A9 | 260259 |  | 4068228 | [13]C6H14N2O2 | 633.23 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -11234.49% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A10 | 241190 |  | 4171737 | [13]C6H14N2O2 | 566.79 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -18992.088% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A11 | 207724 |  | 3438747 | [13]C6H14N2O2 | 824.29 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 6843.197% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_A12 | 218689 |  | 3590722 | [13]C6H14N2O2 | 829.44 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -9931.735% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B1 | 116123 |  | 1602652 | [13]C6H14N2O2 | 970.83 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -34.913% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B2 | 142318 |  | 1985670 | [13]C6H14N2O2 | 1826.48 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 49.944% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B3 | 218950 |  | 2821785 | [13]C6H14N2O2 | 1004.43 | Rsquared 0.617 < R2Threshold 0.99
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B4 | 223978 |  | 3120318 | [13]C6H14N2O2 | 600.10 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 104.815% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B5 | 256648 |  | 3448554 | [13]C6H14N2O2 | 839.22 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -74.159% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B6 | 203955 |  | 3293548 | [13]C6H14N2O2 | 678.40 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 112.719% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B7 | 198774 |  | 3303693 | [13]C6H14N2O2 | 775.04 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -332.494% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B8 | 207219 |  | 3149935 | [13]C6H14N2O2 | 620.81 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 2435.285% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B9 | 215800 |  | 3529531 | [13]C6H14N2O2 | 658.20 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -5097.431% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B10 | 195419 |  | 3570304 | [13]C6H14N2O2 | 638.93 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -8494.113% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B11 | 217014 |  | 3690893 | [13]C6H14N2O2 | 807.62 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -21475.916% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_B12 | 191598 |  | 3296275 | [13]C6H14N2O2 | 641.77 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff -10411.268% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H14N2O2 | N/A
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C1 | 113168 |  | 1502637 | [13]C6H14N2O2 | 521.74 | Rsquared 0.617 < R2Threshold 0.99
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C2 | 116177 |  | 1805311 | [13]C6H14N2O2 | 631.26 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 25.016% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C3 | 192743 |  | 2347717 | [13]C6H14N2O2 | 1059.85 | Rsquared 0.617 < R2Threshold 0.99
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C4 | 151339 |  | 2477654 | [13]C6H14N2O2 | 1281.70 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 160.542% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C5 | 181543 |  | 2629295 | [13]C6H14N2O2 | 652.10 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 256.657% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C6 | 184323 |  | 2899328 | [13]C6H14N2O2 | 1106.46 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 42.85% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C7 | 208834 |  | 2950790 | [13]C6H14N2O2 | 631.48 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 1363.866% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C8 | 220144 |  | 3345535 | [13]C6H14N2O2 | 547.20 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 296.453% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C9 | 188065 |  | 3035857 | [13]C6H14N2O2 | 788.83 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 2891.913% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C10 | 178169 |  | 2810833 | [13]C6H14N2O2 | 648.08 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 8329.048% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C11 | 218332 |  | 3033792 | [13]C6H14N2O2 | 687.62 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 15142.908% (absolute) > max calibration amount diff 20%
Lysine U-13C neg | 13.78 | Target Compound | KD100520_100520_C12 | 170439 |  | 2700625 | [13]C6H14N2O2 | 526.23 | Rsquared 0.617 < R2Threshold 0.99. Calibration Amount Diff 47699.633% (absolute) > max calibration amount diff 20%

## Lysine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A1 | 2309511 |  | 43651880 | [13]C6H14N2O2 | INF | Rsquared 0.506 < R2Threshold 0.99
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A2 | 2915421 |  | 54447964 | [13]C6H14N2O2 | 109107.79 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 52.581% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A3 | 3255871 |  | 56471495 | [13]C6H14N2O2 | 3100.45 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 111.855% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A4 | 3845648 |  | 66100232 | [13]C6H14N2O2 | 6303.94 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 48.736% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A5 | 3828601 |  | 64927800 | [13]C6H14N2O2 | 7646.44 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 313.003% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A6 | 3582899 |  | 64257963 | [13]C6H14N2O2 | 2314.66 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -603.171% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A7 | 3531041 |  | 60918999 | [13]C6H14N2O2 | 2426.85 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -1757.652% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A8 | 3777014 |  | 62813941 | [13]C6H14N2O2 | 3970.18 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -4313.884% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A9 | 3220300 |  | 60159131 | [13]C6H14N2O2 | 2363.87 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -4542.564% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A10 | 3411941 |  | 61656581 | [13]C6H14N2O2 | 2561.57 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -11452.559% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A11 | 3076497 |  | 55908794 | [13]C6H14N2O2 | 2334.43 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 36504.659% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_A12 | 3330754 |  | 56578295 | [13]C6H14N2O2 | 3305.79 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 61033.95% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B1 | 2260606 |  | 42637530 | [13]C6H14N2O2 | 79075.41 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 24.283% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B2 | 3512455 |  | 60475151 | [13]C6H14N2O2 | 11404.74 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -59.041% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B3 | 3311146 |  | 60214319 | [13]C6H14N2O2 | 4274.03 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 57.937% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B4 | 3831668 |  | 65177337 | [13]C6H14N2O2 | 3923.22 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -99.963% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B5 | 3798883 |  | 69353195 | [13]C6H14N2O2 | 2070.76 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -202.059% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B6 | 3809375 |  | 66330932 | [13]C6H14N2O2 | 2582.68 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -177.832% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B7 | 3538738 |  | 63689851 | [13]C6H14N2O2 | 2463.90 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -422.011% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B8 | 3586035 |  | 64585899 | [13]C6H14N2O2 | 1847.55 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 341.541% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B9 | 3196085 |  | 60738754 | [13]C6H14N2O2 | 1530.34 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 1510.456% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B10 | 3623402 |  | 63644457 | [13]C6H14N2O2 | 2004.20 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 1165.051% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B11 | 3021340 |  | 57419905 | [13]C6H14N2O2 | 2359.40 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 40575.873% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_B12 | 3624984 |  | 58590065 | [13]C6H14N2O2 | 1641.36 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 66642.363% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C6H14N2O2 | N/A
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C1 | 2964491 |  | 49262018 | [13]C6H14N2O2 | 93271.30 | Rsquared 0.506 < R2Threshold 0.99
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C2 | 3300491 |  | 53077087 | [13]C6H14N2O2 | 9234.44 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -99.056% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C3 | 3611032 |  | 65043096 | [13]C6H14N2O2 | 5116.17 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -235.435% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C4 | 3467873 |  | 65957221 | [13]C6H14N2O2 | 3590.79 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -419.379% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C5 | 3622117 |  | 66349688 | [13]C6H14N2O2 | 2570.10 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -743.788% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C6 | 3517089 |  | 63601936 | [13]C6H14N2O2 | 1936.89 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff -1465.905% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C7 | 3875509 |  | 65310552 | [13]C6H14N2O2 | 2346.31 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 1045.787% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C8 | 3575296 |  | 64879888 | [13]C6H14N2O2 | 2488.25 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 3366.456% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C9 | 3840291 |  | 67031284 | [13]C6H14N2O2 | 2760.27 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 4251.742% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C10 | 3480371 |  | 61386083 | [13]C6H14N2O2 | 2441.10 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 17287.798% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C11 | 3613864 |  | 60337674 | [13]C6H14N2O2 | 2420.49 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 34312.728% (absolute) > max calibration amount diff 20%
Lysine U-13C pos | 13.78 | Target Compound | KD100520_100520_C12 | 3285671 |  | 58319195 | [13]C6H14N2O2 | 3545.60 | Rsquared 0.506 < R2Threshold 0.99. Calibration Amount Diff 57637.249% (absolute) > max calibration amount diff 20%

## Lysine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A1 | 2456243 |  | 35780343 | [13]C6H14[15]N2O2 | 10142.73
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A2 | 2801521 |  | 44590571 | [13]C6H14[15]N2O2 | 4818.50
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A3 | 3407579 |  | 54405313 | [13]C6H14[15]N2O2 | 4262.04
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A4 | 3397986 |  | 56256477 | [13]C6H14[15]N2O2 | 3249.18
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A5 | 3679395 |  | 63535214 | [13]C6H14[15]N2O2 | 2215.96
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A6 | 3723572 |  | 66720341 | [13]C6H14[15]N2O2 | 2270.23
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A7 | 3493570 |  | 60963860 | [13]C6H14[15]N2O2 | 804.39
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A8 | 3562525 |  | 63520357 | [13]C6H14[15]N2O2 | 994.06
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A9 | 3557468 |  | 64480195 | [13]C6H14[15]N2O2 | 1368.82
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A10 | 3494204 |  | 67237941 | [13]C6H14[15]N2O2 | 1630.57
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A11 | 3206384 |  | 62432208 | [13]C6H14[15]N2O2 | 1664.73
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_A12 | 3505121 |  | 63078425 | [13]C6H14[15]N2O2 | 1729.41
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B1 | 2198145 |  | 35004082 | [13]C6H14[15]N2O2 | 2200.94
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B2 | 2986489 |  | 44915556 | [13]C6H14[15]N2O2 | 3801.97
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B3 | 3486059 |  | 53944625 | [13]C6H14[15]N2O2 | 5646.95
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B4 | 3598340 |  | 59915269 | [13]C6H14[15]N2O2 | 4902.94
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B5 | 3562891 |  | 61688747 | [13]C6H14[15]N2O2 | 2002.28
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B6 | 3134646 |  | 59773720 | [13]C6H14[15]N2O2 | 1276.98
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B7 | 3153662 |  | 58226113 | [13]C6H14[15]N2O2 | 1809.36
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B8 | 3179948 |  | 59429602 | [13]C6H14[15]N2O2 | 3040.89
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B9 | 3245952 |  | 59548531 | [13]C6H14[15]N2O2 | 1418.83
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B10 | 3076962 |  | 60745800 | [13]C6H14[15]N2O2 | 997.49
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B11 | 3372876 |  | 62046851 | [13]C6H14[15]N2O2 | 1857.52
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_B12 | 2888647 |  | 57867937 | [13]C6H14[15]N2O2 | 1657.66
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H14[15]N2O2 | N/A
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C1 | 2415873 |  | 37592728 | [13]C6H14[15]N2O2 | 2861.55
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C2 | 2492413 |  | 39079607 | [13]C6H14[15]N2O2 | 4305.84
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C3 | 2866600 |  | 44476504 | [13]C6H14[15]N2O2 | 4920.95
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C4 | 2606892 |  | 48611546 | [13]C6H14[15]N2O2 | 1170.53
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C5 | 2647000 |  | 49980321 | [13]C6H14[15]N2O2 | 938.50
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C6 | 2884159 |  | 52288830 | [13]C6H14[15]N2O2 | 1685.68
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C7 | 3278366 |  | 56190936 | [13]C6H14[15]N2O2 | 1385.52
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C8 | 3226743 |  | 60099900 | [13]C6H14[15]N2O2 | 1151.66
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C9 | 2880688 |  | 55912005 | [13]C6H14[15]N2O2 | 1991.25
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C10 | 2845014 |  | 52505126 | [13]C6H14[15]N2O2 | 1176.52
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C11 | 3255554 |  | 56402306 | [13]C6H14[15]N2O2 | 2079.65
Lysine U-13C, U-15N neg | 13.78 | Internal Standard | KD100520_100520_C12 | 2606052 |  | 51497763 | [13]C6H14[15]N2O2 | 2196.37

## Lysine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A1 | 36330206 |  | 662559097 | [13]C6H14[15]N2O2 | 112029.50
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A2 | 42356613 |  | 796883799 | [13]C6H14[15]N2O2 | 63708.48
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A3 | 45282024 |  | 808436445 | [13]C6H14[15]N2O2 | 44128.62
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A4 | 53003695 |  | 916757366 | [13]C6H14[15]N2O2 | 62613.59
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A5 | 52907839 |  | 906437411 | [13]C6H14[15]N2O2 | 31256.16
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A6 | 49910119 |  | 864285735 | [13]C6H14[15]N2O2 | 24376.65
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A7 | 46399071 |  | 812099734 | [13]C6H14[15]N2O2 | 9903.64
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A8 | 49724332 |  | 832266222 | [13]C6H14[15]N2O2 | 17187.50
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A9 | 45445555 |  | 807997115 | [13]C6H14[15]N2O2 | 24269.17
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A10 | 45463943 |  | 824608259 | [13]C6H14[15]N2O2 | 26403.46
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A11 | 42139185 |  | 787660083 | [13]C6H14[15]N2O2 | 10929.04
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_A12 | 46033484 |  | 792784837 | [13]C6H14[15]N2O2 | 19822.58
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B1 | 35645799 |  | 654098245 | [13]C6H14[15]N2O2 | 264543.04
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B2 | 45837765 |  | 840180959 | [13]C6H14[15]N2O2 | 95079.44
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B3 | 45505363 |  | 851267352 | [13]C6H14[15]N2O2 | 44545.88
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B4 | 49858375 |  | 888959410 | [13]C6H14[15]N2O2 | 29913.82
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B5 | 50008442 |  | 940555248 | [13]C6H14[15]N2O2 | 34199.23
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B6 | 48411258 |  | 902729870 | [13]C6H14[15]N2O2 | 9284.61
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B7 | 49740006 |  | 864783876 | [13]C6H14[15]N2O2 | 6714.93
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B8 | 45930525 |  | 883609749 | [13]C6H14[15]N2O2 | 11519.00
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B9 | 44657379 |  | 833095638 | [13]C6H14[15]N2O2 | 22149.64
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B10 | 49603422 |  | 869968386 | [13]C6H14[15]N2O2 | 13865.45
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B11 | 43957233 |  | 811923675 | [13]C6H14[15]N2O2 | 10636.26
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_B12 | 49337908 |  | 823039798 | [13]C6H14[15]N2O2 | 8248.64
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C6H14[15]N2O2 | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C1 | 43381855 |  | 741171882 | [13]C6H14[15]N2O2 | 346508.27
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C2 | 44835650 |  | 724224973 | [13]C6H14[15]N2O2 | 98052.71
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C3 | 47913530 |  | 861100051 | [13]C6H14[15]N2O2 | 22275.49
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C4 | 44636472 |  | 868638471 | [13]C6H14[15]N2O2 | 17693.14
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C5 | 45039190 |  | 873570477 | [13]C6H14[15]N2O2 | 14918.46
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C6 | 45719102 |  | 835630670 | [13]C6H14[15]N2O2 | 7141.79
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C7 | 52397605 |  | 905239784 | [13]C6H14[15]N2O2 | 43802.24
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C8 | 49860464 |  | 906818497 | [13]C6H14[15]N2O2 | 13150.42
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C9 | 52846947 |  | 928326985 | [13]C6H14[15]N2O2 | 7443.70
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C10 | 47934703 |  | 863432321 | [13]C6H14[15]N2O2 | 27570.13
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C11 | 49478740 |  | 848382069 | [13]C6H14[15]N2O2 | 9072.32
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD100520_100520_C12 | 44789289 |  | 815938337 | [13]C6H14[15]N2O2 | 16171.27

## Lysine_cali
Compound | Filename | Area | Area_IS | Response ratio | Conc.
Lysine pos | KD100520_100520_A1 | 4855026739 | 662559097 | 7.3277 | 500
Lysine pos | KD100520_100520_A2 | 2856354363 | 796883799 | 3.5844 | 250
Lysine pos | KD100520_100520_A3 | 1838400858 | 808436445 | 2.274 | 125
Lysine pos | KD100520_100520_A4 | 1068603027 | 916757366 | 1.1656 | 62.5
Lysine pos | KD100520_100520_A5 | 710868071 | 906437411 | 0.78424 | 31.25
Lysine pos | KD100520_100520_A6 | 315262601 | 864285735 | 0.36477 | 15.625
Lysine pos | KD100520_100520_A7 | 136135945 | 812099734 | 0.16763 | 7.8125
Lysine pos | KD100520_100520_A8 | 70111862 | 832266222 | 0.084242 | 3.9062
Lysine pos | KD100520_100520_A9 | 36950227 | 807997115 | 0.045731 | 1.9531
Lysine pos | KD100520_100520_A10 | 17134588 | 824608259 | 0.020779 | 0.97656
Lysine pos | KD100520_100520_A11 | 9437678 | 787660083 | 0.011982 | 0.48828
Lysine pos | KD100520_100520_A12 | 3926582 | 792784837 | 0.0049529 | 0.24414
Lysine pos | KD100520_100520_B1 | 5328264135 | 654098245 | 8.146 | 500
Lysine pos | KD100520_100520_B2 | 3228573344 | 840180959 | 3.8427 | 250
Lysine pos | KD100520_100520_B3 | 1878924167 | 851267352 | 2.2072 | 125
Lysine pos | KD100520_100520_B4 | 1082701673 | 888959410 | 1.2179 | 62.5
Lysine pos | KD100520_100520_B5 | 548852811 | 940555248 | 0.58354 | 31.25
Lysine pos | KD100520_100520_B6 | 293676578 | 902729870 | 0.32532 | 15.625
Lysine pos | KD100520_100520_B7 | 143754091 | 864783876 | 0.16623 | 7.8125
Lysine pos | KD100520_100520_B8 | 77876965 | 883609749 | 0.088135 | 3.9062
Lysine pos | KD100520_100520_B9 | 38068491 | 833095638 | 0.045695 | 1.9531
Lysine pos | KD100520_100520_B10 | 18273629 | 869968386 | 0.021005 | 0.97656
Lysine pos | KD100520_100520_B11 | 9050018 | 811923675 | 0.011146 | 0.48828
Lysine pos | KD100520_100520_B12 | 3995167 | 823039798 | 0.0048542 | 0.24414
Lysine pos | KD100520_100520_C1 | 5323255168 | 741171882 | 7.1822 | 500
Lysine pos | KD100520_100520_C2 | 3009002714 | 724224973 | 4.1548 | 250
Lysine pos | KD100520_100520_C3 | 1900365163 | 861100051 | 2.2069 | 125
Lysine pos | KD100520_100520_C4 | 1018535652 | 868638471 | 1.1726 | 62.5
Lysine pos | KD100520_100520_C5 | 554579515 | 873570477 | 0.63484 | 31.25
Lysine pos | KD100520_100520_C6 | 253940090 | 835630670 | 0.30389 | 15.625
Lysine pos | KD100520_100520_C7 | 153177163 | 905239784 | 0.16921 | 7.8125
Lysine pos | KD100520_100520_C8 | 73158384 | 906818497 | 0.080676 | 3.9062
Lysine pos | KD100520_100520_C9 | 39798634 | 928326985 | 0.042871 | 1.9531
Lysine pos | KD100520_100520_C10 | 15364085 | 863432321 | 0.017794 | 0.97656
Lysine pos | KD100520_100520_C11 | 8342456 | 848382069 | 0.0098334 | 0.48828
Lysine pos | KD100520_100520_C12 | 4106539 | 815938337 | 0.0050329 | 0.24414

## Methionine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Methionine neg | 7.14 | Target Compound | KD100520_100520_A1 | 150847205 |  | 1968999177 | C5H11NO2S | 32836.94
Methionine neg | 7.14 | Target Compound | KD100520_100520_A2 | 72813598 |  | 1030096871 | C5H11NO2S | 49074.70
Methionine neg | 7.14 | Target Compound | KD100520_100520_A3 | 47220496 |  | 663911781 | C5H11NO2S | 40406.74
Methionine neg | 7.14 | Target Compound | KD100520_100520_A4 | 26403646 |  | 326639945 | C5H11NO2S | 44958.40
Methionine neg | 7.14 | Target Compound | KD100520_100520_A5 | 16959543 |  | 227181611 | C5H11NO2S | 15936.55 | Calibration Amount Diff 41.588% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_A6 | 8458289 |  | 110013708 | C5H11NO2S | 28680.94
Methionine neg | 7.14 | Target Compound | KD100520_100520_A7 | 4089979 |  | 48085064 | C5H11NO2S | 11123.36 | Calibration Amount Diff -24.042% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_A8 | 2045943 |  | 23920117 | C5H11NO2S | 5351.41 | Calibration Amount Diff -75.057% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_A9 | 1017911 |  | 12413919 | C5H11NO2S | 10088.18 | Calibration Amount Diff -169.014% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_A10 | 395006 |  | 4905598 | C5H11NO2S | 2919.22 | Calibration Amount Diff -386.862% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_A11 | 229595 |  | 2633869 | C5H11NO2S | 28165.05 | Calibration Amount Diff -780.531% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_A12 | 96208 |  | 1068460 | C5H11NO2S | INF | Calibration Amount Diff -1601.04% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_B1 | 159620443 |  | 2121840266 | C5H11NO2S | 67957.68
Methionine neg | 7.14 | Target Compound | KD100520_100520_B2 | 80675365 |  | 1079532175 | C5H11NO2S | 102727.45
Methionine neg | 7.14 | Target Compound | KD100520_100520_B3 | 52852465 |  | 653970142 | C5H11NO2S | 50100.78
Methionine neg | 7.14 | Target Compound | KD100520_100520_B4 | 23505494 |  | 328795963 | C5H11NO2S | 104587.21
Methionine neg | 7.14 | Target Compound | KD100520_100520_B5 | 13340098 |  | 166658356 | C5H11NO2S | 14998.43
Methionine neg | 7.14 | Target Compound | KD100520_100520_B6 | 6575171 |  | 84490587 | C5H11NO2S | 27836.12
Methionine neg | 7.14 | Target Compound | KD100520_100520_B7 | 3496672 |  | 43312616 | C5H11NO2S | 16363.42 | Calibration Amount Diff -28.01% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_B8 | 1726525 |  | 21871438 | C5H11NO2S | 6910.26 | Calibration Amount Diff -72.787% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_B9 | 930255 |  | 11138204 | C5H11NO2S | 4231.39 | Calibration Amount Diff -167.553% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_B10 | 401558 |  | 4412721 | C5H11NO2S | 22886.00 | Calibration Amount Diff -388.955% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_B11 | 165068 |  | 1975529 | C5H11NO2S | 8839.57 | Calibration Amount Diff -796.621% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_B12 | 73991 |  | 879053 | C5H11NO2S | 36999.69 | Calibration Amount Diff -1604.899% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H11NO2S | N/A
Methionine neg | 7.14 | Target Compound | KD100520_100520_C1 | 151149988 |  | 1931539847 | C5H11NO2S | 76982.96
Methionine neg | 7.14 | Target Compound | KD100520_100520_C2 | 79780924 |  | 1052691285 | C5H11NO2S | 57296.03
Methionine neg | 7.14 | Target Compound | KD100520_100520_C3 | 42534783 |  | 577226656 | C5H11NO2S | 63473.54
Methionine neg | 7.14 | Target Compound | KD100520_100520_C4 | 22577238 |  | 288602791 | C5H11NO2S | 32427.70
Methionine neg | 7.14 | Target Compound | KD100520_100520_C5 | 11637690 |  | 147808430 | C5H11NO2S | 15512.07
Methionine neg | 7.14 | Target Compound | KD100520_100520_C6 | 6079990 |  | 76956209 | C5H11NO2S | 15807.35
Methionine neg | 7.14 | Target Compound | KD100520_100520_C7 | 3195045 |  | 41481489 | C5H11NO2S | 9230.84 | Calibration Amount Diff -23.424% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_C8 | 1493951 |  | 19137600 | C5H11NO2S | 4298.67 | Calibration Amount Diff -81.329% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_C9 | 846993 |  | 10153059 | C5H11NO2S | 4643.56 | Calibration Amount Diff -179.726% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_C10 | 279102 |  | 3634638 | C5H11NO2S | 3136.71 | Calibration Amount Diff -403.526% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_C11 | 148174 |  | 1771310 | C5H11NO2S | 4936.38 | Calibration Amount Diff -808.106% (absolute) > max calibration amount diff 20%
Methionine neg | 7.14 | Target Compound | KD100520_100520_C12 | 64047 |  | 775473 | C5H11NO2S | 5615.78 | Calibration Amount Diff -1608.683% (absolute) > max calibration amount diff 20%

## Methionine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Methionine pos | 7.14 | Target Compound | KD100520_100520_A1 | 1384235027 |  | 19943405565 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_A2 | 896262074 |  | 12358349157 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_A3 | 664301694 |  | 9231609696 | C5H11NO2S | 1485256.65
Methionine pos | 7.14 | Target Compound | KD100520_100520_A4 | 364255192 |  | 5026795575 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_A5 | 263823697 |  | 3426832970 | C5H11NO2S | INF | Calibration Amount Diff 44.781% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_A6 | 116160131 |  | 1685469378 | C5H11NO2S | 773435.89
Methionine pos | 7.14 | Target Compound | KD100520_100520_A7 | 61538377 |  | 791301868 | C5H11NO2S | INF | Calibration Amount Diff -22.93% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_A8 | 27712423 |  | 380137680 | C5H11NO2S | 43501.79 | Calibration Amount Diff -85.377% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_A9 | 17108056 |  | 219524795 | C5H11NO2S | 37617.23 | Calibration Amount Diff -194.913% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_A10 | 7496038 |  | 98044444 | C5H11NO2S | 10432.17 | Calibration Amount Diff -450.74% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_A11 | 4341118 |  | 52448564 | C5H11NO2S | 8152.13 | Calibration Amount Diff -928.097% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_A12 | 1894428 |  | 22257200 | C5H11NO2S | 5979.00 | Calibration Amount Diff -1932.645% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_B1 | 1440500307 |  | 21377541312 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_B2 | 960619229 |  | 13135964645 | C5H11NO2S | 2027260.11
Methionine pos | 7.14 | Target Compound | KD100520_100520_B3 | 661820726 |  | 9079853341 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_B4 | 393428117 |  | 5187059904 | C5H11NO2S | 716138.09
Methionine pos | 7.14 | Target Compound | KD100520_100520_B5 | 205700385 |  | 2666033063 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_B6 | 114648170 |  | 1451012370 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_B7 | 56718215 |  | 740536425 | C5H11NO2S | 57754.71 | Calibration Amount Diff -25.883% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_B8 | 30859535 |  | 393466742 | C5H11NO2S | 30598.97 | Calibration Amount Diff -78.137% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_B9 | 15418391 |  | 202304216 | C5H11NO2S | 28012.31 | Calibration Amount Diff -197.288% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_B10 | 7519546 |  | 95761684 | C5H11NO2S | 21563.88 | Calibration Amount Diff -454.327% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_B11 | 3742138 |  | 45771458 | C5H11NO2S | 9262.02 | Calibration Amount Diff -938.564% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_B12 | 1703686 |  | 18846436 | C5H11NO2S | 9630.42 | Calibration Amount Diff -1938.568% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H11NO2S | N/A
Methionine pos | 7.14 | Target Compound | KD100520_100520_C1 | 1227383391 |  | 19061562647 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_C2 | 946266237 |  | 13238417927 | C5H11NO2S | 634642.00
Methionine pos | 7.14 | Target Compound | KD100520_100520_C3 | 526360471 |  | 7830206427 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_C4 | 342088297 |  | 4596632139 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_C5 | 194349987 |  | 2628002181 | C5H11NO2S | INF
Methionine pos | 7.14 | Target Compound | KD100520_100520_C6 | 98238060 |  | 1262918222 | C5H11NO2S | 3963551.23
Methionine pos | 7.14 | Target Compound | KD100520_100520_C7 | 56190283 |  | 802491987 | C5H11NO2S | 87940.79 | Calibration Amount Diff -21.052% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_C8 | 25598241 |  | 354734828 | C5H11NO2S | 58691.93 | Calibration Amount Diff -87.021% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_C9 | 15367786 |  | 197119996 | C5H11NO2S | 21527.74 | Calibration Amount Diff -205.422% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_C10 | 6361139 |  | 80728699 | C5H11NO2S | 11895.76 | Calibration Amount Diff -473.085% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_C11 | 3234921 |  | 41979837 | C5H11NO2S | 6519.63 | Calibration Amount Diff -956.215% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD100520_100520_C12 | 1428295 |  | 19196729 | C5H11NO2S | 7941.75 | Calibration Amount Diff -1935.379% (absolute) > max calibration amount diff 20%

## Methionine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A1 | 199712 |  | 2195155 | [13]C5H11NO2S | INF | Rsquared 0.611 < R2Threshold 0.99
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A2 | 203920 |  | 2647448 | [13]C5H11NO2S | INF | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 34.793% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A3 | 215731 |  | 2729606 | [13]C5H11NO2S | 57657.12 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 187.757% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A4 | 276125 |  | 2933581 | [13]C5H11NO2S | 13937.48 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 234.803% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A5 | 257980 |  | 3148753 | [13]C5H11NO2S | 3230.10 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 455.392% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A6 | 335770 |  | 4024009 | [13]C5H11NO2S | 4668.28 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -1006.588% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A7 | 370144 |  | 3909913 | [13]C5H11NO2S | 29858.69 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -1866.882% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A8 | 364824 |  | 3996095 | [13]C5H11NO2S | 3586.07 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -4413.342% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A9 | 330295 |  | 3990340 | [13]C5H11NO2S | 6894.06 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -9064.386% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A10 | 298714 |  | 3577653 | [13]C5H11NO2S | 7570.15 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -11004.515% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A11 | 292038 |  | 3722816 | [13]C5H11NO2S | 8148.01 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -23616.027% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_A12 | 324474 |  | 3796514 | [13]C5H11NO2S | 5816.51 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -27223.729% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B1 | 171904 |  | 2426663 | [13]C5H11NO2S | INF | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -40.941% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B2 | 212128 |  | 2595202 | [13]C5H11NO2S | 1731.66 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 26.125% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B3 | 287421 |  | 2846125 | [13]C5H11NO2S | 1938.38 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 162.228% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B4 | 229982 |  | 2820492 | [13]C5H11NO2S | 3695.24 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 250.718% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B5 | 337984 |  | 3626354 | [13]C5H11NO2S | 3843.56 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -240.622% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B6 | 302458 |  | 3381984 | [13]C5H11NO2S | 52564.24 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -563.287% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B7 | 296154 |  | 3326401 | [13]C5H11NO2S | 38011.67 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 207.47% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B8 | 288444 |  | 3351097 | [13]C5H11NO2S | 86186.44 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -1147.713% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B9 | 288879 |  | 3219986 | [13]C5H11NO2S | 16661.49 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -92.42% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B10 | 270067 |  | 3030331 | [13]C5H11NO2S | 5888.09 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 3575.524% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B11 | 243681 |  | 2975075 | [13]C5H11NO2S | 4795.20 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 8893.716% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_B12 | 271877 |  | 3308516 | [13]C5H11NO2S | 4567.95 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -37010.092% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H11NO2S | N/A
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C1 | 201810 |  | 2306199 | [13]C5H11NO2S | INF | Rsquared 0.611 < R2Threshold 0.99
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C2 | 185821 |  | 2306015 | [13]C5H11NO2S | 885.98 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 33.874% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C3 | 234703 |  | 2503347 | [13]C5H11NO2S | 3950.66 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 148.773% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C4 | 246984 |  | 2640292 | [13]C5H11NO2S | 13347.24 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 193.381% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C5 | 238998 |  | 2641324 | [13]C5H11NO2S | 2665.92 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 296.789% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C6 | 284103 |  | 3232777 | [13]C5H11NO2S | 6023.66 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -261.807% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C7 | 252977 |  | 3019084 | [13]C5H11NO2S | 10628.35 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 576.495% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C8 | 247804 |  | 2966645 | [13]C5H11NO2S | 5334.98 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 1462.283% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C9 | 294275 |  | 3453129 | [13]C5H11NO2S | 4203.91 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -6150.185% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C10 | 230262 |  | 3000254 | [13]C5H11NO2S | 43968.19 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -3469.955% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C11 | 263165 |  | 3102660 | [13]C5H11NO2S | 6978.34 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff -628.149% (absolute) > max calibration amount diff 20%
Methionine U-13C neg | 7.14 | Target Compound | KD100520_100520_C12 | 230867 |  | 2827362 | [13]C5H11NO2S | 6027.80 | Rsquared 0.611 < R2Threshold 0.99. Calibration Amount Diff 19562.408% (absolute) > max calibration amount diff 20%

## Methionine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A1 | 2071254 |  | 24919809 | [13]C5H11NO2S | INF | Rsquared 0.865 < R2Threshold 0.99
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A2 | 2855871 |  | 35428668 | [13]C5H11NO2S | 32830.22 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 35.704% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A3 | 3868702 |  | 46590019 | [13]C5H11NO2S | 7457.39 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 85.093% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A4 | 4131500 |  | 53031676 | [13]C5H11NO2S | 16066.04 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 151.151% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A5 | 4620580 |  | 56671080 | [13]C5H11NO2S | 6896.12 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 146.251% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A6 | 4713788 |  | 64908123 | [13]C5H11NO2S | 11721.96 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -421.84% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A7 | 5217794 |  | 65600007 | [13]C5H11NO2S | 15806.88 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -520.028% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A8 | 4739525 |  | 64217910 | [13]C5H11NO2S | 10562.50 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -1268.886% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A9 | 5600318 |  | 69719359 | [13]C5H11NO2S | 11734.31 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -4458.739% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A10 | 5061530 |  | 67774556 | [13]C5H11NO2S | 8915.40 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -9382.649% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A11 | 5024062 |  | 64564174 | [13]C5H11NO2S | 6237.63 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -8813.758% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_A12 | 4892493 |  | 64650764 | [13]C5H11NO2S | 10404.16 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -418.677% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B1 | 1906099 |  | 25412985 | [13]C5H11NO2S | INF | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -23.085% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B2 | 3180345 |  | 36103232 | [13]C5H11NO2S | 19810.89 | Rsquared 0.865 < R2Threshold 0.99
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B3 | 3863144 |  | 47835472 | [13]C5H11NO2S | 397420.70 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 59.778% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B4 | 4690241 |  | 55214973 | [13]C5H11NO2S | 17416.29 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 41.367% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B5 | 5385032 |  | 61134480 | [13]C5H11NO2S | 7935.34 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 48.198% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B6 | 4928161 |  | 59396993 | [13]C5H11NO2S | 7035.37 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 56.209% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B7 | 4965222 |  | 62500697 | [13]C5H11NO2S | 10427.04 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -451.483% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B8 | 5017432 |  | 62416124 | [13]C5H11NO2S | 9037.03 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -845.978% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B9 | 4719897 |  | 61328073 | [13]C5H11NO2S | 9108.63 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 208.417% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B10 | 4810955 |  | 65793079 | [13]C5H11NO2S | 8955.26 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -4446.542% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B11 | 4490745 |  | 57852061 | [13]C5H11NO2S | 9131.66 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 4626.258% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_B12 | 4319991 |  | 56313659 | [13]C5H11NO2S | 9747.69 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 11595.297% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H11NO2S | N/A
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C1 | 1726287 |  | 22868911 | [13]C5H11NO2S | INF | Rsquared 0.865 < R2Threshold 0.99
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C2 | 2674208 |  | 32713741 | [13]C5H11NO2S | INF | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 34.914% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C3 | 2826616 |  | 39206526 | [13]C5H11NO2S | 25674.89 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 98.924% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C4 | 4235988 |  | 49928282 | [13]C5H11NO2S | 16527.04 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 69.676% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C5 | 4506140 |  | 57106171 | [13]C5H11NO2S | 8838.06 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -82.674% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C6 | 4556485 |  | 56482332 | [13]C5H11NO2S | 6981.95 | Rsquared 0.865 < R2Threshold 0.99
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C7 | 4616077 |  | 65460678 | [13]C5H11NO2S | 4759.52 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -470.585% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C8 | 4400680 |  | 59688017 | [13]C5H11NO2S | 6587.49 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -671.405% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C9 | 5155161 |  | 63891358 | [13]C5H11NO2S | 6966.99 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -519.484% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C10 | 4763922 |  | 61743689 | [13]C5H11NO2S | 6534.70 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 1888.944% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C11 | 4719952 |  | 61562157 | [13]C5H11NO2S | 8320.45 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff 787.907% (absolute) > max calibration amount diff 20%
Methionine U-13C pos | 7.14 | Target Compound | KD100520_100520_C12 | 4351019 |  | 58522715 | [13]C5H11NO2S | 11091.58 | Rsquared 0.865 < R2Threshold 0.99. Calibration Amount Diff -15703.142% (absolute) > max calibration amount diff 20%

## Methionine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A1 | 20344675 |  | 250679428 | [13]C5H11[15]NO2S | 194759.30
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A2 | 19777703 |  | 275384409 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A3 | 20810234 |  | 288531823 | [13]C5H11[15]NO2S | 206500.83
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A4 | 22435651 |  | 280032928 | [13]C5H11[15]NO2S | 159254.29
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A5 | 20982297 |  | 293815451 | [13]C5H11[15]NO2S | 30770.79
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A6 | 24151324 |  | 313286630 | [13]C5H11[15]NO2S | 118318.43
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A7 | 25129025 |  | 304981982 | [13]C5H11[15]NO2S | 38553.53
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A8 | 24004951 |  | 306788557 | [13]C5H11[15]NO2S | 92982.53
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A9 | 22716213 |  | 305303306 | [13]C5H11[15]NO2S | 114904.48
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A10 | 20245663 |  | 283772878 | [13]C5H11[15]NO2S | 217023.05
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A11 | 21527375 |  | 293992873 | [13]C5H11[15]NO2S | 25339.72
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_A12 | 23946588 |  | 307782631 | [13]C5H11[15]NO2S | 45101.60
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B1 | 17958629 |  | 245239860 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B2 | 20505643 |  | 265903827 | [13]C5H11[15]NO2S | 201740.56
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B3 | 23605257 |  | 294150801 | [13]C5H11[15]NO2S | 1249681.54
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B4 | 19716955 |  | 270974395 | [13]C5H11[15]NO2S | 212042.09
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B5 | 23408464 |  | 297609382 | [13]C5H11[15]NO2S | 39906.11
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B6 | 20117513 |  | 273248666 | [13]C5H11[15]NO2S | 1084522.79
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B7 | 22473194 |  | 283673501 | [13]C5H11[15]NO2S | 164675.44
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B8 | 20893053 |  | 275480427 | [13]C5H11[15]NO2S | 33703.10
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B9 | 21479803 |  | 270875734 | [13]C5H11[15]NO2S | 17061.30
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B10 | 20231476 |  | 260205324 | [13]C5H11[15]NO2S | 80804.75
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B11 | 19120820 |  | 256647625 | [13]C5H11[15]NO2S | 74103.66
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_B12 | 18800541 |  | 264761055 | [13]C5H11[15]NO2S | 23332.76
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C5H11[15]NO2S | N/A
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C1 | 19859386 |  | 254058504 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C2 | 18350050 |  | 239482619 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C3 | 19299672 |  | 255723682 | [13]C5H11[15]NO2S | 128734.51
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C4 | 18822381 |  | 247900852 | [13]C5H11[15]NO2S | 172803.07
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C5 | 17891326 |  | 239005108 | [13]C5H11[15]NO2S | 75688.93
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C6 | 20770416 |  | 268080593 | [13]C5H11[15]NO2S | 33228.27
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C7 | 19064840 |  | 261810921 | [13]C5H11[15]NO2S | 45253.15
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C8 | 19038190 |  | 258499490 | [13]C5H11[15]NO2S | 26729.10
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C9 | 21159638 |  | 272201987 | [13]C5H11[15]NO2S | 43147.24
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C10 | 17011412 |  | 247738684 | [13]C5H11[15]NO2S | 19072.57
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C11 | 19458402 |  | 260602611 | [13]C5H11[15]NO2S | 61148.64
Methionine U-13C, U-15N neg | 7.14 | Internal Standard | KD100520_100520_C12 | 18589979 |  | 244486884 | [13]C5H11[15]NO2S | 15053.44

## Methionine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A1 | 198144890 |  | 2659251861 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A2 | 251250843 |  | 3428098453 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A3 | 304719062 |  | 4098132475 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A4 | 315370954 |  | 4389513855 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A5 | 332172184 |  | 4410840923 | [13]C5H11[15]NO2S | 11863463.44
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A6 | 317510069 |  | 4614109039 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A7 | 341930501 |  | 4719453636 | [13]C5H11[15]NO2S | 1283279.20
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A8 | 320108603 |  | 4579488657 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A9 | 350009188 |  | 4841243638 | [13]C5H11[15]NO2S | 19463628.81
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A10 | 322154046 |  | 4688836444 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A11 | 348664675 |  | 4614060594 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_A12 | 344762587 |  | 4756206239 | [13]C5H11[15]NO2S | 1349459.31
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B1 | 184768487 |  | 2566806311 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B2 | 254483676 |  | 3338641087 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B3 | 288847868 |  | 4098379063 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B4 | 338617906 |  | 4334361524 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B5 | 357941144 |  | 4651934929 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B6 | 334693323 |  | 4448676536 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B7 | 337449057 |  | 4513141133 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B8 | 334806925 |  | 4501812349 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B9 | 333386027 |  | 4533497218 | [13]C5H11[15]NO2S | 17660252.16
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B10 | 334846524 |  | 4702068848 | [13]C5H11[15]NO2S | 923519.76
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B11 | 318977757 |  | 4328791264 | [13]C5H11[15]NO2S | 1832740.84
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_B12 | 295547467 |  | 4230277231 | [13]C5H11[15]NO2S | 1035773.06
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C5H11[15]NO2S | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C1 | 165231539 |  | 2528345304 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C2 | 232721744 |  | 3159611173 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C3 | 237590110 |  | 3499663196 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C4 | 293147687 |  | 3972220204 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C5 | 309640457 |  | 4219584706 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C6 | 316670048 |  | 4199131128 | [13]C5H11[15]NO2S | 15436559.07
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C7 | 322205860 |  | 4722003976 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C8 | 296349897 |  | 4325428572 | [13]C5H11[15]NO2S | 676423.74
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C9 | 346296601 |  | 4675898556 | [13]C5H11[15]NO2S | 17182061.18
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C10 | 335616894 |  | 4607983545 | [13]C5H11[15]NO2S | 18790929.37
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C11 | 315261741 |  | 4545358987 | [13]C5H11[15]NO2S | INF
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD100520_100520_C12 | 282047860 |  | 4195096701 | [13]C5H11[15]NO2S | 1258318.46

## Phenylalanine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A1 | 259153208 |  | 3438467345 | C9H11NO2 | 595091.41
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A2 | 122404068 |  | 1614029540 | C9H11NO2 | 174823.73
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A3 | 73004561 |  | 1025792242 | C9H11NO2 | 77022.11
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A4 | 42170254 |  | 536065708 | C9H11NO2 | 127562.25
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A5 | 30080881 |  | 383788880 | C9H11NO2 | 17892.11 | Calibration Amount Diff 42.149% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A6 | 12709041 |  | 163362489 | C9H11NO2 | 6259.81
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A7 | 6014055 |  | 77157798 | C9H11NO2 | 6823.01 | Calibration Amount Diff -26.233% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A8 | 2703130 |  | 37338877 | C9H11NO2 | 3996.86 | Calibration Amount Diff -84.861% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A9 | 1763387 |  | 21722987 | C9H11NO2 | 4180.08 | Calibration Amount Diff -192.565% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A10 | 738045 |  | 9907788 | C9H11NO2 | 1634.60 | Calibration Amount Diff -448.982% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A11 | 410272 |  | 4862079 | C9H11NO2 | 889.16 | Calibration Amount Diff -925.407% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_A12 | 161978 |  | 1876107 | C9H11NO2 | 568.49 | Calibration Amount Diff -1934.033% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B1 | 294034941 |  | 3777556292 | C9H11NO2 | 1063057.51
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B2 | 127313501 |  | 1793779558 | C9H11NO2 | 191681.93
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B3 | 76345678 |  | 976083294 | C9H11NO2 | 334464.53
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B4 | 36752875 |  | 485923956 | C9H11NO2 | 81224.14
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B5 | 22071850 |  | 282812665 | C9H11NO2 | 27851.68
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B6 | 9968187 |  | 133756943 | C9H11NO2 | 6962.98
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B7 | 5868035 |  | 70267691 | C9H11NO2 | 7924.46 | Calibration Amount Diff -27.453% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B8 | 2728947 |  | 36399090 | C9H11NO2 | 3631.28 | Calibration Amount Diff -79.222% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B9 | 1526154 |  | 18956921 | C9H11NO2 | 2071.85 | Calibration Amount Diff -196.561% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B10 | 687737 |  | 8417109 | C9H11NO2 | 1354.94 | Calibration Amount Diff -456.667% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B11 | 367589 |  | 4257257 | C9H11NO2 | 675.69 | Calibration Amount Diff -935.566% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_B12 | 138108 |  | 1619140 | C9H11NO2 | 357.64 | Calibration Amount Diff -1940.739% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C9H11NO2 | N/A
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C1 | 231009799 |  | 3230603933 | C9H11NO2 | INF
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C2 | 143760685 |  | 1762948584 | C9H11NO2 | 198539.12
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C3 | 69255452 |  | 893917406 | C9H11NO2 | 114694.40
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C4 | 35175472 |  | 477389125 | C9H11NO2 | 59396.85
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C5 | 20059114 |  | 255900944 | C9H11NO2 | 24040.26
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C6 | 9436960 |  | 114822991 | C9H11NO2 | 19243.72
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C7 | 5903403 |  | 71289340 | C9H11NO2 | 5846.58 | Calibration Amount Diff -24.982% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C8 | 2310491 |  | 33099012 | C9H11NO2 | 3543.93 | Calibration Amount Diff -90.012% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C9 | 1397001 |  | 17886366 | C9H11NO2 | 3362.23 | Calibration Amount Diff -207.605% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C10 | 595047 |  | 7035242 | C9H11NO2 | 1284.02 | Calibration Amount Diff -472.24% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C11 | 307109 |  | 3577169 | C9H11NO2 | 906.31 | Calibration Amount Diff -958.841% (absolute) > max calibration amount diff 20%
Phenylalanine neg | 6.19 | Target Compound | KD100520_100520_C12 | 130519 |  | 1549509 | C9H11NO2 | 349.67 | Calibration Amount Diff -1940.574% (absolute) > max calibration amount diff 20%

## Phenylalanine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A1 | 1779707298 |  | 26982212228 | C9H11NO2 | 787472.06
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A2 | 1056662170 |  | 15713525685 | C9H11NO2 | 391378.34
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A3 | 806586261 |  | 11501584732 | C9H11NO2 | 232473.03
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A4 | 440690676 |  | 5913958578 | C9H11NO2 | 161659.17
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A5 | 294935254 |  | 4157636323 | C9H11NO2 | 133490.54 | Calibration Amount Diff 48.871% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A6 | 159867479 |  | 1992490416 | C9H11NO2 | 82957.70 | Calibration Amount Diff 24.867% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A7 | 69671475 |  | 914779569 | C9H11NO2 | 22924.20 | Calibration Amount Diff -24.288% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A8 | 35919754 |  | 448071347 | C9H11NO2 | 18658.78 | Calibration Amount Diff -98.348% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A9 | 17932844 |  | 236777307 | C9H11NO2 | 5647.58 | Calibration Amount Diff -235.545% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A10 | 8968300 |  | 115568684 | C9H11NO2 | 4792.22 | Calibration Amount Diff -546.548% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A11 | 5639038 |  | 69902418 | C9H11NO2 | 3216.67 | Calibration Amount Diff -1123.835% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_A12 | 2352639 |  | 28320251 | C9H11NO2 | 1397.22 | Calibration Amount Diff -2368.532% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B1 | 1942505327 |  | 28298682984 | C9H11NO2 | 1069621.71
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B2 | 1188969974 |  | 17281974812 | C9H11NO2 | 583571.83
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B3 | 813715860 |  | 10966407765 | C9H11NO2 | 173885.02
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B4 | 473784183 |  | 6034530286 | C9H11NO2 | 429535.85
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B5 | 249028003 |  | 3314674799 | C9H11NO2 | 86455.98
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B6 | 137665171 |  | 1767804489 | C9H11NO2 | 68070.39
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B7 | 67665713 |  | 884766431 | C9H11NO2 | 23961.27 | Calibration Amount Diff -27.378% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B8 | 33760941 |  | 438271001 | C9H11NO2 | 17178.75 | Calibration Amount Diff -89.649% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B9 | 17355414 |  | 225328212 | C9H11NO2 | 7591.94 | Calibration Amount Diff -233.021% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B10 | 8244529 |  | 113210238 | C9H11NO2 | 5706.63 | Calibration Amount Diff -546.883% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B11 | 3962446 |  | 54760351 | C9H11NO2 | 1858.21 | Calibration Amount Diff -1145.554% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_B12 | 2241325 |  | 26256717 | C9H11NO2 | 1419.34 | Calibration Amount Diff -2370.656% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C9H11NO2 | N/A
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C1 | 2034347844 |  | 27205830335 | C9H11NO2 | INF
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C2 | 1274695174 |  | 17506952647 | C9H11NO2 | 434573.92
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C3 | 727935830 |  | 9883093619 | C9H11NO2 | 549487.13
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C4 | 393797536 |  | 5628235356 | C9H11NO2 | 181024.18
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C5 | 234990143 |  | 3188502999 | C9H11NO2 | 96911.49
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C6 | 116370278 |  | 1497806822 | C9H11NO2 | 72401.59
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C7 | 67277872 |  | 897775868 | C9H11NO2 | 29477.24 | Calibration Amount Diff -21.905% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C8 | 32344418 |  | 439158288 | C9H11NO2 | 12395.81 | Calibration Amount Diff -100.552% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C9 | 17006060 |  | 238088205 | C9H11NO2 | 7845.48 | Calibration Amount Diff -242.76% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C10 | 6778695 |  | 95619964 | C9H11NO2 | 4745.65 | Calibration Amount Diff -568.33% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C11 | 4322533 |  | 53402058 | C9H11NO2 | 3000.35 | Calibration Amount Diff -1162.441% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD100520_100520_C12 | 1997365 |  | 26009982 | C9H11NO2 | 1359.25 | Calibration Amount Diff -2373.128% (absolute) > max calibration amount diff 20%

## Phenylalanine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A1 | 189014 |  | 2228158 | [13]C9H11NO2 | 23226.45 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -22.034% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A2 | 180398 |  | 2333088 | [13]C9H11NO2 | 3073.96 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 31.322% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A3 | 187508 |  | 2535329 | [13]C9H11NO2 | 1067.40 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 127.126% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A4 | 271241 |  | 3134585 | [13]C9H11NO2 | 133560.92 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 110.407% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A5 | 305541 |  | 3519703 | [13]C9H11NO2 | 4781.15 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 151.925% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A6 | 313979 |  | 3373407 | [13]C9H11NO2 | 10637.40 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 53.027% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A7 | 304739 |  | 3607581 | [13]C9H11NO2 | 3570.65 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -252.385% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A8 | 276646 |  | 3703102 | [13]C9H11NO2 | 4237.66 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -2583.426% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A9 | 346908 |  | 3970234 | [13]C9H11NO2 | 6212.34 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -5964.501% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A10 | 295357 |  | 4044413 | [13]C9H11NO2 | 2233.40 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -13879.756% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A11 | 279470 |  | 3437521 | [13]C9H11NO2 | 1130.11 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -10593.427% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_A12 | 291123 |  | 3484508 | [13]C9H11NO2 | 2666.01 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -19165.623% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B1 | 199197 |  | 2243352 | [13]C9H11NO2 | 2493.55 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -25.31% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B2 | 188359 |  | 2398871 | [13]C9H11NO2 | INF | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 34.035% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B3 | 229154 |  | 2402883 | [13]C9H11NO2 | 109433.82 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 137.591% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B4 | 187316 |  | 2227196 | [13]C9H11NO2 | 5709.21 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 367.139% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B5 | 318848 |  | 3677567 | [13]C9H11NO2 | 4776.59 | Rsquared 0.727 < R2Threshold 0.99
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B6 | 251126 |  | 2993604 | [13]C9H11NO2 | 12041.14 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 310.45% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B7 | 309573 |  | 3474257 | [13]C9H11NO2 | 5769.42 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -925.883% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B8 | 248215 |  | 3266896 | [13]C9H11NO2 | 2101.12 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -834.989% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B9 | 283352 |  | 3396017 | [13]C9H11NO2 | 1997.04 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -3144.472% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B10 | 277287 |  | 3339108 | [13]C9H11NO2 | 9990.22 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -3540.496% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B11 | 283779 |  | 3352434 | [13]C9H11NO2 | 2386.08 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -19248.499% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_B12 | 247163 |  | 3120578 | [13]C9H11NO2 | 3595.08 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -7774.35% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C9H11NO2 | N/A
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C1 | 142234 |  | 1835603 | [13]C9H11NO2 | INF | Rsquared 0.727 < R2Threshold 0.99
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C2 | 207614 |  | 2372194 | [13]C9H11NO2 | 7139.45 | Rsquared 0.727 < R2Threshold 0.99
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C3 | 199518 |  | 2298187 | [13]C9H11NO2 | 237643.55 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 109.814% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C4 | 209451 |  | 2497925 | [13]C9H11NO2 | 2994.28 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 240.569% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C5 | 270800 |  | 2989890 | [13]C9H11NO2 | 6492.57 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 22.157% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C6 | 264586 |  | 2744055 | [13]C9H11NO2 | 2902.38 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 364.263% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C7 | 313972 |  | 3247670 | [13]C9H11NO2 | 7611.46 | Rsquared 0.727 < R2Threshold 0.99
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C8 | 233723 |  | 3247547 | [13]C9H11NO2 | 4479.40 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -1129.775% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C9 | 273605 |  | 3407139 | [13]C9H11NO2 | 1587.53 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -2157.912% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C10 | 270255 |  | 3231950 | [13]C9H11NO2 | 5120.31 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -5669.569% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C11 | 272442 |  | 3286345 | [13]C9H11NO2 | 2161.96 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -11749.82% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C neg | 6.19 | Target Compound | KD100520_100520_C12 | 243719 |  | 2953045 | [13]C9H11NO2 | 2572.06 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -3419.643% (absolute) > max calibration amount diff 20%

## Phenylalanine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A1 | 1788063 |  | 21936041 | [13]C9H11NO2 | INF | Rsquared 0.799 < R2Threshold 0.99
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A2 | 2131054 |  | 28692435 | [13]C9H11NO2 | 81672.80 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 72.675% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A3 | 3118111 |  | 39674845 | [13]C9H11NO2 | 16081.62 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 78.087% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A4 | 3376756 |  | 42778085 | [13]C9H11NO2 | 6977.71 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 83.872% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A5 | 3375461 |  | 44255359 | [13]C9H11NO2 | 2859.77 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 246.077% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A6 | 4077549 |  | 49502116 | [13]C9H11NO2 | 2884.07 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -568.48% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A7 | 3675428 |  | 47269641 | [13]C9H11NO2 | 4719.51 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -403.462% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A8 | 3604911 |  | 46107708 | [13]C9H11NO2 | 2642.91 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -991.206% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A9 | 3623944 |  | 45485121 | [13]C9H11NO2 | 1891.85 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -3563.495% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A10 | 3842422 |  | 47557309 | [13]C9H11NO2 | 3951.24 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -10349.6% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A11 | 3698126 |  | 45596938 | [13]C9H11NO2 | 3373.95 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -1910.925% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_A12 | 3683269 |  | 45086351 | [13]C9H11NO2 | 4259.55 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -18637.791% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B1 | 1697411 |  | 24245834 | [13]C9H11NO2 | INF | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -30.003% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B2 | 2112669 |  | 31950014 | [13]C9H11NO2 | 3849.60 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 37.459% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B3 | 3026428 |  | 38541698 | [13]C9H11NO2 | 3417.22 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 88.45% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B4 | 3376827 |  | 38890062 | [13]C9H11NO2 | 2558.17 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 223.868% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B5 | 3844625 |  | 48929518 | [13]C9H11NO2 | 3982.16 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 74.336% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B6 | 3689839 |  | 47006646 | [13]C9H11NO2 | 4797.34 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -230.42% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B7 | 3877685 |  | 48464969 | [13]C9H11NO2 | 2729.05 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -1144.901% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B8 | 3391539 |  | 43186233 | [13]C9H11NO2 | 3700.90 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -1453.841% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B9 | 3185621 |  | 40878722 | [13]C9H11NO2 | 2481.77 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -276.805% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B10 | 3292072 |  | 45228016 | [13]C9H11NO2 | 2593.73 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -5295.24% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B11 | 2936905 |  | 40057194 | [13]C9H11NO2 | 3279.65 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 93.428% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_B12 | 3805185 |  | 42265583 | [13]C9H11NO2 | 2538.92 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -17177.191% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C9H11NO2 | N/A
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C1 | 1922340 |  | 24978745 | [13]C9H11NO2 | INF | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -20.746% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C2 | 2558046 |  | 31750535 | [13]C9H11NO2 | 19181.11 | Rsquared 0.799 < R2Threshold 0.99
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C3 | 2896967 |  | 34427607 | [13]C9H11NO2 | 1676.31 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 96.254% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C4 | 3022280 |  | 40223033 | [13]C9H11NO2 | 3171.89 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 66.683% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C5 | 3588897 |  | 45005509 | [13]C9H11NO2 | 7998.85 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -113.257% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C6 | 3458337 |  | 42908622 | [13]C9H11NO2 | 3935.48 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -113.24% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C7 | 3520721 |  | 45083045 | [13]C9H11NO2 | 4866.83 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -154.548% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C8 | 3477670 |  | 44660160 | [13]C9H11NO2 | 3313.40 | Rsquared 0.799 < R2Threshold 0.99
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C9 | 3284685 |  | 45502136 | [13]C9H11NO2 | 3051.38 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 238.937% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C10 | 3129003 |  | 42487030 | [13]C9H11NO2 | 2191.98 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 555.723% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C11 | 3501729 |  | 43568506 | [13]C9H11NO2 | 6733.93 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff -1000.499% (absolute) > max calibration amount diff 20%
Phenylalanine U-13C pos | 6.19 | Target Compound | KD100520_100520_C12 | 3116400 |  | 40611220 | [13]C9H11NO2 | 2612.42 | Rsquared 0.799 < R2Threshold 0.99. Calibration Amount Diff 11082.612% (absolute) > max calibration amount diff 20%

## Phenylalanine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A1 | 35137710 |  | 448255461 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A2 | 31777634 |  | 431304956 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A3 | 32122702 |  | 442780896 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A4 | 36645534 |  | 460119652 | [13]C9H11[15]NO2 | 1491782.81
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A5 | 38851003 |  | 489606544 | [13]C9H11[15]NO2 | 144616.15
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A6 | 35745216 |  | 445059739 | [13]C9H11[15]NO2 | 149347.02
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A7 | 34789174 |  | 460441982 | [13]C9H11[15]NO2 | 105400.85
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A8 | 30294634 |  | 438664323 | [13]C9H11[15]NO2 | 91320.03
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A9 | 34560084 |  | 463446356 | [13]C9H11[15]NO2 | 23669.81
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A10 | 32414477 |  | 464339612 | [13]C9H11[15]NO2 | 16699.40
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A11 | 30155854 |  | 423587415 | [13]C9H11[15]NO2 | 43962.53
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_A12 | 31382893 |  | 431153176 | [13]C9H11[15]NO2 | 19666.05
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B1 | 34619456 |  | 440952175 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B2 | 33878002 |  | 447465887 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B3 | 33026417 |  | 426598745 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B4 | 30030450 |  | 392907286 | [13]C9H11[15]NO2 | 547144.89
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B5 | 38177945 |  | 486793493 | [13]C9H11[15]NO2 | 31657.97
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B6 | 30803550 |  | 410482416 | [13]C9H11[15]NO2 | 63405.50
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B7 | 33418271 |  | 423162705 | [13]C9H11[15]NO2 | 54623.45
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B8 | 30393799 |  | 410680226 | [13]C9H11[15]NO2 | 31855.12
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B9 | 31352266 |  | 415461537 | [13]C9H11[15]NO2 | 25941.63
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B10 | 29881918 |  | 417920752 | [13]C9H11[15]NO2 | 44438.51
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B11 | 30113510 |  | 398321667 | [13]C9H11[15]NO2 | 22976.05
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_B12 | 29089580 |  | 395828576 | [13]C9H11[15]NO2 | 20010.16
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C9H11[15]NO2 | N/A
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C1 | 30431652 |  | 419425100 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C2 | 32706143 |  | 403656809 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C3 | 30985976 |  | 390833527 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C4 | 31709412 |  | 400789169 | [13]C9H11[15]NO2 | 170168.35
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C5 | 31247930 |  | 399826154 | [13]C9H11[15]NO2 | 49489.00
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C6 | 30040386 |  | 379382986 | [13]C9H11[15]NO2 | 57563.45
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C7 | 34693722 |  | 421501708 | [13]C9H11[15]NO2 | 11786.33
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C8 | 28415093 |  | 404077783 | [13]C9H11[15]NO2 | 8720.61
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C9 | 30068240 |  | 423947144 | [13]C9H11[15]NO2 | 12961.32
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C10 | 29678320 |  | 397137115 | [13]C9H11[15]NO2 | 14116.90
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C11 | 30252077 |  | 402961004 | [13]C9H11[15]NO2 | 15835.98
Phenylalanine U-13C, U-15N neg | 6.19 | Internal Standard | KD100520_100520_C12 | 27279671 |  | 378212285 | [13]C9H11[15]NO2 | 10117.82

## Phenylalanine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A1 | 242520235 |  | 3797038088 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A2 | 321003735 |  | 4626564134 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A3 | 377485447 |  | 5335854068 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A4 | 389330405 |  | 5300199250 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A5 | 384502115 |  | 5456220444 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A6 | 431529353 |  | 5391683045 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A7 | 408163068 |  | 5317685525 | [13]C9H11[15]NO2 | 1554771.74
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A8 | 401952091 |  | 5149024626 | [13]C9H11[15]NO2 | 10463176.42
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A9 | 384091780 |  | 4971921104 | [13]C9H11[15]NO2 | 867096.34
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A10 | 394036197 |  | 5091694370 | [13]C9H11[15]NO2 | 999104.40
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A11 | 404390523 |  | 5180662769 | [13]C9H11[15]NO2 | 979132.78
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_A12 | 381708570 |  | 5000620699 | [13]C9H11[15]NO2 | 917164.71
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B1 | 258888766 |  | 3627525161 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B2 | 318226416 |  | 4753601407 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B3 | 382324669 |  | 5237303342 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B4 | 402493940 |  | 5147812505 | [13]C9H11[15]NO2 | 596852.69
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B5 | 436981155 |  | 5805789012 | [13]C9H11[15]NO2 | 2007882.57
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B6 | 408238864 |  | 5299820642 | [13]C9H11[15]NO2 | 14227271.04
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B7 | 379819646 |  | 5250242802 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B8 | 369829161 |  | 4766013980 | [13]C9H11[15]NO2 | 7704905.44
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B9 | 345727787 |  | 4661414103 | [13]C9H11[15]NO2 | 1092668.23
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B10 | 367777252 |  | 4998256913 | [13]C9H11[15]NO2 | 1206339.29
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B11 | 323210204 |  | 4581290335 | [13]C9H11[15]NO2 | 1046803.54
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_B12 | 385597380 |  | 4698724006 | [13]C9H11[15]NO2 | 1675999.51
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C9H11[15]NO2 | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C1 | 302747790 |  | 3896414140 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C2 | 312768236 |  | 4400105000 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C3 | 369046984 |  | 4715136798 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C4 | 360214395 |  | 4944787853 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C5 | 379516680 |  | 5129611278 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C6 | 379095924 |  | 4897445840 | [13]C9H11[15]NO2 | 11762091.91
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C7 | 375464498 |  | 5138037800 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C8 | 389741030 |  | 5120179861 | [13]C9H11[15]NO2 | INF
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C9 | 378102255 |  | 5224032491 | [13]C9H11[15]NO2 | 9937651.11
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C10 | 346435212 |  | 4877168737 | [13]C9H11[15]NO2 | 1149624.30
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C11 | 380104703 |  | 4964985621 | [13]C9H11[15]NO2 | 1430671.84
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD100520_100520_C12 | 359540216 |  | 4728724630 | [13]C9H11[15]NO2 | 823093.61

## Proline neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Proline neg | 8.15 | Target Compound | KD100520_100520_A1 | 74912052 |  | 1040789363 | C5H9NO2 | 42513.74
Proline neg | 8.15 | Target Compound | KD100520_100520_A2 | 36912345 |  | 505972617 | C5H9NO2 | 39585.46
Proline neg | 8.15 | Target Compound | KD100520_100520_A3 | 22935101 |  | 305605608 | C5H9NO2 | 19828.89
Proline neg | 8.15 | Target Compound | KD100520_100520_A4 | 11181934 |  | 150721031 | C5H9NO2 | 8271.63
Proline neg | 8.15 | Target Compound | KD100520_100520_A5 | 7804995 |  | 101893811 | C5H9NO2 | 9610.33 | Calibration Amount Diff 36.438% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_A6 | 3621371 |  | 46826963 | C5H9NO2 | 4328.17
Proline neg | 8.15 | Target Compound | KD100520_100520_A7 | 1700112 |  | 20776185 | C5H9NO2 | 3129.98
Proline neg | 8.15 | Target Compound | KD100520_100520_A8 | 801344 |  | 9651372 | C5H9NO2 | 1775.16 | Calibration Amount Diff -62.758% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_A9 | 395804 |  | 5088945 | C5H9NO2 | 689.51 | Calibration Amount Diff -122.638% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_A10 | 175538 |  | 1926510 | C5H9NO2 | 343.89 | Calibration Amount Diff -285.876% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_A11 | 76669 |  | 1052899 | C5H9NO2 | 219.42 | Calibration Amount Diff -546.289% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_A12 | 34571 |  | 410817 | C5H9NO2 | 95.03 | Calibration Amount Diff -1105.71% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_B1 | 81817700 |  | 1109263087 | C5H9NO2 | 34176.03
Proline neg | 8.15 | Target Compound | KD100520_100520_B2 | 40015272 |  | 544630186 | C5H9NO2 | 28271.28
Proline neg | 8.15 | Target Compound | KD100520_100520_B3 | 22300274 |  | 301892368 | C5H9NO2 | 23534.12
Proline neg | 8.15 | Target Compound | KD100520_100520_B4 | 12749206 |  | 167248265 | C5H9NO2 | 8820.35
Proline neg | 8.15 | Target Compound | KD100520_100520_B5 | 5800637 |  | 74593079 | C5H9NO2 | 8558.81
Proline neg | 8.15 | Target Compound | KD100520_100520_B6 | 2906173 |  | 39253208 | C5H9NO2 | 3177.35
Proline neg | 8.15 | Target Compound | KD100520_100520_B7 | 1696948 |  | 20834237 | C5H9NO2 | 1946.04
Proline neg | 8.15 | Target Compound | KD100520_100520_B8 | 797669 |  | 9830557 | C5H9NO2 | 1178.99 | Calibration Amount Diff -55.478% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_B9 | 385969 |  | 4914085 | C5H9NO2 | 636.40 | Calibration Amount Diff -123.769% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_B10 | 119279 |  | 1624863 | C5H9NO2 | 251.95 | Calibration Amount Diff -298.781% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_B11 | 94307 |  | 1037311 | C5H9NO2 | 181.67 | Calibration Amount Diff -543.993% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_B12 | 31928 |  | 451329 | C5H9NO2 | 59.19 | Calibration Amount Diff -1097.959% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H9NO2 | N/A
Proline neg | 8.15 | Target Compound | KD100520_100520_C1 | 79644692 |  | 1048769780 | C5H9NO2 | 33094.16
Proline neg | 8.15 | Target Compound | KD100520_100520_C2 | 38905599 |  | 554593541 | C5H9NO2 | 23720.71
Proline neg | 8.15 | Target Compound | KD100520_100520_C3 | 20251713 |  | 278546133 | C5H9NO2 | 8604.04
Proline neg | 8.15 | Target Compound | KD100520_100520_C4 | 10836414 |  | 144303415 | C5H9NO2 | 5752.22
Proline neg | 8.15 | Target Compound | KD100520_100520_C5 | 5798526 |  | 74817638 | C5H9NO2 | 5492.28
Proline neg | 8.15 | Target Compound | KD100520_100520_C6 | 2632379 |  | 34819802 | C5H9NO2 | 2548.79
Proline neg | 8.15 | Target Compound | KD100520_100520_C7 | 1631024 |  | 20416898 | C5H9NO2 | 1861.15
Proline neg | 8.15 | Target Compound | KD100520_100520_C8 | 809968 |  | 9474165 | C5H9NO2 | 1294.76 | Calibration Amount Diff -62.532% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_C9 | 356911 |  | 4347996 | C5H9NO2 | 586.64 | Calibration Amount Diff -139% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_C10 | 148814 |  | 1587193 | C5H9NO2 | 275.93 | Calibration Amount Diff -297.637% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_C11 | 78972 |  | 947036 | C5H9NO2 | 86.48 | Calibration Amount Diff -558.592% (absolute) > max calibration amount diff 20%
Proline neg | 8.15 | Target Compound | KD100520_100520_C12 | 30058 |  | 438753 | C5H9NO2 | 43.34 | Calibration Amount Diff -1091.929% (absolute) > max calibration amount diff 20%

## Proline pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Proline pos | 8.15 | Target Compound | KD100520_100520_A1 | 2003675455 |  | 34309446421 | C5H9NO2 | 163713.52
Proline pos | 8.15 | Target Compound | KD100520_100520_A2 | 1424874933 |  | 24230449839 | C5H9NO2 | 199323.21
Proline pos | 8.15 | Target Compound | KD100520_100520_A3 | 1115078494 |  | 18080934257 | C5H9NO2 | 128964.15
Proline pos | 8.15 | Target Compound | KD100520_100520_A4 | 704961824 |  | 11043765872 | C5H9NO2 | 144926.20
Proline pos | 8.15 | Target Compound | KD100520_100520_A5 | 503713712 |  | 7850747508 | C5H9NO2 | 128137.08 | Calibration Amount Diff 51.155% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_A6 | 260473404 |  | 4017508026 | C5H9NO2 | 65250.43 | Calibration Amount Diff 23.835% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_A7 | 127966025 |  | 1956645261 | C5H9NO2 | 49357.29 | Calibration Amount Diff -27.589% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_A8 | 67750118 |  | 1024603165 | C5H9NO2 | 22673.77 | Calibration Amount Diff -111.619% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_A9 | 36881161 |  | 545254627 | C5H9NO2 | 8144.12 | Calibration Amount Diff -278.151% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_A10 | 17490267 |  | 257524690 | C5H9NO2 | 5952.81 | Calibration Amount Diff -658.476% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_A11 | 10134241 |  | 150917709 | C5H9NO2 | 2950.57 | Calibration Amount Diff -1361.004% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_A12 | 4670899 |  | 66678366 | C5H9NO2 | 1408.66 | Calibration Amount Diff -2865.169% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_B1 | 2149697932 |  | 36923367902 | C5H9NO2 | 169100.13
Proline pos | 8.15 | Target Compound | KD100520_100520_B2 | 1511564233 |  | 25494441099 | C5H9NO2 | 211857.84
Proline pos | 8.15 | Target Compound | KD100520_100520_B3 | 1092310502 |  | 17610618472 | C5H9NO2 | 137760.08
Proline pos | 8.15 | Target Compound | KD100520_100520_B4 | 750167217 |  | 11762294222 | C5H9NO2 | 134327.04 | Calibration Amount Diff 20.102% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_B5 | 410670991 |  | 6394552934 | C5H9NO2 | 85361.49
Proline pos | 8.15 | Target Compound | KD100520_100520_B6 | 229062313 |  | 3635038071 | C5H9NO2 | 59869.11
Proline pos | 8.15 | Target Compound | KD100520_100520_B7 | 124180484 |  | 1900540142 | C5H9NO2 | 40072.86 | Calibration Amount Diff -32.803% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_B8 | 65292577 |  | 1020868860 | C5H9NO2 | 15578.16 | Calibration Amount Diff -104.547% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_B9 | 35317928 |  | 535925867 | C5H9NO2 | 12675.08 | Calibration Amount Diff -279.528% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_B10 | 16199376 |  | 242557238 | C5H9NO2 | 4891.04 | Calibration Amount Diff -661.85% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_B11 | 8610412 |  | 134714420 | C5H9NO2 | 2581.76 | Calibration Amount Diff -1372.657% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_B12 | 4519453 |  | 64808162 | C5H9NO2 | 1416.19 | Calibration Amount Diff -2860.08% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H9NO2 | N/A
Proline pos | 8.15 | Target Compound | KD100520_100520_C1 | 1913639121 |  | 34489910884 | C5H9NO2 | 200167.50
Proline pos | 8.15 | Target Compound | KD100520_100520_C2 | 1496927803 |  | 26197403645 | C5H9NO2 | 151542.96
Proline pos | 8.15 | Target Compound | KD100520_100520_C3 | 1035201507 |  | 17277385684 | C5H9NO2 | 210192.26
Proline pos | 8.15 | Target Compound | KD100520_100520_C4 | 668974489 |  | 10413608683 | C5H9NO2 | 121489.56
Proline pos | 8.15 | Target Compound | KD100520_100520_C5 | 397672982 |  | 6214949874 | C5H9NO2 | 81887.96
Proline pos | 8.15 | Target Compound | KD100520_100520_C6 | 226436945 |  | 3255727642 | C5H9NO2 | 59027.92
Proline pos | 8.15 | Target Compound | KD100520_100520_C7 | 125586454 |  | 1889776921 | C5H9NO2 | 45791.24 | Calibration Amount Diff -27.142% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_C8 | 65016233 |  | 959816333 | C5H9NO2 | 15002.70 | Calibration Amount Diff -118.065% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_C9 | 33610685 |  | 502196974 | C5H9NO2 | 13369.89 | Calibration Amount Diff -290.038% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_C10 | 14343357 |  | 219941159 | C5H9NO2 | 3519.02 | Calibration Amount Diff -680.472% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_C11 | 8377692 |  | 127684490 | C5H9NO2 | 2967.97 | Calibration Amount Diff -1399.593% (absolute) > max calibration amount diff 20%
Proline pos | 8.15 | Target Compound | KD100520_100520_C12 | 4373085 |  | 63482136 | C5H9NO2 | 1356.71 | Calibration Amount Diff -2859.192% (absolute) > max calibration amount diff 20%

## Proline U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A1 | 1458848 |  | 17841800 | [13]C5H9NO2 | 5567.40 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 29.309% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A2 | 1527965 |  | 19658651 | [13]C5H9NO2 | 2834.64 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 27.97% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A3 | 1529268 |  | 19838262 | [13]C5H9NO2 | 12361.88 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 125.769% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A4 | 1574740 |  | 19820796 | [13]C5H9NO2 | 22933.42 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 339.835% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A5 | 1689775 |  | 21142508 | [13]C5H9NO2 | 2842.18 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 300.885% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A6 | 1895497 |  | 23610256 | [13]C5H9NO2 | 7258.76 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -1549.672% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A7 | 1813131 |  | 22680049 | [13]C5H9NO2 | 6396.81 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -951.988% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A8 | 1786339 |  | 23196462 | [13]C5H9NO2 | 7027.65 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -2322.429% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A9 | 1795783 |  | 22644891 | [13]C5H9NO2 | 3994.39 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -2508.733% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A10 | 1909237 |  | 23801448 | [13]C5H9NO2 | 15197.46 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -21581.395% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A11 | 1678572 |  | 21969422 | [13]C5H9NO2 | 8678.07 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 19725.669% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_A12 | 1783030 |  | 22006825 | [13]C5H9NO2 | 12827.55 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 34628.004% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B1 | 1433312 |  | 17926699 | [13]C5H9NO2 | INF | Rsquared 0.533 < R2Threshold 0.99
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B2 | 1592999 |  | 19848674 | [13]C5H9NO2 | 10155.80 | Rsquared 0.533 < R2Threshold 0.99
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B3 | 1596303 |  | 20682567 | [13]C5H9NO2 | 3748.26 | Rsquared 0.533 < R2Threshold 0.99
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B4 | 1809717 |  | 22744916 | [13]C5H9NO2 | 6105.59 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -280.881% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B5 | 1789733 |  | 21934735 | [13]C5H9NO2 | 3872.11 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -83.232% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B6 | 1562318 |  | 20995676 | [13]C5H9NO2 | 4450.69 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 220.179% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B7 | 1866430 |  | 22995478 | [13]C5H9NO2 | 3959.35 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -1446.248% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B8 | 1651019 |  | 21636350 | [13]C5H9NO2 | 3936.77 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 243.567% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B9 | 1674194 |  | 21874030 | [13]C5H9NO2 | 6024.66 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -86.704% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B10 | 1809427 |  | 23001702 | [13]C5H9NO2 | 14797.70 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -5431.756% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B11 | 1690483 |  | 21056807 | [13]C5H9NO2 | 3499.44 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 22683.118% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_B12 | 1737696 |  | 21884506 | [13]C5H9NO2 | 5272.10 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 30807.597% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H9NO2 | N/A
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C1 | 1555313 |  | 18795194 | [13]C5H9NO2 | 22889.33 | Rsquared 0.533 < R2Threshold 0.99
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C2 | 1584045 |  | 20517795 | [13]C5H9NO2 | 3238.61 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -155.955% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C3 | 1489887 |  | 19567064 | [13]C5H9NO2 | 8076.12 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -71.036% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C4 | 1615921 |  | 20319696 | [13]C5H9NO2 | 2520.53 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -150.833% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C5 | 1719330 |  | 20846039 | [13]C5H9NO2 | 5345.76 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -632.362% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C6 | 1525852 |  | 20451428 | [13]C5H9NO2 | 6811.80 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff -927.788% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C7 | 1689737 |  | 21417348 | [13]C5H9NO2 | 6909.17 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 1252.848% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C8 | 1809993 |  | 21681247 | [13]C5H9NO2 | 5986.27 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 3404.061% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C9 | 1811536 |  | 21849445 | [13]C5H9NO2 | 9734.67 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 6320.183% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C10 | 1728207 |  | 20972847 | [13]C5H9NO2 | 5930.79 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 20095.028% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C11 | 1789613 |  | 22421546 | [13]C5H9NO2 | 7600.11 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 34903.733% (absolute) > max calibration amount diff 20%
Proline U-13C neg | 8.15 | Target Compound | KD100520_100520_C12 | 1651361 |  | 19669293 | [13]C5H9NO2 | 7048.49 | Rsquared 0.533 < R2Threshold 0.99. Calibration Amount Diff 42380.901% (absolute) > max calibration amount diff 20%

## Proline U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A1 | 49553320 |  | 804040905 | [13]C5H9NO2 | INF | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 168.259% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A2 | 71399416 |  | 1180379595 | [13]C5H9NO2 | 71806.79 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 109.26% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A3 | 86806012 |  | 1386773982 | [13]C5H9NO2 | 30671.98 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 418.54% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A4 | 105948351 |  | 1635204707 | [13]C5H9NO2 | 125508.24 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 679.753% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A5 | 109466427 |  | 1681233871 | [13]C5H9NO2 | 69594.19 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 1818.819% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A6 | 117417273 |  | 1850852156 | [13]C5H9NO2 | 77274.57 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -8352.828% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A7 | 117398348 |  | 1839349035 | [13]C5H9NO2 | 48459.11 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -10793.596% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A8 | 120219481 |  | 1830775180 | [13]C5H9NO2 | 77579.25 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -19633.041% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A9 | 114373037 |  | 1757291083 | [13]C5H9NO2 | 65830.68 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -31614.524% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A10 | 114068179 |  | 1763170035 | [13]C5H9NO2 | 82812.88 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -60243.772% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A11 | 106207471 |  | 1637565180 | [13]C5H9NO2 | 76816.31 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 239276.351% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_A12 | 103659638 |  | 1643672812 | [13]C5H9NO2 | 53342.24 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 501265.065% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B1 | 51558573 |  | 827253083 | [13]C5H9NO2 | INF | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -120.681% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B2 | 72844525 |  | 1170286851 | [13]C5H9NO2 | INF | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -114.052% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B3 | 88985785 |  | 1401582255 | [13]C5H9NO2 | 110648.39 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -351.315% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B4 | 110683744 |  | 1710106859 | [13]C5H9NO2 | 55325.06 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -723.524% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B5 | 119320772 |  | 1869510611 | [13]C5H9NO2 | 76556.75 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -975.166% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B6 | 115269537 |  | 1855506424 | [13]C5H9NO2 | 148904.22 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -2722.467% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B7 | 116662564 |  | 1819203068 | [13]C5H9NO2 | 109623.75 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -3701.251% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B8 | 107746448 |  | 1711800217 | [13]C5H9NO2 | 60358.40 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 3870.617% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B9 | 108410701 |  | 1713671158 | [13]C5H9NO2 | 74328.40 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -1355.546% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B10 | 108966034 |  | 1665844196 | [13]C5H9NO2 | 63426.66 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 3153.472% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B11 | 94489909 |  | 1539103279 | [13]C5H9NO2 | 122352.45 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 267357.613% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_B12 | 101751886 |  | 1558455154 | [13]C5H9NO2 | 79343.84 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 477437.91% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H9NO2 | N/A
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C1 | 47977335 |  | 803292188 | [13]C5H9NO2 | INF | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 300.595% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C2 | 68590936 |  | 1149626096 | [13]C5H9NO2 | 81446.21 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -662.595% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C3 | 89055894 |  | 1458909721 | [13]C5H9NO2 | INF | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -1701.136% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C4 | 104076760 |  | 1606063904 | [13]C5H9NO2 | 120841.23 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -2565.635% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C5 | 111085621 |  | 1741504839 | [13]C5H9NO2 | 88883.30 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -5681.55% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C6 | 120736023 |  | 1794423313 | [13]C5H9NO2 | 124533.47 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff -9956.419% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C7 | 110264960 |  | 1708199525 | [13]C5H9NO2 | 52032.22 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 7381.662% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C8 | 109097964 |  | 1692787228 | [13]C5H9NO2 | 126061.56 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 30729.939% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C9 | 105208462 |  | 1639645732 | [13]C5H9NO2 | 114453.63 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 69185.868% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C10 | 100327238 |  | 1637192063 | [13]C5H9NO2 | 70175.47 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 124513.736% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C11 | 108087345 |  | 1696203599 | [13]C5H9NO2 | 113683.39 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 189180.797% (absolute) > max calibration amount diff 20%
Proline U-13C pos | 8.15 | Target Compound | KD100520_100520_C12 | 97836328 |  | 1517435054 | [13]C5H9NO2 | 59000.39 | Rsquared 0.02 < R2Threshold 0.99. Calibration Amount Diff 500166.878% (absolute) > max calibration amount diff 20%

## Proline U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A1 | 9615375 |  | 128338109 | [13]C5H9[15]NO2 | 44638.84
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A2 | 9772606 |  | 131095652 | [13]C5H9[15]NO2 | 15894.35
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A3 | 9783732 |  | 131188791 | [13]C5H9[15]NO2 | 33336.29
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A4 | 9586018 |  | 130861392 | [13]C5H9[15]NO2 | 24116.04
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A5 | 10431885 |  | 135119161 | [13]C5H9[15]NO2 | 12809.36
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A6 | 10706016 |  | 140328218 | [13]C5H9[15]NO2 | 23451.00
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A7 | 11012548 |  | 139230606 | [13]C5H9[15]NO2 | 9436.70
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A8 | 10768516 |  | 141810862 | [13]C5H9[15]NO2 | 23871.44
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A9 | 10721341 |  | 139574807 | [13]C5H9[15]NO2 | 10838.44
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A10 | 11419663 |  | 141934464 | [13]C5H9[15]NO2 | 33923.42
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A11 | 10385052 |  | 139552100 | [13]C5H9[15]NO2 | 9118.45
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_A12 | 10943725 |  | 139433488 | [13]C5H9[15]NO2 | 29147.01
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B1 | 9273838 |  | 124561637 | [13]C5H9[15]NO2 | 70850.80
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B2 | 9840206 |  | 130363503 | [13]C5H9[15]NO2 | 17678.29
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B3 | 9881375 |  | 132677011 | [13]C5H9[15]NO2 | 57406.01
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B4 | 10497032 |  | 138307483 | [13]C5H9[15]NO2 | 12311.39
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B5 | 10787834 |  | 136672124 | [13]C5H9[15]NO2 | 59477.88
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B6 | 9674898 |  | 132054728 | [13]C5H9[15]NO2 | 18785.76
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B7 | 11140677 |  | 140055460 | [13]C5H9[15]NO2 | 19969.69
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B8 | 10170150 |  | 135043357 | [13]C5H9[15]NO2 | 13068.12
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B9 | 10404461 |  | 136152435 | [13]C5H9[15]NO2 | 21300.00
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B10 | 10777952 |  | 141626511 | [13]C5H9[15]NO2 | 26964.86
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B11 | 10523109 |  | 134166848 | [13]C5H9[15]NO2 | 12388.60
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_B12 | 10990108 |  | 138384858 | [13]C5H9[15]NO2 | 12724.01
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C5H9[15]NO2 | N/A
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C1 | 10155432 |  | 131124757 | [13]C5H9[15]NO2 | INF
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C2 | 8906146 |  | 124086840 | [13]C5H9[15]NO2 | 28589.97
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C3 | 8960733 |  | 122712123 | [13]C5H9[15]NO2 | 39061.88
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C4 | 9473588 |  | 125639075 | [13]C5H9[15]NO2 | 11525.65
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C5 | 9877929 |  | 125399820 | [13]C5H9[15]NO2 | 19292.19
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C6 | 9118451 |  | 123950155 | [13]C5H9[15]NO2 | 19116.82
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C7 | 10398745 |  | 136304658 | [13]C5H9[15]NO2 | 17337.47
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C8 | 11174898 |  | 138906684 | [13]C5H9[15]NO2 | 18045.63
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C9 | 11132454 |  | 139640382 | [13]C5H9[15]NO2 | 18408.26
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C10 | 10726473 |  | 136132802 | [13]C5H9[15]NO2 | 17111.04
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C11 | 11196644 |  | 144703294 | [13]C5H9[15]NO2 | 27809.26
Proline U-13C, U-15N neg | 8.15 | Internal Standard | KD100520_100520_C12 | 10070336 |  | 125125149 | [13]C5H9[15]NO2 | 12815.87

## Proline U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A1 | 309077346 |  | 5051569630 | [13]C5H9[15]NO2 | INF
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A2 | 440790862 |  | 7257464718 | [13]C5H9[15]NO2 | INF
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A3 | 541104832 |  | 8554409319 | [13]C5H9[15]NO2 | 4092562.25
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A4 | 648815178 |  | 10044521555 | [13]C5H9[15]NO2 | 6224610.01
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A5 | 666935874 |  | 10357623987 | [13]C5H9[15]NO2 | 1077296.92
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A6 | 690951966 |  | 10865190643 | [13]C5H9[15]NO2 | 929653.95
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A7 | 707934121 |  | 10921366291 | [13]C5H9[15]NO2 | 1255354.35
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A8 | 713522777 |  | 10890389647 | [13]C5H9[15]NO2 | 2909501.72
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A9 | 684103403 |  | 10492387852 | [13]C5H9[15]NO2 | 10150320.36
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A10 | 677709780 |  | 10534941037 | [13]C5H9[15]NO2 | 9586710.99
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A11 | 667365321 |  | 10241031738 | [13]C5H9[15]NO2 | 1142241.26
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_A12 | 653509030 |  | 10294323270 | [13]C5H9[15]NO2 | 776781.57
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B1 | 308398619 |  | 5004366351 | [13]C5H9[15]NO2 | INF
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B2 | 435368939 |  | 7091955421 | [13]C5H9[15]NO2 | INF
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B3 | 536947813 |  | 8433011892 | [13]C5H9[15]NO2 | 643494.42
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B4 | 667267781 |  | 10269495253 | [13]C5H9[15]NO2 | 1063319.15
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B5 | 718729324 |  | 11260139181 | [13]C5H9[15]NO2 | 768236.28
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B6 | 700198581 |  | 11136944607 | [13]C5H9[15]NO2 | INF
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B7 | 707696878 |  | 10954926180 | [13]C5H9[15]NO2 | 981937.00
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B8 | 657667074 |  | 10424604174 | [13]C5H9[15]NO2 | 630007.10
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B9 | 662578713 |  | 10387710516 | [13]C5H9[15]NO2 | 650046.64
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B10 | 661174210 |  | 10112467628 | [13]C5H9[15]NO2 | 515734.65
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B11 | 598040674 |  | 9660453374 | [13]C5H9[15]NO2 | 729786.77
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_B12 | 638299067 |  | 9745524437 | [13]C5H9[15]NO2 | 1045133.10
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C5H9[15]NO2 | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C1 | 299037888 |  | 5137632667 | [13]C5H9[15]NO2 | INF
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C2 | 394801898 |  | 6729093779 | [13]C5H9[15]NO2 | INF
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C3 | 512657735 |  | 8414887004 | [13]C5H9[15]NO2 | 436485.49
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C4 | 604451135 |  | 9369400864 | [13]C5H9[15]NO2 | 516066.05
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C5 | 644217006 |  | 10108603269 | [13]C5H9[15]NO2 | 1743432.94
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C6 | 711197406 |  | 10468493900 | [13]C5H9[15]NO2 | 1346079.58
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C7 | 675916749 |  | 10519596218 | [13]C5H9[15]NO2 | 987220.46
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C8 | 686587285 |  | 10596536080 | [13]C5H9[15]NO2 | 2496387.27
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C9 | 663796540 |  | 10304708783 | [13]C5H9[15]NO2 | 950776.02
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C10 | 636114512 |  | 10252136767 | [13]C5H9[15]NO2 | 1965851.69
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C11 | 673347624 |  | 10539268728 | [13]C5H9[15]NO2 | 1279730.23
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD100520_100520_C12 | 618231080 |  | 9503018389 | [13]C5H9[15]NO2 | 1282184.28

## Serine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Serine neg | 10.04 | Target Compound | KD100520_100520_A1 | 26121228 |  | 415763174 | C3H7NO3 | 46387.68
Serine neg | 10.04 | Target Compound | KD100520_100520_A2 | 13524591 |  | 205697727 | C3H7NO3 | 12052.69
Serine neg | 10.04 | Target Compound | KD100520_100520_A3 | 8416800 |  | 133126840 | C3H7NO3 | 5156.62
Serine neg | 10.04 | Target Compound | KD100520_100520_A4 | 4298280 |  | 68983097 | C3H7NO3 | 3223.25
Serine neg | 10.04 | Target Compound | KD100520_100520_A5 | 3004778 |  | 48460428 | C3H7NO3 | 3509.92 | Calibration Amount Diff 32.177% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_A6 | 1263505 |  | 21153221 | C3H7NO3 | 1752.52
Serine neg | 10.04 | Target Compound | KD100520_100520_A7 | 526794 |  | 8463608 | C3H7NO3 | 812.19
Serine neg | 10.04 | Target Compound | KD100520_100520_A8 | 256846 |  | 4024626 | C3H7NO3 | 395.51 | Calibration Amount Diff -23.349% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_A9 | 100486 |  | 1721702 | C3H7NO3 | 234.87 | Calibration Amount Diff -47.469% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_A10 | 43250 |  | 573796 | C3H7NO3 | 78.50 | Calibration Amount Diff -93.637% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_A11 | 28134 |  | 475056 | C3H7NO3 | 58.96 | Calibration Amount Diff -104.825% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_A12 | N/F |  | N/F | C3H7NO3 | N/A
Serine neg | 10.04 | Target Compound | KD100520_100520_B1 | 27619898 |  | 446026361 | C3H7NO3 | INF
Serine neg | 10.04 | Target Compound | KD100520_100520_B2 | 13610282 |  | 224635865 | C3H7NO3 | 11138.88
Serine neg | 10.04 | Target Compound | KD100520_100520_B3 | 8122258 |  | 133857271 | C3H7NO3 | 5361.84
Serine neg | 10.04 | Target Compound | KD100520_100520_B4 | 4550157 |  | 75531941 | C3H7NO3 | 4454.92
Serine neg | 10.04 | Target Compound | KD100520_100520_B5 | 2080749 |  | 34213839 | C3H7NO3 | 2744.26
Serine neg | 10.04 | Target Compound | KD100520_100520_B6 | 1095144 |  | 18381506 | C3H7NO3 | 1185.08
Serine neg | 10.04 | Target Compound | KD100520_100520_B7 | 531724 |  | 8799442 | C3H7NO3 | 817.33
Serine neg | 10.04 | Target Compound | KD100520_100520_B8 | 250391 |  | 3932834 | C3H7NO3 | 577.73 | Calibration Amount Diff -22.142% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_B9 | 110106 |  | 1709793 | C3H7NO3 | 257.81 | Calibration Amount Diff -47.961% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_B10 | 31224 |  | 540467 | C3H7NO3 | 80.83 | Calibration Amount Diff -95.513% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_B11 | 22725 |  | 388492 | C3H7NO3 | 40.38 | Calibration Amount Diff -115.666% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_B12 | 18701 |  | 278552 | C3H7NO3 | 69.82 | Calibration Amount Diff -172.532% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C3H7NO3 | N/A
Serine neg | 10.04 | Target Compound | KD100520_100520_C1 | 25432011 |  | 390351978 | C3H7NO3 | INF
Serine neg | 10.04 | Target Compound | KD100520_100520_C2 | 14846976 |  | 235851374 | C3H7NO3 | 31256.94
Serine neg | 10.04 | Target Compound | KD100520_100520_C3 | 7601949 |  | 126798364 | C3H7NO3 | 4748.80
Serine neg | 10.04 | Target Compound | KD100520_100520_C4 | 4361487 |  | 68082861 | C3H7NO3 | 3357.30
Serine neg | 10.04 | Target Compound | KD100520_100520_C5 | 2137713 |  | 34996636 | C3H7NO3 | 2714.81
Serine neg | 10.04 | Target Compound | KD100520_100520_C6 | 911964 |  | 15755875 | C3H7NO3 | 1283.05
Serine neg | 10.04 | Target Compound | KD100520_100520_C7 | 546288 |  | 8901551 | C3H7NO3 | 617.53
Serine neg | 10.04 | Target Compound | KD100520_100520_C8 | 196925 |  | 2990853 | C3H7NO3 | 339.09 | Calibration Amount Diff -44.819% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_C9 | 97544 |  | 1438722 | C3H7NO3 | 185.02 | Calibration Amount Diff -58.273% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_C10 | 27692 |  | 441046 | C3H7NO3 | 67.37 | Calibration Amount Diff -102.961% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_C11 | 25193 |  | 423521 | C3H7NO3 | 56.09 | Calibration Amount Diff -108.923% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD100520_100520_C12 | 13303 |  | 253350 | C3H7NO3 | 52.79 | Calibration Amount Diff -176.449% (absolute) > max calibration amount diff 20%

## Serine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Serine pos | 10.04 | Target Compound | KD100520_100520_A1 | 56520253 |  | 898951671 | C3H7NO3 | INF
Serine pos | 10.04 | Target Compound | KD100520_100520_A2 | 33663330 |  | 497055159 | C3H7NO3 | INF
Serine pos | 10.04 | Target Compound | KD100520_100520_A3 | 21615284 |  | 335152406 | C3H7NO3 | INF
Serine pos | 10.04 | Target Compound | KD100520_100520_A4 | 13435608 |  | 194281829 | C3H7NO3 | 42608.01
Serine pos | 10.04 | Target Compound | KD100520_100520_A5 | 9329142 |  | 139719394 | C3H7NO3 | INF | Calibration Amount Diff 24.931% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_A6 | 4736395 |  | 70241078 | C3H7NO3 | 4640.18
Serine pos | 10.04 | Target Compound | KD100520_100520_A7 | 2193505 |  | 29861020 | C3H7NO3 | 1545.72
Serine pos | 10.04 | Target Compound | KD100520_100520_A8 | 1059342 |  | 15540198 | C3H7NO3 | 961.52 | Calibration Amount Diff 46.046% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_A9 | 451022 |  | 6830528 | C3H7NO3 | 457.95 | Calibration Amount Diff 105.696% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_A10 | 183746 |  | 2748004 | C3H7NO3 | 876.14 | Calibration Amount Diff 224.703% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_A11 | 128353 |  | 1894534 | C3H7NO3 | 362.85 | Calibration Amount Diff 516.706% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_A12 | 79787 |  | 999710 | C3H7NO3 | 127.19 | Calibration Amount Diff 1057.794% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_B1 | 63953114 |  | 969472108 | C3H7NO3 | INF
Serine pos | 10.04 | Target Compound | KD100520_100520_B2 | 34565458 |  | 539020577 | C3H7NO3 | INF
Serine pos | 10.04 | Target Compound | KD100520_100520_B3 | 22126547 |  | 333122245 | C3H7NO3 | INF
Serine pos | 10.04 | Target Compound | KD100520_100520_B4 | 14249106 |  | 210390389 | C3H7NO3 | 26646.15
Serine pos | 10.04 | Target Compound | KD100520_100520_B5 | 7318427 |  | 109687111 | C3H7NO3 | 4652.68
Serine pos | 10.04 | Target Compound | KD100520_100520_B6 | 4065954 |  | 63490309 | C3H7NO3 | 2115.20
Serine pos | 10.04 | Target Compound | KD100520_100520_B7 | 1933077 |  | 28686928 | C3H7NO3 | 1412.22
Serine pos | 10.04 | Target Compound | KD100520_100520_B8 | 1063215 |  | 15867349 | C3H7NO3 | 1000.92 | Calibration Amount Diff 49.444% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_B9 | 503308 |  | 7381345 | C3H7NO3 | 335.38 | Calibration Amount Diff 108.5% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_B10 | 169142 |  | 2611693 | C3H7NO3 | 256.32 | Calibration Amount Diff 219.442% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_B11 | 97989 |  | 1522057 | C3H7NO3 | 139.53 | Calibration Amount Diff 501.3% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_B12 | 91293 |  | 1149454 | C3H7NO3 | 162.14 | Calibration Amount Diff 1071.813% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C3H7NO3 | N/A
Serine pos | 10.04 | Target Compound | KD100520_100520_C1 | 60008483 |  | 917349144 | C3H7NO3 | INF
Serine pos | 10.04 | Target Compound | KD100520_100520_C2 | 37915546 |  | 580422712 | C3H7NO3 | INF
Serine pos | 10.04 | Target Compound | KD100520_100520_C3 | 22011738 |  | 333796694 | C3H7NO3 | 11186.14
Serine pos | 10.04 | Target Compound | KD100520_100520_C4 | 12956543 |  | 200255540 | C3H7NO3 | 8690.70
Serine pos | 10.04 | Target Compound | KD100520_100520_C5 | 7184396 |  | 119315470 | C3H7NO3 | 2361.44
Serine pos | 10.04 | Target Compound | KD100520_100520_C6 | 3719275 |  | 56615456 | C3H7NO3 | 2608.90
Serine pos | 10.04 | Target Compound | KD100520_100520_C7 | 2211919 |  | 32752027 | C3H7NO3 | 4120.42
Serine pos | 10.04 | Target Compound | KD100520_100520_C8 | 1113815 |  | 14291696 | C3H7NO3 | 818.14 | Calibration Amount Diff 37.001% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_C9 | 476498 |  | 6559065 | C3H7NO3 | 683.00 | Calibration Amount Diff 99.816% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_C10 | 151362 |  | 1994880 | C3H7NO3 | 264.86 | Calibration Amount Diff 209.616% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_C11 | 145902 |  | 1994612 | C3H7NO3 | 462.14 | Calibration Amount Diff 516.986% (absolute) > max calibration amount diff 20%
Serine pos | 10.04 | Target Compound | KD100520_100520_C12 | 66768 |  | 1010756 | C3H7NO3 | 107.08 | Calibration Amount Diff 1064.264% (absolute) > max calibration amount diff 20%

## Serine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A1 | 218094 |  | 3238823 | [13]C3H7NO3 | INF | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -29.167% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A2 | 243353 |  | 3585662 | [13]C3H7NO3 | 1511.46 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 39.762% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A3 | 297608 |  | 4540306 | [13]C3H7NO3 | 1470.90 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 58.295% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A4 | 342425 |  | 5118249 | [13]C3H7NO3 | 1412.24 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 123.715% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A5 | 375672 |  | 5486052 | [13]C3H7NO3 | 930.86 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 344.899% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A6 | 378416 |  | 6445955 | [13]C3H7NO3 | 998.39 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -424.033% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A7 | 396307 |  | 6437061 | [13]C3H7NO3 | 3369.78 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -752.879% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A8 | 431170 |  | 6726202 | [13]C3H7NO3 | 1646.49 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -2191.808% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A9 | 416249 |  | 6965206 | [13]C3H7NO3 | 3183.16 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -5639.566% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A10 | 431590 |  | 7094754 | [13]C3H7NO3 | 889.21 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -11478.11% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A11 | 416066 |  | 6846112 | [13]C3H7NO3 | 1172.95 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -10075.71% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_A12 | 383228 |  | 6402921 | [13]C3H7NO3 | 779.73 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -1802.948% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B1 | 180162 |  | 2851941 | [13]C3H7NO3 | 2705.40 | Rsquared 0.755 < R2Threshold 0.99
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B2 | 252048 |  | 3540057 | [13]C3H7NO3 | 33857.38 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 41.022% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B3 | 278465 |  | 4201549 | [13]C3H7NO3 | 1962.87 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 125.566% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B4 | 336347 |  | 5002652 | [13]C3H7NO3 | 4034.98 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 190.783% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B5 | 371000 |  | 5828528 | [13]C3H7NO3 | 1576.82 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 78.215% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B6 | 400936 |  | 6078732 | [13]C3H7NO3 | 874.21 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -65.008% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B7 | 397373 |  | 6536746 | [13]C3H7NO3 | 1259.19 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -942.503% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B8 | 420813 |  | 6343728 | [13]C3H7NO3 | 1255.86 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -1475.444% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B9 | 420110 |  | 6830058 | [13]C3H7NO3 | 897.90 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -4333.02% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B10 | 412981 |  | 6669306 | [13]C3H7NO3 | 1775.32 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -5904.09% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B11 | 383411 |  | 6305512 | [13]C3H7NO3 | 1259.50 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -2609.536% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_B12 | 394826 |  | 6578185 | [13]C3H7NO3 | 1000.83 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -12170.391% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C3H7NO3 | N/A
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C1 | 196812 |  | 2755986 | [13]C3H7NO3 | 3431.22 | Rsquared 0.755 < R2Threshold 0.99
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C2 | 243912 |  | 3600208 | [13]C3H7NO3 | 4122.67 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 23.937% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C3 | 239141 |  | 3710138 | [13]C3H7NO3 | 817.07 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 169.837% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C4 | 350540 |  | 5000481 | [13]C3H7NO3 | 621.02 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 142.933% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C5 | 344539 |  | 5343773 | [13]C3H7NO3 | 2029.61 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 55.117% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C6 | 334217 |  | 5472011 | [13]C3H7NO3 | 1086.69 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -175.253% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C7 | 382557 |  | 6070467 | [13]C3H7NO3 | 1395.72 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 275.824% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C8 | 353175 |  | 5902048 | [13]C3H7NO3 | 1956.03 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 1486.494% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C9 | 396627 |  | 6529342 | [13]C3H7NO3 | 1618.36 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -2695.432% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C10 | 385682 |  | 6259256 | [13]C3H7NO3 | 2041.11 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -1314.181% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C11 | 396764 |  | 6121868 | [13]C3H7NO3 | 1281.46 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff 2710.273% (absolute) > max calibration amount diff 20%
Serine U-13C neg | 10.04 | Target Compound | KD100520_100520_C12 | 379113 |  | 6327918 | [13]C3H7NO3 | 1029.95 | Rsquared 0.755 < R2Threshold 0.99. Calibration Amount Diff -19296.879% (absolute) > max calibration amount diff 20%

## Serine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A1 | 433531 |  | 5318480 | [13]C3H7NO3 | INF | Rsquared 0.536 < R2Threshold 0.99
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A2 | 606453 |  | 7129191 | [13]C3H7NO3 | 7173.83 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 82.214% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A3 | 745324 |  | 9558399 | [13]C3H7NO3 | 2165.92 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 148.048% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A4 | 1030356 |  | 12448171 | [13]C3H7NO3 | 3050.02 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 330.694% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A5 | 1096019 |  | 14290838 | [13]C3H7NO3 | 2757.82 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 690.783% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A6 | 1469745 |  | 19995936 | [13]C3H7NO3 | 3093.00 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 40.935% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A7 | 1735387 |  | 22294688 | [13]C3H7NO3 | 5087.37 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -778.871% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A8 | 1808068 |  | 26056928 | [13]C3H7NO3 | 6266.03 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -3326.054% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A9 | 1606964 |  | 26254810 | [13]C3H7NO3 | 2995.81 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -8591.733% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A10 | 1765854 |  | 27406845 | [13]C3H7NO3 | 3532.32 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -17145.422% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A11 | 1684247 |  | 26048547 | [13]C3H7NO3 | 2830.84 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -21414.005% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_A12 | 1972139 |  | 28312246 | [13]C3H7NO3 | 7461.23 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -56625.115% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B1 | 532882 |  | 5548828 | [13]C3H7NO3 | 24905.11 | Rsquared 0.536 < R2Threshold 0.99
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B2 | 672900 |  | 9470255 | [13]C3H7NO3 | 1215.49 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -21.294% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B3 | 615720 |  | 9308873 | [13]C3H7NO3 | 6006.39 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 176.95% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B4 | 959335 |  | 12227949 | [13]C3H7NO3 | 2357.98 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 419.328% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B5 | 1264686 |  | 17574984 | [13]C3H7NO3 | 4034.34 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 317.661% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B6 | 1406218 |  | 21108915 | [13]C3H7NO3 | 1550.06 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -36.192% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B7 | 1306845 |  | 19862631 | [13]C3H7NO3 | 3874.97 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 1075% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B8 | 1640431 |  | 23910599 | [13]C3H7NO3 | 2833.09 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -1160.334% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B9 | 1722942 |  | 25555408 | [13]C3H7NO3 | 3313.46 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -4057.022% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B10 | 1861918 |  | 28908399 | [13]C3H7NO3 | 4595.05 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -17467.186% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B11 | 1668921 |  | 26415945 | [13]C3H7NO3 | 2887.64 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -22506.1% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_B12 | 1768018 |  | 25897122 | [13]C3H7NO3 | 1791.47 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -25383.391% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C3H7NO3 | N/A
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C1 | 577271 |  | 7015398 | [13]C3H7NO3 | 5298.16 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -47.857% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C2 | 494014 |  | 7413715 | [13]C3H7NO3 | 616.93 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 56.88% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C3 | 682652 |  | 10213623 | [13]C3H7NO3 | 1026.68 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 124.26% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C4 | 938928 |  | 13604505 | [13]C3H7NO3 | 1756.72 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 266.375% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C5 | 1239568 |  | 18246173 | [13]C3H7NO3 | 2437.87 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 104.691% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C6 | 1289876 |  | 19096063 | [13]C3H7NO3 | 1737.43 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 310.931% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C7 | 1688606 |  | 22008968 | [13]C3H7NO3 | 2089.11 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff 719.376% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C8 | 1829964 |  | 25519494 | [13]C3H7NO3 | 2342.09 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -1313.739% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C9 | 1661790 |  | 25787846 | [13]C3H7NO3 | 2517.81 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -4117.543% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C10 | 1716588 |  | 25323930 | [13]C3H7NO3 | 1592.31 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -5560.247% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C11 | 1906380 |  | 27979189 | [13]C3H7NO3 | 3355.55 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -27983.846% (absolute) > max calibration amount diff 20%
Serine U-13C pos | 10.04 | Target Compound | KD100520_100520_C12 | 1678318 |  | 25290481 | [13]C3H7NO3 | 2532.49 | Rsquared 0.536 < R2Threshold 0.99. Calibration Amount Diff -34052.462% (absolute) > max calibration amount diff 20%

## Serine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A1 | 3216598 |  | 47191723 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A2 | 3519501 |  | 51960185 | [13]C3H7[15]NO3 | 11174.41
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A3 | 3663726 |  | 56048377 | [13]C3H7[15]NO3 | 20511.72
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A4 | 3683455 |  | 59790387 | [13]C3H7[15]NO3 | 12156.74
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A5 | 3949178 |  | 64040120 | [13]C3H7[15]NO3 | 8574.09
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A6 | 3822097 |  | 64021832 | [13]C3H7[15]NO3 | 10887.85
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A7 | 3603486 |  | 63914438 | [13]C3H7[15]NO3 | 14840.57
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A8 | 3980702 |  | 65211543 | [13]C3H7[15]NO3 | 12166.95
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A9 | 3812655 |  | 66183242 | [13]C3H7[15]NO3 | 3867.74
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A10 | 3854026 |  | 67263578 | [13]C3H7[15]NO3 | 4934.04
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A11 | 3905111 |  | 68100611 | [13]C3H7[15]NO3 | 6765.65
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_A12 | 3689541 |  | 66006985 | [13]C3H7[15]NO3 | 3414.78
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B1 | 2983794 |  | 44751554 | [13]C3H7[15]NO3 | 18936.61
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B2 | 3311347 |  | 51485485 | [13]C3H7[15]NO3 | 203463.03
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B3 | 3572678 |  | 56511490 | [13]C3H7[15]NO3 | 57010.14
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B4 | 3660672 |  | 60797313 | [13]C3H7[15]NO3 | 23655.77
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B5 | 3777783 |  | 63171733 | [13]C3H7[15]NO3 | 11225.33
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B6 | 3703477 |  | 63161073 | [13]C3H7[15]NO3 | 7319.04
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B7 | 3756520 |  | 64156926 | [13]C3H7[15]NO3 | 14428.52
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B8 | 3737494 |  | 62853540 | [13]C3H7[15]NO3 | 8160.09
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B9 | 3849084 |  | 66169435 | [13]C3H7[15]NO3 | 4143.66
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B10 | 3722549 |  | 65925175 | [13]C3H7[15]NO3 | 2531.72
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B11 | 3685492 |  | 64576299 | [13]C3H7[15]NO3 | 4332.69
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_B12 | 3945007 |  | 66441800 | [13]C3H7[15]NO3 | 5596.58
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C3H7[15]NO3 | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C1 | 3054452 |  | 45165913 | [13]C3H7[15]NO3 | 12324.70
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C2 | 3187038 |  | 49905478 | [13]C3H7[15]NO3 | 24644.69
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C3 | 3327291 |  | 53027171 | [13]C3H7[15]NO3 | 23108.24
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C4 | 3713954 |  | 59070954 | [13]C3H7[15]NO3 | 9064.63
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C5 | 3333806 |  | 57561201 | [13]C3H7[15]NO3 | 29694.85
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C6 | 3226894 |  | 56062399 | [13]C3H7[15]NO3 | 4147.36
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C7 | 3655020 |  | 64340135 | [13]C3H7[15]NO3 | 12585.12
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C8 | 3739639 |  | 64314987 | [13]C3H7[15]NO3 | 6639.65
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C9 | 3586237 |  | 64847057 | [13]C3H7[15]NO3 | 35075.40
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C10 | 3742818 |  | 64122539 | [13]C3H7[15]NO3 | 18493.00
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C11 | 3894133 |  | 64043812 | [13]C3H7[15]NO3 | 9432.11
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD100520_100520_C12 | 3499285 |  | 63037370 | [13]C3H7[15]NO3 | 4153.99

## Serine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A1 | 6321938 |  | 88497767 | [13]C3H7[15]NO3 | INF | 10.090:Peak area 88497766.727 is out of bounds (ISTD Minimum Recovery 103306760.455 and ISTD Max Recovery 309920281.366)
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A2 | 8378274 |  | 121000402 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A3 | 9270348 |  | 136602136 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A4 | 11634009 |  | 170339731 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A5 | 12757705 |  | 191168941 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A6 | 14040846 |  | 217680870 | [13]C3H7[15]NO3 | 8163.59
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A7 | 15737435 |  | 228517157 | [13]C3H7[15]NO3 | 32686.95
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A8 | 16408688 |  | 252717483 | [13]C3H7[15]NO3 | 135575.48
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A9 | 14517352 |  | 247374891 | [13]C3H7[15]NO3 | 28393.29
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A10 | 15957554 |  | 258101914 | [13]C3H7[15]NO3 | 50575.60
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A11 | 16185538 |  | 256799119 | [13]C3H7[15]NO3 | 11625.93
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_A12 | 17454175 |  | 272201616 | [13]C3H7[15]NO3 | 136635.90
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B1 | 7155903 |  | 89169606 | [13]C3H7[15]NO3 | INF | 10.103:Peak area 89169605.695 is out of bounds (ISTD Minimum Recovery 103306760.455 and ISTD Max Recovery 309920281.366)
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B2 | 8029448 |  | 120515368 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B3 | 9712683 |  | 138468428 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B4 | 12332455 |  | 177553042 | [13]C3H7[15]NO3 | 78121.46
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B5 | 13418572 |  | 210189807 | [13]C3H7[15]NO3 | 11273.85
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B6 | 14058729 |  | 227528910 | [13]C3H7[15]NO3 | 40264.50
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B7 | 14066439 |  | 229475496 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B8 | 16189834 |  | 247318907 | [13]C3H7[15]NO3 | 97171.95
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B9 | 16745298 |  | 257085506 | [13]C3H7[15]NO3 | 17450.04
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B10 | 16907161 |  | 271631651 | [13]C3H7[15]NO3 | 34476.54
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B11 | 15674360 |  | 259385585 | [13]C3H7[15]NO3 | 83084.54
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_B12 | 17520570 |  | 263665025 | [13]C3H7[15]NO3 | 26409.01
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C3H7[15]NO3 | N/A
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C1 | 7122047 |  | 95160123 | [13]C3H7[15]NO3 | INF | 10.077:Peak area 95160123.240 is out of bounds (ISTD Minimum Recovery 103306760.455 and ISTD Max Recovery 309920281.366)
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C2 | 8783579 |  | 116327978 | [13]C3H7[15]NO3 | 167099.48
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C3 | 9739478 |  | 141400397 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C4 | 11689905 |  | 178693956 | [13]C3H7[15]NO3 | INF
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C5 | 13060644 |  | 205771976 | [13]C3H7[15]NO3 | 14977.68
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C6 | 13684432 |  | 215401334 | [13]C3H7[15]NO3 | 29989.39
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C7 | 16466761 |  | 248218085 | [13]C3H7[15]NO3 | 756909.31
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C8 | 17600884 |  | 262723225 | [13]C3H7[15]NO3 | 18562.19
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C9 | 15247849 |  | 259189640 | [13]C3H7[15]NO3 | 142721.68
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C10 | 16363887 |  | 259504499 | [13]C3H7[15]NO3 | 18446.44
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C11 | 17437340 |  | 269363189 | [13]C3H7[15]NO3 | 18162.61
Serine U-13C, U-15N pos | 10.04 | Internal Standard | KD100520_100520_C12 | 15398062 |  | 253342993 | [13]C3H7[15]NO3 | 11929.34

## Threonine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Threonine neg | 9.15 | Target Compound | KD100520_100520_A1 | 54257376 |  | 830479551 | C4H9NO3 | 56731.81
Threonine neg | 9.15 | Target Compound | KD100520_100520_A2 | 30370652 |  | 450838593 | C4H9NO3 | 19108.58
Threonine neg | 9.15 | Target Compound | KD100520_100520_A3 | 21846910 |  | 338418990 | C4H9NO3 | 7626.16
Threonine neg | 9.15 | Target Compound | KD100520_100520_A4 | 15403007 |  | 216941337 | C4H9NO3 | 11612.25
Threonine neg | 9.15 | Target Compound | KD100520_100520_A5 | 10614101 |  | 160171813 | C4H9NO3 | 7901.84 | Calibration Amount Diff 34.728% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_A6 | 5387006 |  | 80929773 | C4H9NO3 | 5825.16
Threonine neg | 9.15 | Target Compound | KD100520_100520_A7 | 2513163 |  | 38645703 | C4H9NO3 | 4038.24
Threonine neg | 9.15 | Target Compound | KD100520_100520_A8 | 1330707 |  | 20004211 | C4H9NO3 | 1504.22
Threonine neg | 9.15 | Target Compound | KD100520_100520_A9 | 756900 |  | 11160523 | C4H9NO3 | 1071.90 | Calibration Amount Diff -35.98% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_A10 | 354415 |  | 5260666 | C4H9NO3 | 578.41 | Calibration Amount Diff -107.479% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_A11 | 180110 |  | 2812017 | C4H9NO3 | 388.85 | Calibration Amount Diff -221.07% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_A12 | 72734 |  | 1172794 | C4H9NO3 | 198.00 | Calibration Amount Diff -491.15% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_B1 | 52182789 |  | 816361935 | C4H9NO3 | 42050.30
Threonine neg | 9.15 | Target Compound | KD100520_100520_B2 | 28718706 |  | 418043987 | C4H9NO3 | 14162.67
Threonine neg | 9.15 | Target Compound | KD100520_100520_B3 | 20075681 |  | 307180072 | C4H9NO3 | 6978.68
Threonine neg | 9.15 | Target Compound | KD100520_100520_B4 | 13145552 |  | 200010001 | C4H9NO3 | 9255.41
Threonine neg | 9.15 | Target Compound | KD100520_100520_B5 | 8331434 |  | 126334570 | C4H9NO3 | 6216.91
Threonine neg | 9.15 | Target Compound | KD100520_100520_B6 | 4329808 |  | 65745549 | C4H9NO3 | 3058.81
Threonine neg | 9.15 | Target Compound | KD100520_100520_B7 | 2467527 |  | 37014938 | C4H9NO3 | 3688.51
Threonine neg | 9.15 | Target Compound | KD100520_100520_B8 | 1193772 |  | 19776363 | C4H9NO3 | 2044.67
Threonine neg | 9.15 | Target Compound | KD100520_100520_B9 | 657184 |  | 10280534 | C4H9NO3 | 1318.48 | Calibration Amount Diff -36.537% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_B10 | 306831 |  | 4580201 | C4H9NO3 | 625.59 | Calibration Amount Diff -115.379% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_B11 | 158531 |  | 2183204 | C4H9NO3 | 321.13 | Calibration Amount Diff -240.768% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_B12 | 64542 |  | 966710 | C4H9NO3 | 144.04 | Calibration Amount Diff -505.943% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C4H9NO3 | N/A
Threonine neg | 9.15 | Target Compound | KD100520_100520_C1 | 43832384 |  | 674838872 | C4H9NO3 | 178771.55
Threonine neg | 9.15 | Target Compound | KD100520_100520_C2 | 27294053 |  | 396487070 | C4H9NO3 | 8301.97
Threonine neg | 9.15 | Target Compound | KD100520_100520_C3 | 18084992 |  | 264114180 | C4H9NO3 | 8462.86
Threonine neg | 9.15 | Target Compound | KD100520_100520_C4 | 12128364 |  | 177049206 | C4H9NO3 | 9028.63
Threonine neg | 9.15 | Target Compound | KD100520_100520_C5 | 7037944 |  | 110140274 | C4H9NO3 | 6635.71
Threonine neg | 9.15 | Target Compound | KD100520_100520_C6 | 3544203 |  | 57865970 | C4H9NO3 | 4366.82
Threonine neg | 9.15 | Target Compound | KD100520_100520_C7 | 2352247 |  | 35547022 | C4H9NO3 | 2959.61
Threonine neg | 9.15 | Target Compound | KD100520_100520_C8 | 1163780 |  | 17692338 | C4H9NO3 | 2176.87 | Calibration Amount Diff -20.836% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_C9 | 635154 |  | 9072236 | C4H9NO3 | 805.73 | Calibration Amount Diff -48.492% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_C10 | 241212 |  | 3675252 | C4H9NO3 | 411.80 | Calibration Amount Diff -132.599% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_C11 | 138026 |  | 2099355 | C4H9NO3 | 293.28 | Calibration Amount Diff -246.117% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD100520_100520_C12 | 56736 |  | 939031 | C4H9NO3 | 156.94 | Calibration Amount Diff -501.263% (absolute) > max calibration amount diff 20%

## Threonine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Threonine pos | 9.15 | Target Compound | KD100520_100520_A1 | 314930183 |  | 3776772377 | C4H9NO3 | INF
Threonine pos | 9.15 | Target Compound | KD100520_100520_A2 | 185651253 |  | 2446084499 | C4H9NO3 | 26513.67
Threonine pos | 9.15 | Target Compound | KD100520_100520_A3 | 130811078 |  | 1766095394 | C4H9NO3 | 33885.33
Threonine pos | 9.15 | Target Compound | KD100520_100520_A4 | 65505212 |  | 952172969 | C4H9NO3 | 14570.00
Threonine pos | 9.15 | Target Compound | KD100520_100520_A5 | 50122228 |  | 718506091 | C4H9NO3 | 21176.19 | Calibration Amount Diff 36.428% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_A6 | 26559179 |  | 360917664 | C4H9NO3 | 11620.11
Threonine pos | 9.15 | Target Compound | KD100520_100520_A7 | 12468622 |  | 176023162 | C4H9NO3 | 4095.69
Threonine pos | 9.15 | Target Compound | KD100520_100520_A8 | 5913572 |  | 87352936 | C4H9NO3 | 2149.80 | Calibration Amount Diff -38.964% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_A9 | 3730697 |  | 50288660 | C4H9NO3 | 1245.27 | Calibration Amount Diff -93.062% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_A10 | 1799484 |  | 24270463 | C4H9NO3 | 648.01 | Calibration Amount Diff -233.251% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_A11 | 933660 |  | 13058925 | C4H9NO3 | 486.80 | Calibration Amount Diff -478.907% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_A12 | 402120 |  | 5401669 | C4H9NO3 | 162.01 | Calibration Amount Diff -1031.903% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_B1 | 356311405 |  | 3999794085 | C4H9NO3 | 85445.16
Threonine pos | 9.15 | Target Compound | KD100520_100520_B2 | 205638410 |  | 2529864316 | C4H9NO3 | 25372.18
Threonine pos | 9.15 | Target Compound | KD100520_100520_B3 | 117911302 |  | 1690732259 | C4H9NO3 | 16436.79
Threonine pos | 9.15 | Target Compound | KD100520_100520_B4 | 79905079 |  | 1071063559 | C4H9NO3 | 15801.61
Threonine pos | 9.15 | Target Compound | KD100520_100520_B5 | 41983570 |  | 575947417 | C4H9NO3 | 14256.59
Threonine pos | 9.15 | Target Compound | KD100520_100520_B6 | 23397964 |  | 329341696 | C4H9NO3 | 9093.84
Threonine pos | 9.15 | Target Compound | KD100520_100520_B7 | 11734931 |  | 175514419 | C4H9NO3 | 4916.31
Threonine pos | 9.15 | Target Compound | KD100520_100520_B8 | 7032994 |  | 97291361 | C4H9NO3 | 3695.30 | Calibration Amount Diff -34.597% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_B9 | 3462026 |  | 50269425 | C4H9NO3 | 1358.62 | Calibration Amount Diff -89.07% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_B10 | 1620141 |  | 22422329 | C4H9NO3 | 625.65 | Calibration Amount Diff -234.884% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_B11 | 878059 |  | 12694197 | C4H9NO3 | 311.38 | Calibration Amount Diff -486.178% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_B12 | 471749 |  | 6089992 | C4H9NO3 | 479.27 | Calibration Amount Diff -1011.842% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C4H9NO3 | N/A
Threonine pos | 9.15 | Target Compound | KD100520_100520_C1 | 346824510 |  | 4128607372 | C4H9NO3 | 110061.26
Threonine pos | 9.15 | Target Compound | KD100520_100520_C2 | 187213337 |  | 2594265806 | C4H9NO3 | 29803.18
Threonine pos | 9.15 | Target Compound | KD100520_100520_C3 | 127690979 |  | 1619174848 | C4H9NO3 | 32485.40
Threonine pos | 9.15 | Target Compound | KD100520_100520_C4 | 65450067 |  | 960399968 | C4H9NO3 | 25808.47
Threonine pos | 9.15 | Target Compound | KD100520_100520_C5 | 39986621 |  | 572310138 | C4H9NO3 | 14524.83
Threonine pos | 9.15 | Target Compound | KD100520_100520_C6 | 20522336 |  | 301044450 | C4H9NO3 | 7418.72
Threonine pos | 9.15 | Target Compound | KD100520_100520_C7 | 11430374 |  | 174004931 | C4H9NO3 | 4275.14
Threonine pos | 9.15 | Target Compound | KD100520_100520_C8 | 6318609 |  | 87848746 | C4H9NO3 | 2270.26 | Calibration Amount Diff -45.908% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_C9 | 3363917 |  | 47810028 | C4H9NO3 | 1248.78 | Calibration Amount Diff -102.145% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_C10 | 1491932 |  | 20119964 | C4H9NO3 | 754.26 | Calibration Amount Diff -247.448% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_C11 | 834002 |  | 11419441 | C4H9NO3 | 429.30 | Calibration Amount Diff -492.922% (absolute) > max calibration amount diff 20%
Threonine pos | 9.15 | Target Compound | KD100520_100520_C12 | 426146 |  | 5575191 | C4H9NO3 | 245.31 | Calibration Amount Diff -1022.277% (absolute) > max calibration amount diff 20%

## Threonine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A1 | 372534 |  | 5222648 | [13]C4H9NO3 | 740.73 | Rsquared 0.923 < R2Threshold 0.99
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A2 | 569888 |  | 7334797 | [13]C4H9NO3 | 832.22 | Rsquared 0.923 < R2Threshold 0.99
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A3 | 716180 |  | 10461611 | [13]C4H9NO3 | 6672.67 | Rsquared 0.923 < R2Threshold 0.99
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A4 | 1029057 |  | 13460775 | [13]C4H9NO3 | 6226.05 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 111.135% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A5 | 1067012 |  | 15740672 | [13]C4H9NO3 | 2493.12 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 130.759% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A6 | 1275042 |  | 18834098 | [13]C4H9NO3 | 2251.05 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -277.042% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A7 | 1271872 |  | 19948999 | [13]C4H9NO3 | 1942.08 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -643.928% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A8 | 1313976 |  | 20428766 | [13]C4H9NO3 | 2271.79 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -1634.062% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A9 | 1366327 |  | 20942664 | [13]C4H9NO3 | 1551.64 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -1930.298% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A10 | 1389031 |  | 21245200 | [13]C4H9NO3 | 1247.14 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -3740.463% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A11 | 1251653 |  | 20111621 | [13]C4H9NO3 | 1507.26 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 27.807% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_A12 | 1233830 |  | 19674037 | [13]C4H9NO3 | 1095.14 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 3766.587% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B1 | 312478 |  | 4337070 | [13]C4H9NO3 | 1520.89 | Rsquared 0.923 < R2Threshold 0.99
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B2 | 473731 |  | 5616461 | [13]C4H9NO3 | 2717.79 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 44.705% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B3 | 701525 |  | 9405178 | [13]C4H9NO3 | 1434.73 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 32.517% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B4 | 861305 |  | 12238755 | [13]C4H9NO3 | 8170.78 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 40.005% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B5 | 1199480 |  | 17813143 | [13]C4H9NO3 | 3399.89 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -92.608% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B6 | 1112231 |  | 16660397 | [13]C4H9NO3 | 2155.62 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -22.758% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B7 | 1267124 |  | 19016156 | [13]C4H9NO3 | 2233.49 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -224.021% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B8 | 1135456 |  | 19083498 | [13]C4H9NO3 | 1222.33 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -462.125% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B9 | 1163832 |  | 19095195 | [13]C4H9NO3 | 1159.41 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -952.54% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B10 | 1254070 |  | 19782228 | [13]C4H9NO3 | 1283.09 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -3415.651% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B11 | 1100627 |  | 18085351 | [13]C4H9NO3 | 1145.81 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 4825.259% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_B12 | 1174659 |  | 18831958 | [13]C4H9NO3 | 1758.06 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 1548.503% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C4H9NO3 | N/A
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C1 | 290734 |  | 3770503 | [13]C4H9NO3 | INF | Rsquared 0.923 < R2Threshold 0.99
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C2 | 420583 |  | 5212373 | [13]C4H9NO3 | 1760.08 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 21.9% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C3 | 592260 |  | 7913821 | [13]C4H9NO3 | 1593.97 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 36.879% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C4 | 869580 |  | 11864175 | [13]C4H9NO3 | 2265.74 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -56.86% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C5 | 969122 |  | 14134935 | [13]C4H9NO3 | 2467.05 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -56.604% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C6 | 1026132 |  | 16384396 | [13]C4H9NO3 | 2194.91 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -351.392% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C7 | 1151507 |  | 17270237 | [13]C4H9NO3 | 1689.14 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 194.328% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C8 | 1193795 |  | 18370690 | [13]C4H9NO3 | 1702.69 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 140.295% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C9 | 1186447 |  | 17581501 | [13]C4H9NO3 | 1162.47 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 2738.709% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C10 | 1081666 |  | 17760772 | [13]C4H9NO3 | 1641.75 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 4215.61% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C11 | 1178991 |  | 18684493 | [13]C4H9NO3 | 2005.40 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff -811.005% (absolute) > max calibration amount diff 20%
Threonine U-13C neg | 9.15 | Target Compound | KD100520_100520_C12 | 1056109 |  | 17238277 | [13]C4H9NO3 | 1626.66 | Rsquared 0.923 < R2Threshold 0.99. Calibration Amount Diff 7317.24% (absolute) > max calibration amount diff 20%

## Threonine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A1 | 1772045 |  | 20202488 | [13]C4H9NO3 | 171.09 | Rsquared 0.895 < R2Threshold 0.99
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A2 | 3605223 |  | 38798050 | [13]C4H9NO3 | 420.56 | Rsquared 0.895 < R2Threshold 0.99
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A3 | 4039170 |  | 45850224 | [13]C4H9NO3 | 526.84 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 118.22% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A4 | 4369611 |  | 55337516 | [13]C4H9NO3 | 1092.80 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 153.889% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A5 | 4928270 |  | 66487724 | [13]C4H9NO3 | 1651.49 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 222.002% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A6 | 5957872 |  | 78026026 | [13]C4H9NO3 | 2433.71 | Rsquared 0.895 < R2Threshold 0.99
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A7 | 6252678 |  | 86553116 | [13]C4H9NO3 | 5201.33 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -427.681% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A8 | 5616970 |  | 83236968 | [13]C4H9NO3 | 6430.80 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -1186.244% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A9 | 6300436 |  | 87424021 | [13]C4H9NO3 | 9387.09 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -2160.767% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A10 | 6622823 |  | 91374262 | [13]C4H9NO3 | 1496.40 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -6734.181% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A11 | 5852965 |  | 84020387 | [13]C4H9NO3 | 33828.16 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -7730.023% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_A12 | 5452684 |  | 81551826 | [13]C4H9NO3 | 2621.48 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -7996.464% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B1 | 2109341 |  | 21253758 | [13]C4H9NO3 | 273.56 | Rsquared 0.895 < R2Threshold 0.99
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B2 | 3473178 |  | 34832815 | [13]C4H9NO3 | 431.02 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 22.683% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B3 | 3898281 |  | 47662865 | [13]C4H9NO3 | 599.68 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 61.378% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B4 | 5119142 |  | 62178463 | [13]C4H9NO3 | 1664.63 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 84.134% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B5 | 5744925 |  | 77007512 | [13]C4H9NO3 | 2213.71 | Rsquared 0.895 < R2Threshold 0.99
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B6 | 6001017 |  | 79869915 | [13]C4H9NO3 | 2135.07 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 50.082% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B7 | 6000462 |  | 87950507 | [13]C4H9NO3 | 6356.67 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -554.718% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B8 | 6244227 |  | 87934501 | [13]C4H9NO3 | 8255.37 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -588.18% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B9 | 5721365 |  | 84551336 | [13]C4H9NO3 | 6189.23 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -1862.432% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B10 | 5849339 |  | 83073062 | [13]C4H9NO3 | 2562.40 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -3103.035% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B11 | 5683383 |  | 83128844 | [13]C4H9NO3 | 895.47 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 217.602% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_B12 | 5382981 |  | 78343111 | [13]C4H9NO3 | 1279.72 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -1882.106% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_blank | 847 |  | 3189 | [13]C4H9NO3 | INF
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C1 | 2336417 |  | 26471698 | [13]C4H9NO3 | 188.16 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -22.001% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C2 | 2668804 |  | 32099080 | [13]C4H9NO3 | 704.58 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 30.536% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C3 | 4450711 |  | 47355554 | [13]C4H9NO3 | 1413.02 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 29.271% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C4 | 4667549 |  | 60819507 | [13]C4H9NO3 | 839.21 | Rsquared 0.895 < R2Threshold 0.99
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C5 | 5267033 |  | 71253642 | [13]C4H9NO3 | 1565.35 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -77.871% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C6 | 5578841 |  | 80727656 | [13]C4H9NO3 | 2612.44 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -252.023% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C7 | 5380967 |  | 79818196 | [13]C4H9NO3 | 4523.94 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff 236.594% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C8 | 6225887 |  | 85410502 | [13]C4H9NO3 | 5250.64 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -65.263% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C9 | 6044747 |  | 86042560 | [13]C4H9NO3 | 8662.49 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -115.093% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C10 | 5560726 |  | 81506937 | [13]C4H9NO3 | 13312.40 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -1499.339% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C11 | 5252680 |  | 79409147 | [13]C4H9NO3 | 12053.41 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -2649.083% (absolute) > max calibration amount diff 20%
Threonine U-13C pos | 9.15 | Target Compound | KD100520_100520_C12 | 5156105 |  | 77057447 | [13]C4H9NO3 | 1192.77 | Rsquared 0.895 < R2Threshold 0.99. Calibration Amount Diff -738.44% (absolute) > max calibration amount diff 20%

## Threonine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A1 | 6986604 |  | 98860483 | [13]C4H9[15]NO3 | INF | 9.240:Peak area 98860483.245 is out of bounds (ISTD Minimum Recovery 100873248.744 and ISTD Max Recovery 302619746.232)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A2 | 8525449 |  | 121220390 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A3 | 9681857 |  | 151022216 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A4 | 13379396 |  | 195187444 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A5 | 13790830 |  | 216929099 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A6 | 15658697 |  | 239741969 | [13]C4H9[15]NO3 | 50071.20
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A7 | 16009835 |  | 251083321 | [13]C4H9[15]NO3 | 27583.12
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A8 | 15985027 |  | 253776010 | [13]C4H9[15]NO3 | 10178.93
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A9 | 16907850 |  | 264943373 | [13]C4H9[15]NO3 | 12275.57
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A10 | 16860206 |  | 268807197 | [13]C4H9[15]NO3 | 17547.61
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A11 | 15973586 |  | 261668102 | [13]C4H9[15]NO3 | 26916.42
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_A12 | 15573014 |  | 257751177 | [13]C4H9[15]NO3 | 37391.20
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B1 | 5978874 |  | 88794650 | [13]C4H9[15]NO3 | INF | 9.226:Peak area 88794649.984 is out of bounds (ISTD Minimum Recovery 100873248.744 and ISTD Max Recovery 302619746.232)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B2 | 7403947 |  | 101844230 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B3 | 9420875 |  | 140505730 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B4 | 11084309 |  | 170846317 | [13]C4H9[15]NO3 | 29665.53
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B5 | 14851843 |  | 232069145 | [13]C4H9[15]NO3 | 59968.07
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B6 | 13605760 |  | 218723082 | [13]C4H9[15]NO3 | 15565.95
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B7 | 15664472 |  | 245435232 | [13]C4H9[15]NO3 | 66175.13
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B8 | 14323563 |  | 245455917 | [13]C4H9[15]NO3 | 30614.65
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B9 | 15139067 |  | 245131440 | [13]C4H9[15]NO3 | 9696.40
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B10 | 15944833 |  | 250902444 | [13]C4H9[15]NO3 | 33283.68
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B11 | 15002520 |  | 239694191 | [13]C4H9[15]NO3 | 18914.28
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_B12 | 15608080 |  | 245671783 | [13]C4H9[15]NO3 | 14077.81
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C4H9[15]NO3 | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C1 | 5461958 |  | 81973617 | [13]C4H9[15]NO3 | INF | 9.256:Peak area 81973616.769 is out of bounds (ISTD Minimum Recovery 100873248.744 and ISTD Max Recovery 302619746.232)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C2 | 6407719 |  | 88986161 | [13]C4H9[15]NO3 | INF | 9.181:Peak area 88986161.219 is out of bounds (ISTD Minimum Recovery 100873248.744 and ISTD Max Recovery 302619746.232)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C3 | 8174740 |  | 118807790 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C4 | 10963046 |  | 157609446 | [13]C4H9[15]NO3 | 179185.44
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C5 | 11941916 |  | 185786832 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C6 | 12596707 |  | 206721532 | [13]C4H9[15]NO3 | 18351.55
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C7 | 14241002 |  | 228701668 | [13]C4H9[15]NO3 | 32943.59
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C8 | 15529207 |  | 240666725 | [13]C4H9[15]NO3 | 20565.41
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C9 | 15607670 |  | 238996928 | [13]C4H9[15]NO3 | 8876.64
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C10 | 14123191 |  | 238842378 | [13]C4H9[15]NO3 | 14383.95
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C11 | 15407887 |  | 242324139 | [13]C4H9[15]NO3 | 28158.26
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD100520_100520_C12 | 13716251 |  | 227391752 | [13]C4H9[15]NO3 | 16690.88

## Threonine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A1 | 41135803 |  | 474595186 | [13]C4H9[15]NO3 | INF | 9.227:Peak area 474595186.206 is out of bounds (ISTD Minimum Recovery 484415178.752 and ISTD Max Recovery 1453245536.257)
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A2 | 52674757 |  | 677307833 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A3 | 63228827 |  | 808529667 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A4 | 60969609 |  | 868300772 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A5 | 68944766 |  | 987869149 | [13]C4H9[15]NO3 | 143252.17
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A6 | 78496367 |  | 1074126352 | [13]C4H9[15]NO3 | 656884.15
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A7 | 79436667 |  | 1152909291 | [13]C4H9[15]NO3 | 96038.75
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A8 | 74009626 |  | 1093649363 | [13]C4H9[15]NO3 | 57508.18
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A9 | 80053394 |  | 1150693456 | [13]C4H9[15]NO3 | 77685.43
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A10 | 83943402 |  | 1179215402 | [13]C4H9[15]NO3 | 44041.94
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A11 | 75198045 |  | 1108599528 | [13]C4H9[15]NO3 | 120771.43
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_A12 | 70849348 |  | 1091961908 | [13]C4H9[15]NO3 | 21089.07
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B1 | 41269575 |  | 465501200 | [13]C4H9[15]NO3 | INF | 9.240:Peak area 465501199.508 is out of bounds (ISTD Minimum Recovery 484415178.752 and ISTD Max Recovery 1453245536.257)
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B2 | 54746986 |  | 637736158 | [13]C4H9[15]NO3 | 642800.84
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B3 | 58769390 |  | 780315609 | [13]C4H9[15]NO3 | 120279.15
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B4 | 70168167 |  | 936221466 | [13]C4H9[15]NO3 | 62806.30
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B5 | 75376829 |  | 1070977368 | [13]C4H9[15]NO3 | 67685.17
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B6 | 78240960 |  | 1108466355 | [13]C4H9[15]NO3 | 39255.98
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B7 | 77385926 |  | 1162066463 | [13]C4H9[15]NO3 | 123889.31
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B8 | 81764598 |  | 1177617762 | [13]C4H9[15]NO3 | 49065.56
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B9 | 74720636 |  | 1118160659 | [13]C4H9[15]NO3 | 89477.30
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B10 | 75978063 |  | 1103173313 | [13]C4H9[15]NO3 | 18472.67
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B11 | 76612536 |  | 1132507990 | [13]C4H9[15]NO3 | 11515.36
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_B12 | 71262736 |  | 1062056556 | [13]C4H9[15]NO3 | 23698.20
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C4H9[15]NO3 | N/A
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C1 | 47109433 |  | 534854654 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C2 | 45698418 |  | 600981356 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C3 | 62759161 |  | 745149278 | [13]C4H9[15]NO3 | 201472.66
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C4 | 59105329 |  | 865272983 | [13]C4H9[15]NO3 | INF
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C5 | 67522498 |  | 975112814 | [13]C4H9[15]NO3 | 101238.54
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C6 | 73746254 |  | 1076944574 | [13]C4H9[15]NO3 | 53755.43
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C7 | 72725023 |  | 1110449934 | [13]C4H9[15]NO3 | 74103.83
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C8 | 79632271 |  | 1163403862 | [13]C4H9[15]NO3 | 73002.93
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C9 | 77930773 |  | 1170392661 | [13]C4H9[15]NO3 | 77635.09
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C10 | 72386497 |  | 1096417365 | [13]C4H9[15]NO3 | 65542.90
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C11 | 68791215 |  | 1069289020 | [13]C4H9[15]NO3 | 70587.21
Threonine U-13C, U-15N pos | 9.15 | Internal Standard | KD100520_100520_C12 | 68839622 |  | 1047065562 | [13]C4H9[15]NO3 | 28412.06

## Tryptophan neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A1 | 228127508 |  | 3589178369 | C11H12N2O2 | 1286058.06
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A2 | 136054475 |  | 1931426776 | C11H12N2O2 | 13471896.07
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A3 | 79723944 |  | 1164204326 | C11H12N2O2 | 127710.26
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A4 | 35928007 |  | 533435291 | C11H12N2O2 | 4422873.40
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A5 | 29605228 |  | 397267066 | C11H12N2O2 | 21333.32 | Calibration Amount Diff 49.599% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A6 | 11595062 |  | 162925905 | C11H12N2O2 | 8194.52
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A7 | 4373755 |  | 63975084 | C11H12N2O2 | 7132.17 | Calibration Amount Diff -40.377% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A8 | 2380149 |  | 36011850 | C11H12N2O2 | 6484.97 | Calibration Amount Diff -125.038% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A9 | 1304926 |  | 17089672 | C11H12N2O2 | 4962.38 | Calibration Amount Diff -294.632% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A10 | 596478 |  | 7521261 | C11H12N2O2 | 1400.94 | Calibration Amount Diff -660.985% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A11 | 340660 |  | 3808601 | C11H12N2O2 | 768.46 | Calibration Amount Diff -1350.208% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_A12 | 127366 |  | 1459402 | C11H12N2O2 | 612.76 | Calibration Amount Diff -2757.438% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B1 | 251145288 |  | 3877542873 | C11H12N2O2 | 2837919.29
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B2 | 139524765 |  | 2055549701 | C11H12N2O2 | 10961276.80
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B3 | 73057371 |  | 1079469624 | C11H12N2O2 | 301930.56
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B4 | 41256184 |  | 553544789 | C11H12N2O2 | 68999.79 | Calibration Amount Diff 22.011% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B5 | 16382616 |  | 251371915 | C11H12N2O2 | 15804.62
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B6 | 9161247 |  | 127673489 | C11H12N2O2 | 8934.00
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B7 | 4855884 |  | 68271605 | C11H12N2O2 | 5957.34 | Calibration Amount Diff -42.911% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B8 | 2564093 |  | 33962329 | C11H12N2O2 | 5089.76 | Calibration Amount Diff -123.482% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B9 | 1126973 |  | 15939028 | C11H12N2O2 | 2134.40 | Calibration Amount Diff -295.188% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B10 | 538643 |  | 6759281 | C11H12N2O2 | 2520.93 | Calibration Amount Diff -665.364% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B11 | 294585 |  | 3390896 | C11H12N2O2 | 661.39 | Calibration Amount Diff -1355.639% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_B12 | 98935 |  | 1168585 | C11H12N2O2 | 269.22 | Calibration Amount Diff -2766.477% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C11H12N2O2 | N/A
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C1 | 257225171 |  | 3735274550 | C11H12N2O2 | 1161960.98
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C2 | 128171051 |  | 1906668649 | C11H12N2O2 | 3681594.16
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C3 | 72575962 |  | 1015669474 | C11H12N2O2 | 189616.32
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C4 | 34728710 |  | 473313735 | C11H12N2O2 | 503767.54
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C5 | 15746883 |  | 240423449 | C11H12N2O2 | 15194.27
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C6 | 8284505 |  | 117168401 | C11H12N2O2 | 8901.09
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C7 | 4638020 |  | 63145569 | C11H12N2O2 | 4120.04 | Calibration Amount Diff -39.167% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C8 | 2271201 |  | 28858495 | C11H12N2O2 | 3406.35 | Calibration Amount Diff -134.675% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C9 | 1021980 |  | 14476689 | C11H12N2O2 | 1147.01 | Calibration Amount Diff -308.563% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C10 | 397556 |  | 4727471 | C11H12N2O2 | 991.79 | Calibration Amount Diff -692.276% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C11 | 217727 |  | 2301065 | C11H12N2O2 | 1122.39 | Calibration Amount Diff -1387.773% (absolute) > max calibration amount diff 20%
Tryptophan neg | 7.3 | Target Compound | KD100520_100520_C12 | 77250 |  | 968995 | C11H12N2O2 | 333.50 | Calibration Amount Diff -2775.621% (absolute) > max calibration amount diff 20%

## Tryptophan pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A1 | 923094001 |  | 15088950429 | C11H12N2O2 | 133463.00
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A2 | 516618807 |  | 8357356418 | C11H12N2O2 | INF
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A3 | 350332777 |  | 5650455023 | C11H12N2O2 | 373142.55
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A4 | 201609640 |  | 2991905389 | C11H12N2O2 | 2885221.66
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A5 | 124284243 |  | 1933242493 | C11H12N2O2 | 2572502.39 | Calibration Amount Diff 42.62% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A6 | 60153203 |  | 925559490 | C11H12N2O2 | 364859.99
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A7 | 28087314 |  | 404742927 | C11H12N2O2 | 33707.72 | Calibration Amount Diff -20.925% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A8 | 12912888 |  | 189044502 | C11H12N2O2 | 10471.48 | Calibration Amount Diff -75.74% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A9 | 7661668 |  | 104050890 | C11H12N2O2 | 9003.27 | Calibration Amount Diff -169.881% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A10 | 3381190 |  | 43788679 | C11H12N2O2 | 4454.58 | Calibration Amount Diff -398.469% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A11 | 1611360 |  | 21778195 | C11H12N2O2 | 2444.69 | Calibration Amount Diff -811.608% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_A12 | 725901 |  | 9384036 | C11H12N2O2 | 774.29 | Calibration Amount Diff -1658.503% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B1 | 1001353094 |  | 15798275676 | C11H12N2O2 | 1149937.80
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B2 | 518023756 |  | 8750962302 | C11H12N2O2 | 1150683.50
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B3 | 385656409 |  | 5585821726 | C11H12N2O2 | 401723.49
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B4 | 181276170 |  | 3000333299 | C11H12N2O2 | 205213.50
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B5 | 96342440 |  | 1463478577 | C11H12N2O2 | INF
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B6 | 47916669 |  | 755259550 | C11H12N2O2 | 198210.27
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B7 | 25135087 |  | 379630880 | C11H12N2O2 | 69807.08 | Calibration Amount Diff -22.991% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B8 | 14726330 |  | 202941678 | C11H12N2O2 | 8509.11 | Calibration Amount Diff -67.489% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B9 | 7725745 |  | 108121169 | C11H12N2O2 | 11574.06 | Calibration Amount Diff -170.46% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B10 | 3433713 |  | 44396166 | C11H12N2O2 | 4183.26 | Calibration Amount Diff -396.895% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B11 | 1604741 |  | 19307969 | C11H12N2O2 | 2064.21 | Calibration Amount Diff -818.01% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_B12 | 625509 |  | 8026518 | C11H12N2O2 | 952.37 | Calibration Amount Diff -1663.285% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_blank | 2563 |  | 32591 | C11H12N2O2 | 379.38
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C1 | 948544666 |  | 14693635072 | C11H12N2O2 | 968844.42
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C2 | 625455956 |  | 9425201158 | C11H12N2O2 | 236710.60
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C3 | 317275430 |  | 4983454267 | C11H12N2O2 | 245638.86
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C4 | 179936950 |  | 2699540428 | C11H12N2O2 | 4455046.71
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C5 | 98141407 |  | 1507540235 | C11H12N2O2 | 97704.84
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C6 | 46664495 |  | 658476413 | C11H12N2O2 | 82454.52
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C7 | 30103050 |  | 431350022 | C11H12N2O2 | 14692.53
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C8 | 12580175 |  | 182514168 | C11H12N2O2 | 6464.97 | Calibration Amount Diff -79.599% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C9 | 7034907 |  | 98176616 | C11H12N2O2 | 3947.21 | Calibration Amount Diff -181.011% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C10 | 3073197 |  | 39520452 | C11H12N2O2 | 3172.04 | Calibration Amount Diff -408.966% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C11 | 1527728 |  | 19319386 | C11H12N2O2 | 1638.79 | Calibration Amount Diff -828% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD100520_100520_C12 | 663258 |  | 8038764 | C11H12N2O2 | 1620.84 | Calibration Amount Diff -1662.958% (absolute) > max calibration amount diff 20%

## Tryptophan U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A1 | 31758 |  | 243187 | [13]C11H12N2O2 | 5288.25 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 30.672% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A2 | 53646 |  | 560937 | [13]C11H12N2O2 | 1714.20 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -98.863% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A3 | 48601 |  | 532830 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -107.239% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A4 | 31694 |  | 404710 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 157.622% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A5 | 48418 |  | 522882 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -140.34% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A6 | 46994 |  | 485210 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -717.227% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A7 | 27391 |  | 393960 | [13]C11H12N2O2 | 6109.87 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -190.101% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A8 | 31489 |  | 434834 | [13]C11H12N2O2 | 9648.32 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 55.945% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A9 | 39953 |  | 431203 | [13]C11H12N2O2 | 13383.59 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -3688.929% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A10 | 32614 |  | 410452 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 3706.257% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A11 | 37315 |  | 463897 | [13]C11H12N2O2 | 1113.69 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -20424.212% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_A12 | 44379 |  | 440617 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -8931.889% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B1 | 48832 |  | 333375 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B2 | 44142 |  | 427880 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B3 | 40116 |  | 436621 | [13]C11H12N2O2 | 6590.07 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 22.564% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B4 | 39689 |  | 405524 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 151.638% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B5 | 35046 |  | 445915 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -157.59% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B6 | 42209 |  | 403492 | [13]C11H12N2O2 | 6829.70 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -210.016% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B7 | 29894 |  | 410049 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 471.456% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B8 | 36563 |  | 434540 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -2171.832% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B9 | 32333 |  | 376952 | [13]C11H12N2O2 | 1122.83 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 732.285% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B10 | 34454 |  | 386389 | [13]C11H12N2O2 | 1541.01 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 2321.921% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B11 | 30227 |  | 385019 | [13]C11H12N2O2 | 22576.88 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 9028.225% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_B12 | 33305 |  | 365592 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 22532.765% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C11H12N2O2 | N/A
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C1 | 37086 |  | 247828 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 34.757% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C2 | 37715 |  | 465559 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -89.744% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C3 | 37663 |  | 469798 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -104.791% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C4 | 27434 |  | 362543 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 132.344% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C5 | 29780 |  | 355519 | [13]C11H12N2O2 | 483.53 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 209.478% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C6 | 25507 |  | 357000 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 678.848% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C7 | 29260 |  | 360201 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 858.35% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C8 | 40245 |  | 396340 | [13]C11H12N2O2 | 10333.63 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff -1919.794% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C9 | 28332 |  | 380541 | [13]C11H12N2O2 | 10901.10 | Rsquared 0.683 < R2Threshold 0.99
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C10 | 23042 |  | 308211 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 15999.638% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C11 | 32182 |  | 355263 | [13]C11H12N2O2 | INF | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 9225.96% (absolute) > max calibration amount diff 20%
Tryptophan U-13C neg | 7.3 | Target Compound | KD100520_100520_C12 | 24222 |  | 323944 | [13]C11H12N2O2 | 3520.95 | Rsquared 0.683 < R2Threshold 0.99. Calibration Amount Diff 43550.574% (absolute) > max calibration amount diff 20%

## Tryptophan U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A1 | 90397 |  | 395528 | [13]C11H12N2O2 | INF | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 28.152% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A2 | 212042 |  | 2467560 | [13]C11H12N2O2 | INF | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -88.149% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A3 | 204196 |  | 2471108 | [13]C11H12N2O2 | 6874.28 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -29.555% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A4 | 184108 |  | 2810893 | [13]C11H12N2O2 | 862.65 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -61.099% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A5 | 205811 |  | 2695333 | [13]C11H12N2O2 | 9102.92 | Rsquared 0.774 < R2Threshold 0.99
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A6 | 206195 |  | 3062201 | [13]C11H12N2O2 | 1172.40 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -364.389% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A7 | 207866 |  | 2778424 | [13]C11H12N2O2 | 356.32 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -25.363% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A8 | 197360 |  | 2877147 | [13]C11H12N2O2 | 1668.71 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -1566.634% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A9 | 218147 |  | 2716716 | [13]C11H12N2O2 | 14472.23 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 423.187% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A10 | 253797 |  | 2971190 | [13]C11H12N2O2 | 707.85 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -6048.179% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A11 | 190898 |  | 2818285 | [13]C11H12N2O2 | 1349.46 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -4837.214% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_A12 | 201199 |  | 3061622 | [13]C11H12N2O2 | 1305.22 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -22856.088% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B1 | 100977 |  | 343884 | [13]C11H12N2O2 | INF | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 29.782% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B2 | 150457 |  | 1597603 | [13]C11H12N2O2 | INF | Rsquared 0.774 < R2Threshold 0.99
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B3 | 228712 |  | 2845584 | [13]C11H12N2O2 | 2641.27 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -100.844% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B4 | 167448 |  | 2236549 | [13]C11H12N2O2 | 3738.23 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 137.184% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B5 | 198524 |  | 2475486 | [13]C11H12N2O2 | 529.87 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 240.564% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B6 | 185739 |  | 2642796 | [13]C11H12N2O2 | 916.09 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -73.081% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B7 | 221400 |  | 2695382 | [13]C11H12N2O2 | 308.33 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -202.979% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B8 | 201880 |  | 2658871 | [13]C11H12N2O2 | 505.49 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 221.323% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B9 | 167500 |  | 2706940 | [13]C11H12N2O2 | 375.51 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 2202.165% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B10 | 220898 |  | 2870351 | [13]C11H12N2O2 | 282.29 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -3128.476% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B11 | 208638 |  | 2689259 | [13]C11H12N2O2 | 544.09 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -7546.57% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_B12 | 158004 |  | 2023806 | [13]C11H12N2O2 | 199.78 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 68184.61% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C11H12N2O2 | N/A
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C1 | 169698 |  | 1102687 | [13]C11H12N2O2 | INF | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -24.247% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C2 | 182423 |  | 1977064 | [13]C11H12N2O2 | INF | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -35.618% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C3 | 202671 |  | 2187907 | [13]C11H12N2O2 | 15566.52 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -23.013% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C4 | 169019 |  | 2228451 | [13]C11H12N2O2 | 1537.43 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 75.994% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C5 | 176359 |  | 2208592 | [13]C11H12N2O2 | 413.70 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 354.404% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C6 | 188128 |  | 2514626 | [13]C11H12N2O2 | 307.29 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -137.018% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C7 | 229598 |  | 2618853 | [13]C11H12N2O2 | 839.74 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 957.452% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C8 | 194734 |  | 2566297 | [13]C11H12N2O2 | 500.69 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 662.8% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C9 | 231374 |  | 2906850 | [13]C11H12N2O2 | 372.58 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff -1218.532% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C10 | 190909 |  | 2726896 | [13]C11H12N2O2 | 262.09 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 495.36% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C11 | 203162 |  | 2657700 | [13]C11H12N2O2 | 385.35 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 9677.822% (absolute) > max calibration amount diff 20%
Tryptophan U-13C pos | 7.3 | Target Compound | KD100520_100520_C12 | 189513 |  | 2419164 | [13]C11H12N2O2 | 830.24 | Rsquared 0.774 < R2Threshold 0.99. Calibration Amount Diff 18494.903% (absolute) > max calibration amount diff 20%

## Tryptophan U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A1 | 40913565 |  | 597026870 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A2 | 48255730 |  | 649685465 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A3 | 44233365 |  | 611230731 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A4 | 36799739 |  | 537817117 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A5 | 44114652 |  | 598107907 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A6 | 35384819 |  | 520052814 | [13]C11H12[15]N2O2 | 92954.81
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A7 | 28859817 |  | 452659425 | [13]C11H12[15]N2O2 | 79223.08
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A8 | 32234655 |  | 504964904 | [13]C11H12[15]N2O2 | 55557.20
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A9 | 33067915 |  | 471503110 | [13]C11H12[15]N2O2 | 45706.67
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A10 | 33507233 |  | 489031093 | [13]C11H12[15]N2O2 | 40007.51
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A11 | 35674860 |  | 496187402 | [13]C11H12[15]N2O2 | 54748.10
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_A12 | 36463418 |  | 500418771 | [13]C11H12[15]N2O2 | 37344.13
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B1 | 39702905 |  | 582428401 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B2 | 46026100 |  | 643435524 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B3 | 38321717 |  | 576032505 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B4 | 40275411 |  | 537027157 | [13]C11H12[15]N2O2 | 91139.05
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B5 | 33321059 |  | 507872501 | [13]C11H12[15]N2O2 | 1106121.30
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B6 | 32072466 |  | 459850953 | [13]C11H12[15]N2O2 | 24921.31
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B7 | 32841672 |  | 491614455 | [13]C11H12[15]N2O2 | 22880.05
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B8 | 33610597 |  | 471238206 | [13]C11H12[15]N2O2 | 7150145.93
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B9 | 28902522 |  | 441399159 | [13]C11H12[15]N2O2 | 67566.58
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B10 | 30337564 |  | 455217979 | [13]C11H12[15]N2O2 | 94165.49
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B11 | 32214192 |  | 461565635 | [13]C11H12[15]N2O2 | 20276.12
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_B12 | 31757818 |  | 442246121 | [13]C11H12[15]N2O2 | 40479.20
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C11H12[15]N2O2 | N/A
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C1 | 45577468 |  | 630589416 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C2 | 40087247 |  | 549389180 | [13]C11H12[15]N2O2 | 106465.13
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C3 | 39301060 |  | 540254094 | [13]C11H12[15]N2O2 | 110689.38
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C4 | 34872523 |  | 474793794 | [13]C11H12[15]N2O2 | 65877.09
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C5 | 29192777 |  | 445743986 | [13]C11H12[15]N2O2 | 81863.95
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C6 | 32211557 |  | 457651028 | [13]C11H12[15]N2O2 | 48958.68
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C7 | 30696563 |  | 443108980 | [13]C11H12[15]N2O2 | 65373.17
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C8 | 30503191 |  | 433053298 | [13]C11H12[15]N2O2 | 38210.87
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C9 | 27806865 |  | 440483590 | [13]C11H12[15]N2O2 | 789415.23
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C10 | 26663738 |  | 408165609 | [13]C11H12[15]N2O2 | 96731.94
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C11 | 31005654 |  | 426239550 | [13]C11H12[15]N2O2 | 28245.79
Tryptophan U-13C, U-15N neg | 7.3 | Internal Standard | KD100520_100520_C12 | 23955158 |  | 409683365 | [13]C11H12[15]N2O2 | 25038.73

## Tryptophan U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A1 | 152386598 |  | 2428705056 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A2 | 183961767 |  | 2801865886 | [13]C11H12[15]N2O2 | 201100.07
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A3 | 190975088 |  | 3043026600 | [13]C11H12[15]N2O2 | 381219.33
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A4 | 213626795 |  | 3169250883 | [13]C11H12[15]N2O2 | 644656.84
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A5 | 195825413 |  | 3068949698 | [13]C11H12[15]N2O2 | 271993.15
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A6 | 198101079 |  | 3176533508 | [13]C11H12[15]N2O2 | 112419.63
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A7 | 208469121 |  | 3057801026 | [13]C11H12[15]N2O2 | 243366.83
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A8 | 190845368 |  | 2927583219 | [13]C11H12[15]N2O2 | 95743.15
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A9 | 213868881 |  | 3006939264 | [13]C11H12[15]N2O2 | 141083.19
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A10 | 204859116 |  | 3020289516 | [13]C11H12[15]N2O2 | 686238.37
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A11 | 195392835 |  | 2989786513 | [13]C11H12[15]N2O2 | 166393.21
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_A12 | 204525450 |  | 3121866325 | [13]C11H12[15]N2O2 | 6775267.43
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B1 | 153258720 |  | 2243484192 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B2 | 163100670 |  | 2699730453 | [13]C11H12[15]N2O2 | INF
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B3 | 222448373 |  | 3104078689 | [13]C11H12[15]N2O2 | 627516.12
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B4 | 180473849 |  | 3016805783 | [13]C11H12[15]N2O2 | 249239.47
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B5 | 201669963 |  | 3131635422 | [13]C11H12[15]N2O2 | 81660.97
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B6 | 179531891 |  | 2902437594 | [13]C11H12[15]N2O2 | 134495.06
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B7 | 184973946 |  | 2914133439 | [13]C11H12[15]N2O2 | 145179.66
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B8 | 210962771 |  | 2951867729 | [13]C11H12[15]N2O2 | 149580.54
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B9 | 211551917 |  | 3137850513 | [13]C11H12[15]N2O2 | 132074.77
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B10 | 204767162 |  | 3020665610 | [13]C11H12[15]N2O2 | 7015970.80
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B11 | 191049892 |  | 2806668822 | [13]C11H12[15]N2O2 | 114260.84
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_B12 | 190934958 |  | 2811695510 | [13]C11H12[15]N2O2 | 101539.82
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C11H12[15]N2O2 | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C1 | 158570366 |  | 2343181282 | [13]C11H12[15]N2O2 | 261832.55
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C2 | 190495819 |  | 2721612406 | [13]C11H12[15]N2O2 | 79816.71
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C3 | 176762423 |  | 2726531854 | [13]C11H12[15]N2O2 | 385102.46
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C4 | 186760123 |  | 2834154444 | [13]C11H12[15]N2O2 | 93169.49
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C5 | 188249890 |  | 2949936704 | [13]C11H12[15]N2O2 | 81146.59
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C6 | 193264787 |  | 2726537833 | [13]C11H12[15]N2O2 | 127069.66
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C7 | 223401553 |  | 3199805077 | [13]C11H12[15]N2O2 | 244893.28
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C8 | 190308282 |  | 2914632718 | [13]C11H12[15]N2O2 | 209211.04
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C9 | 208370352 |  | 3088638702 | [13]C11H12[15]N2O2 | 178509.06
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C10 | 202194234 |  | 3001035261 | [13]C11H12[15]N2O2 | 421363.78
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C11 | 205234217 |  | 3092370659 | [13]C11H12[15]N2O2 | 139113.57
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD100520_100520_C12 | 187753447 |  | 2805832577 | [13]C11H12[15]N2O2 | 245171.52

## Tyrosine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A1 | 97357558 |  | 1640408843 | C9H11NO3 | 107595.14
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A2 | 60918250 |  | 896855591 | C9H11NO3 | 53058.15
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A3 | 40709864 |  | 571640157 | C9H11NO3 | 100541.45
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A4 | 23102331 |  | 310757595 | C9H11NO3 | 26974.59
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A5 | 16123882 |  | 224028415 | C9H11NO3 | 13908.79 | Calibration Amount Diff 36.668% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A6 | 9101798 |  | 117035047 | C9H11NO3 | 9847.28
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A7 | 4624974 |  | 56657438 | C9H11NO3 | 8931.02
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A8 | 2200164 |  | 28013979 | C9H11NO3 | 4924.22 | Calibration Amount Diff -49.866% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A9 | 1285869 |  | 15243300 | C9H11NO3 | 3545.99 | Calibration Amount Diff -105.084% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A10 | 487108 |  | 6061437 | C9H11NO3 | 1442.76 | Calibration Amount Diff -267.324% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A11 | 290859 |  | 3499105 | C9H11NO3 | 933.55 | Calibration Amount Diff -517.352% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_A12 | 111410 |  | 1281945 | C9H11NO3 | 360.11 | Calibration Amount Diff -1068.41% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B1 | 107815645 |  | 1766055898 | C9H11NO3 | 61901.22
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B2 | 62152262 |  | 973977532 | C9H11NO3 | 61686.67
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B3 | 38230986 |  | 555066165 | C9H11NO3 | 47235.32
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B4 | 24196787 |  | 337568172 | C9H11NO3 | 44755.80
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B5 | 11836520 |  | 169036068 | C9H11NO3 | 16699.64
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B6 | 8044716 |  | 98846369 | C9H11NO3 | 9670.80
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B7 | 4207302 |  | 55914659 | C9H11NO3 | 11173.20
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B8 | 2415243 |  | 29328892 | C9H11NO3 | 6136.29 | Calibration Amount Diff -42.489% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B9 | 1287828 |  | 15739445 | C9H11NO3 | 3933.50 | Calibration Amount Diff -107.085% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B10 | 513714 |  | 6059259 | C9H11NO3 | 1096.72 | Calibration Amount Diff -265.138% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B11 | 220793 |  | 2901037 | C9H11NO3 | 612.98 | Calibration Amount Diff -526.494% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_B12 | 126415 |  | 1465018 | C9H11NO3 | 582.54 | Calibration Amount Diff -1058.259% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C9H11NO3 | N/A
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C1 | 102852433 |  | 1697716203 | C9H11NO3 | 97874.69
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C2 | 69031741 |  | 1015031847 | C9H11NO3 | 46909.70
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C3 | 36314637 |  | 531549139 | C9H11NO3 | 40706.24
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C4 | 22441991 |  | 303864528 | C9H11NO3 | 17197.63
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C5 | 11950221 |  | 165310540 | C9H11NO3 | 9768.44
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C6 | 6814344 |  | 85607234 | C9H11NO3 | 8284.42
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C7 | 4062451 |  | 53721441 | C9H11NO3 | 6083.54
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C8 | 2053144 |  | 26814521 | C9H11NO3 | 3927.37 | Calibration Amount Diff -52.782% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C9 | 950361 |  | 12727782 | C9H11NO3 | 1374.86 | Calibration Amount Diff -123.27% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C10 | 440507 |  | 4883930 | C9H11NO3 | 1191.35 | Calibration Amount Diff -281.257% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C11 | 213235 |  | 2786827 | C9H11NO3 | 770.89 | Calibration Amount Diff -539.498% (absolute) > max calibration amount diff 20%
Tyrosine neg | 8.28 | Target Compound | KD100520_100520_C12 | 94140 |  | 1018863 | C9H11NO3 | 275.00 | Calibration Amount Diff -1076.677% (absolute) > max calibration amount diff 20%

## Tyrosine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A1 | 350129644 |  | 6145127880 | C9H11NO3 | 33545.98
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A2 | 247680893 |  | 3782429639 | C9H11NO3 | 47139.91
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A3 | 171534483 |  | 2469850842 | C9H11NO3 | 22016.95
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A4 | 92963612 |  | 1413071098 | C9H11NO3 | 83918.19
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A5 | 75147950 |  | 1033275711 | C9H11NO3 | 50484.42 | Calibration Amount Diff 33.247% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A6 | 29290223 |  | 462778590 | C9H11NO3 | 22507.13
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A7 | 17095809 |  | 227064675 | C9H11NO3 | 7555.13
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A8 | 8655818 |  | 111090330 | C9H11NO3 | 4500.90 | Calibration Amount Diff -31.676% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A9 | 4205361 |  | 54651176 | C9H11NO3 | 2767.58 | Calibration Amount Diff -67.815% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A10 | 1983454 |  | 22886215 | C9H11NO3 | 2010.31 | Calibration Amount Diff -164.731% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A11 | 844398 |  | 12025728 | C9H11NO3 | 631.27 | Calibration Amount Diff -317.377% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_A12 | 383437 |  | 4967627 | C9H11NO3 | 807.78 | Calibration Amount Diff -640.895% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B1 | 401056554 |  | 6801066346 | C9H11NO3 | 47565.16
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B2 | 266821760 |  | 4120633506 | C9H11NO3 | INF
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B3 | 171743603 |  | 2618770092 | C9H11NO3 | 39503.28
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B4 | 106680081 |  | 1559188894 | C9H11NO3 | 57296.61
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B5 | 52519909 |  | 791953099 | C9H11NO3 | 47024.16
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B6 | 31315867 |  | 468052507 | C9H11NO3 | 21523.94
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B7 | 17128672 |  | 232015999 | C9H11NO3 | 5062.31
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B8 | 8823076 |  | 121721766 | C9H11NO3 | 4626.64 | Calibration Amount Diff -23.592% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B9 | 4515364 |  | 56392544 | C9H11NO3 | 5625.74 | Calibration Amount Diff -66.877% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B10 | 1519813 |  | 20311545 | C9H11NO3 | 1196.82 | Calibration Amount Diff -169.881% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B11 | 677648 |  | 9782108 | C9H11NO3 | 685.09 | Calibration Amount Diff -328.394% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_B12 | 288844 |  | 4158560 | C9H11NO3 | 284.54 | Calibration Amount Diff -647.479% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C9H11NO3 | N/A
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C1 | 360665427 |  | 6431653009 | C9H11NO3 | 85714.72
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C2 | 305050553 |  | 4480276403 | C9H11NO3 | 2641800.20
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C3 | 183103335 |  | 2592927469 | C9H11NO3 | 121013.95
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C4 | 104348631 |  | 1407495141 | C9H11NO3 | 90181.64
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C5 | 57475787 |  | 799980783 | C9H11NO3 | 42371.17
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C6 | 24066994 |  | 385613504 | C9H11NO3 | 11469.26
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C7 | 15687606 |  | 225832333 | C9H11NO3 | 12616.06
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C8 | 7520753 |  | 108104074 | C9H11NO3 | 5072.95 | Calibration Amount Diff -34.724% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C9 | 4077986 |  | 56767353 | C9H11NO3 | 2448.18 | Calibration Amount Diff -74.055% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C10 | 1364955 |  | 18650573 | C9H11NO3 | 1033.91 | Calibration Amount Diff -181.949% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C11 | 785692 |  | 10229342 | C9H11NO3 | 684.86 | Calibration Amount Diff -335.146% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD100520_100520_C12 | 339058 |  | 4370126 | C9H11NO3 | 386.30 | Calibration Amount Diff -646.525% (absolute) > max calibration amount diff 20%

## Tyrosine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A1 | 65361 |  | 686056 | [13]C9H11NO3 | 419.25 | Rsquared 0.727 < R2Threshold 0.99
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A2 | 102133 |  | 928853 | [13]C9H11NO3 | 685.98 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 36.638% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A3 | 82116 |  | 1184708 | [13]C9H11NO3 | 365.74 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -47.52% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A4 | 109275 |  | 1274358 | [13]C9H11NO3 | 1681.01 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 85.152% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A5 | 107571 |  | 1368608 | [13]C9H11NO3 | 416.42 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 347.292% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A6 | 144288 |  | 1751754 | [13]C9H11NO3 | 1013.05 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -227.428% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A7 | 147470 |  | 1677062 | [13]C9H11NO3 | 1625.13 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 1295.88% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A8 | 154189 |  | 1834596 | [13]C9H11NO3 | 557.70 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 1039.647% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A9 | 197266 |  | 2038413 | [13]C9H11NO3 | 1806.49 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -6139.96% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A10 | 169081 |  | 2239993 | [13]C9H11NO3 | 611.57 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -19194.064% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A11 | 165500 |  | 2116201 | [13]C9H11NO3 | 859.38 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -14370.327% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_A12 | 177498 |  | 1871387 | [13]C9H11NO3 | 1132.80 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 19211.026% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B1 | 83673 |  | 763338 | [13]C9H11NO3 | 948.94 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -28.304% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B2 | 77640 |  | 922913 | [13]C9H11NO3 | 565.14 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 38.583% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B3 | 98575 |  | 1196147 | [13]C9H11NO3 | 511.24 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -46.556% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B4 | 100429 |  | 1352008 | [13]C9H11NO3 | 356.56 | Rsquared 0.727 < R2Threshold 0.99
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B5 | 103604 |  | 1510698 | [13]C9H11NO3 | 733.82 | Rsquared 0.727 < R2Threshold 0.99
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B6 | 148841 |  | 1697169 | [13]C9H11NO3 | 926.13 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -413.221% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B7 | 137720 |  | 1813795 | [13]C9H11NO3 | 775.88 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 335.631% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B8 | 185643 |  | 2030209 | [13]C9H11NO3 | 4232.37 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -3528.598% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B9 | 180753 |  | 2064226 | [13]C9H11NO3 | 811.16 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -3293.706% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B10 | 164460 |  | 2122171 | [13]C9H11NO3 | 1898.10 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -13990.91% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B11 | 128439 |  | 1682173 | [13]C9H11NO3 | 566.94 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 24393.643% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_B12 | 171327 |  | 2082817 | [13]C9H11NO3 | 809.91 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -35377.684% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C9H11NO3 | N/A
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C1 | 71024 |  | 676604 | [13]C9H11NO3 | 494.24 | Rsquared 0.727 < R2Threshold 0.99
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C2 | 82634 |  | 982167 | [13]C9H11NO3 | 366.01 | Rsquared 0.727 < R2Threshold 0.99
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C3 | 77565 |  | 1060435 | [13]C9H11NO3 | 277.95 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 42.774% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C4 | 115541 |  | 1244564 | [13]C9H11NO3 | 1249.25 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 57.861% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C5 | 96404 |  | 1175876 | [13]C9H11NO3 | 2078.83 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 574.314% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C6 | 118100 |  | 1424938 | [13]C9H11NO3 | 2233.68 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 530.708% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C7 | 146106 |  | 1732657 | [13]C9H11NO3 | 465.27 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -130.599% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C8 | 154995 |  | 1886206 | [13]C9H11NO3 | 5009.58 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -690.159% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C9 | 122873 |  | 1664676 | [13]C9H11NO3 | 581.37 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 7177.072% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C10 | 155537 |  | 2018150 | [13]C9H11NO3 | 1217.24 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff -11459.735% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C11 | 159446 |  | 1992067 | [13]C9H11NO3 | 758.82 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 4553.733% (absolute) > max calibration amount diff 20%
Tyrosine U-13C neg | 8.28 | Target Compound | KD100520_100520_C12 | 140583 |  | 1623869 | [13]C9H11NO3 | 2936.85 | Rsquared 0.727 < R2Threshold 0.99. Calibration Amount Diff 29141.884% (absolute) > max calibration amount diff 20%

## Tyrosine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A1 | 179657 |  | 1859843 | [13]C9H11NO3 | 51.40 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 36.859% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A2 | 300608 |  | 4089210 | [13]C9H11NO3 | 149.17 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -50.305% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A3 | 303282 |  | 3035974 | [13]C9H11NO3 | 92.57 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 368.93% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A4 | 344382 |  | 5113643 | [13]C9H11NO3 | 73.12 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 54.978% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A5 | 479070 |  | 6219013 | [13]C9H11NO3 | 128.01 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -128.326% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A6 | 415391 |  | 5706016 | [13]C9H11NO3 | 120.01 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 642.533% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A7 | 508304 |  | 6288133 | [13]C9H11NO3 | 250.85 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 1518.207% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A8 | 661616 |  | 7430545 | [13]C9H11NO3 | 216.96 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -3529.189% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A9 | 500791 |  | 6073809 | [13]C9H11NO3 | 172.78 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 6503.364% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A10 | 601136 |  | 6828434 | [13]C9H11NO3 | 213.62 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -2620.837% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A11 | 528054 |  | 7349522 | [13]C9H11NO3 | 235.69 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -22036.368% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_A12 | 644269 |  | 7851292 | [13]C9H11NO3 | 251.51 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -114543.939% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B1 | 204894 |  | 2002442 | [13]C9H11NO3 | 364.87 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 28.396% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B2 | 382793 |  | 4744966 | [13]C9H11NO3 | 151.07 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -155.928% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B3 | 354317 |  | 4409133 | [13]C9H11NO3 | 80.65 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 86.004% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B4 | 360445 |  | 4984672 | [13]C9H11NO3 | 53.65 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 263.825% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B5 | 412500 |  | 5477968 | [13]C9H11NO3 | 118.35 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 659.862% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B6 | 561382 |  | 6775217 | [13]C9H11NO3 | 218.01 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 306.964% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B7 | 629078 |  | 7328965 | [13]C9H11NO3 | 284.55 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -997.235% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B8 | 562836 |  | 7042454 | [13]C9H11NO3 | 184.61 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -390.475% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B9 | 655596 |  | 7707672 | [13]C9H11NO3 | 358.73 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -11917.584% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B10 | 512217 |  | 6403424 | [13]C9H11NO3 | 376.46 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -1967.976% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B11 | 421797 |  | 5910632 | [13]C9H11NO3 | 186.07 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 23884.019% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_B12 | 438559 |  | 6108823 | [13]C9H11NO3 | 211.17 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 1418.209% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C9H11NO3 | N/A
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C1 | 258446 |  | 2660739 | [13]C9H11NO3 | 45.61 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -27.967% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C2 | 366319 |  | 4197452 | [13]C9H11NO3 | 230.36 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -68.489% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C3 | 443587 |  | 5008300 | [13]C9H11NO3 | 66.44 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -100.751% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C4 | 439745 |  | 5180993 | [13]C9H11NO3 | 86.65 | Rsquared 0.464 < R2Threshold 0.99
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C5 | 383426 |  | 5428610 | [13]C9H11NO3 | 77.89 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 316.086% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C6 | 447236 |  | 5817061 | [13]C9H11NO3 | 144.18 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 393.18% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C7 | 524578 |  | 6920313 | [13]C9H11NO3 | 217.38 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -597.785% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C8 | 544413 |  | 6630632 | [13]C9H11NO3 | 226.42 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 1386.96% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C9 | 566661 |  | 7517501 | [13]C9H11NO3 | 340.50 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -2976.621% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C10 | 501629 |  | 6189361 | [13]C9H11NO3 | 201.23 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff 14379.049% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C11 | 605186 |  | 7632291 | [13]C9H11NO3 | 225.48 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -17011.18% (absolute) > max calibration amount diff 20%
Tyrosine U-13C pos | 8.28 | Target Compound | KD100520_100520_C12 | 503056 |  | 6707675 | [13]C9H11NO3 | 256.92 | Rsquared 0.464 < R2Threshold 0.99. Calibration Amount Diff -37624.487% (absolute) > max calibration amount diff 20%

## Tyrosine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A1 | 11790974 |  | 189357251 | [13]C9H11[15]NO3 | 36827.26
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A2 | 15733827 |  | 225018847 | [13]C9H11[15]NO3 | 24398.52
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A3 | 16425762 |  | 230980846 | [13]C9H11[15]NO3 | 17648.17
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A4 | 19199302 |  | 257590543 | [13]C9H11[15]NO3 | 42657.47
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A5 | 20485081 |  | 281609156 | [13]C9H11[15]NO3 | 25511.77
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A6 | 25307546 |  | 322060521 | [13]C9H11[15]NO3 | 21123.31
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A7 | 27123086 |  | 337343246 | [13]C9H11[15]NO3 | 16414.42
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A8 | 26442682 |  | 352433323 | [13]C9H11[15]NO3 | 21308.89
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A9 | 28226236 |  | 351761526 | [13]C9H11[15]NO3 | 13075.82
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A10 | 26821378 |  | 370634745 | [13]C9H11[15]NO3 | 16254.82
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A11 | 28562383 |  | 376575519 | [13]C9H11[15]NO3 | 10545.54
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_A12 | 28468541 |  | 360154486 | [13]C9H11[15]NO3 | 17163.83
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B1 | 11709679 |  | 187706864 | [13]C9H11[15]NO3 | 44172.52
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B2 | 14788388 |  | 224539443 | [13]C9H11[15]NO3 | 27653.87
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B3 | 16557971 |  | 233409986 | [13]C9H11[15]NO3 | 21352.80
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B4 | 18965965 |  | 264606137 | [13]C9H11[15]NO3 | 27142.62
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B5 | 19573133 |  | 287085890 | [13]C9H11[15]NO3 | 15731.55
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B6 | 24737987 |  | 306098990 | [13]C9H11[15]NO3 | 30128.56
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B7 | 25838303 |  | 345911488 | [13]C9H11[15]NO3 | 28847.09
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B8 | 27752322 |  | 346880290 | [13]C9H11[15]NO3 | 11056.64
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B9 | 28615914 |  | 369057720 | [13]C9H11[15]NO3 | 22556.06
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B10 | 27367690 |  | 362190369 | [13]C9H11[15]NO3 | 19771.19
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B11 | 25277257 |  | 340975869 | [13]C9H11[15]NO3 | 21948.16
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_B12 | 28008579 |  | 366743577 | [13]C9H11[15]NO3 | 24892.38
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C9H11[15]NO3 | N/A
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C1 | 12643839 |  | 201029189 | [13]C9H11[15]NO3 | 22274.35
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C2 | 15354880 |  | 220775568 | [13]C9H11[15]NO3 | 18366.41
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C3 | 15985406 |  | 224686887 | [13]C9H11[15]NO3 | 20733.52
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C4 | 18823488 |  | 248460839 | [13]C9H11[15]NO3 | 23018.75
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C5 | 17382583 |  | 255480168 | [13]C9H11[15]NO3 | 12642.32
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C6 | 21793142 |  | 284446388 | [13]C9H11[15]NO3 | 21201.31
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C7 | 24450543 |  | 322315402 | [13]C9H11[15]NO3 | 19012.54
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C8 | 26385811 |  | 346054601 | [13]C9H11[15]NO3 | 13416.90
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C9 | 24506280 |  | 343128997 | [13]C9H11[15]NO3 | 17359.27
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C10 | 26334710 |  | 349792203 | [13]C9H11[15]NO3 | 11983.91
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C11 | 29142647 |  | 376956678 | [13]C9H11[15]NO3 | 12054.23
Tyrosine U-13C, U-15N neg | 8.28 | Internal Standard | KD100520_100520_C12 | 24716339 |  | 317896433 | [13]C9H11[15]NO3 | 9939.06

## Tyrosine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A1 | 43372681 |  | 727041980 | [13]C9H11[15]NO3 | INF
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A2 | 71476042 |  | 1002332406 | [13]C9H11[15]NO3 | INF
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A3 | 77168296 |  | 1074792915 | [13]C9H11[15]NO3 | INF
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A4 | 81640070 |  | 1230995400 | [13]C9H11[15]NO3 | 90847.83
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A5 | 102931454 |  | 1400275009 | [13]C9H11[15]NO3 | 73013.10
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A6 | 88790717 |  | 1391026716 | [13]C9H11[15]NO3 | 57027.53
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A7 | 110630645 |  | 1543578859 | [13]C9H11[15]NO3 | 63826.76
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A8 | 115301907 |  | 1554137510 | [13]C9H11[15]NO3 | 55175.20
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A9 | 102150954 |  | 1493490979 | [13]C9H11[15]NO3 | 76453.06
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A10 | 113177450 |  | 1522795960 | [13]C9H11[15]NO3 | 63893.88
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A11 | 104765097 |  | 1561048131 | [13]C9H11[15]NO3 | 58684.64
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_A12 | 107812042 |  | 1516895482 | [13]C9H11[15]NO3 | 76150.55
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B1 | 46985572 |  | 749122613 | [13]C9H11[15]NO3 | INF
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B2 | 65214077 |  | 989131248 | [13]C9H11[15]NO3 | INF
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B3 | 80151362 |  | 1164728520 | [13]C9H11[15]NO3 | INF
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B4 | 91738233 |  | 1311949251 | [13]C9H11[15]NO3 | 57756.58
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B5 | 96533850 |  | 1452235887 | [13]C9H11[15]NO3 | 76704.12
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B6 | 108019974 |  | 1596245806 | [13]C9H11[15]NO3 | 72185.03
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B7 | 115145849 |  | 1590595914 | [13]C9H11[15]NO3 | 64647.24
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B8 | 109724016 |  | 1583265826 | [13]C9H11[15]NO3 | 108621.72
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B9 | 107558998 |  | 1528004106 | [13]C9H11[15]NO3 | 110192.51
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B10 | 93745411 |  | 1433555353 | [13]C9H11[15]NO3 | 45681.95
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B11 | 96104217 |  | 1441883824 | [13]C9H11[15]NO3 | 86288.30
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_B12 | 89928503 |  | 1386111390 | [13]C9H11[15]NO3 | 93262.26
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C9H11[15]NO3 | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C1 | 46224400 |  | 773757248 | [13]C9H11[15]NO3 | 138384.01
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C2 | 72501370 |  | 998631080 | [13]C9H11[15]NO3 | INF
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C3 | 84931922 |  | 1133155323 | [13]C9H11[15]NO3 | INF
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C4 | 89780485 |  | 1218775853 | [13]C9H11[15]NO3 | 103383.87
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C5 | 99179959 |  | 1335791612 | [13]C9H11[15]NO3 | 53755.86
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C6 | 88285942 |  | 1382424182 | [13]C9H11[15]NO3 | 53091.24
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C7 | 102621616 |  | 1530062959 | [13]C9H11[15]NO3 | 70311.37
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C8 | 104585967 |  | 1556693646 | [13]C9H11[15]NO3 | 56421.12
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C9 | 112413372 |  | 1645016210 | [13]C9H11[15]NO3 | 74591.77
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C10 | 103665843 |  | 1534716363 | [13]C9H11[15]NO3 | 82332.01
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C11 | 115174310 |  | 1644378833 | [13]C9H11[15]NO3 | 64135.95
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD100520_100520_C12 | 92813253 |  | 1437568944 | [13]C9H11[15]NO3 | 45083.21

## Tyr_cali
Compound | Filename | Area | Area_IS | Response ratio | Conc.
Tyrosine pos | KD100520_100520_A1 | 6145127880 | 727041980 | 8.4522 | 500
Tyrosine pos | KD100520_100520_A2 | 3782429639 | 1002332406 | 3.7736 | 250
Tyrosine pos | KD100520_100520_A3 | 2469850842 | 1074792915 | 2.298 | 125
Tyrosine pos | KD100520_100520_A4 | 1413071098 | 1230995400 | 1.1479 | 62.5
Tyrosine pos | KD100520_100520_A5 | 1033275711 | 1400275009 | 0.73791 | 31.25
Tyrosine pos | KD100520_100520_A6 | 462778590 | 1391026716 | 0.33269 | 15.625
Tyrosine pos | KD100520_100520_A7 | 227064675 | 1543578859 | 0.1471 | 7.8125
Tyrosine pos | KD100520_100520_A8 | 111090330 | 1554137510 | 0.07148 | 3.9062
Tyrosine pos | KD100520_100520_A9 | 54651176 | 1493490979 | 0.036593 | 1.9531
Tyrosine pos | KD100520_100520_A10 | 22886215 | 1522795960 | 0.015029 | 0.97656
Tyrosine pos | KD100520_100520_A11 | 12025728 | 1561048131 | 0.0077036 | 0.48828
Tyrosine pos | KD100520_100520_A12 | 4967627 | 1516895482 | 0.0032749 | 0.24414
Tyrosine pos | KD100520_100520_B1 | 6801066346 | 749122613 | 9.0787 | 500
Tyrosine pos | KD100520_100520_B2 | 4120633506 | 989131248 | 4.1659 | 250
Tyrosine pos | KD100520_100520_B3 | 2618770092 | 1164728520 | 2.2484 | 125
Tyrosine pos | KD100520_100520_B4 | 1559188894 | 1311949251 | 1.1885 | 62.5
Tyrosine pos | KD100520_100520_B5 | 791953099 | 1452235887 | 0.54533 | 31.25
Tyrosine pos | KD100520_100520_B6 | 468052507 | 1596245806 | 0.29322 | 15.625
Tyrosine pos | KD100520_100520_B7 | 232015999 | 1590595914 | 0.14587 | 7.8125
Tyrosine pos | KD100520_100520_B8 | 121721766 | 1583265826 | 0.07688 | 3.9062
Tyrosine pos | KD100520_100520_B9 | 56392544 | 1528004106 | 0.036906 | 1.9531
Tyrosine pos | KD100520_100520_B10 | 20311545 | 1433555353 | 0.014169 | 0.97656
Tyrosine pos | KD100520_100520_B11 | 9782108 | 1441883824 | 0.0067843 | 0.48828
Tyrosine pos | KD100520_100520_B12 | 4158560 | 1386111390 | 0.0030002 | 0.24414
Tyrosine pos | KD100520_100520_C1 | 6431653009 | 773757248 | 8.3122 | 500
Tyrosine pos | KD100520_100520_C2 | 4480276403 | 998631080 | 4.4864 | 250
Tyrosine pos | KD100520_100520_C3 | 2592927469 | 1133155323 | 2.2882 | 125
Tyrosine pos | KD100520_100520_C4 | 1407495141 | 1218775853 | 1.1548 | 62.5
Tyrosine pos | KD100520_100520_C5 | 799980783 | 1335791612 | 0.59888 | 31.25
Tyrosine pos | KD100520_100520_C6 | 385613504 | 1382424182 | 0.27894 | 15.625
Tyrosine pos | KD100520_100520_C7 | 225832333 | 1530062959 | 0.1476 | 7.8125
Tyrosine pos | KD100520_100520_C8 | 108104074 | 1556693646 | 0.069445 | 3.9062
Tyrosine pos | KD100520_100520_C9 | 56767353 | 1645016210 | 0.034509 | 1.9531
Tyrosine pos | KD100520_100520_C10 | 18650573 | 1534716363 | 0.012152 | 0.97656
Tyrosine pos | KD100520_100520_C11 | 10229342 | 1644378833 | 0.0062208 | 0.48828
Tyrosine pos | KD100520_100520_C12 | 4370126 | 1437568944 | 0.0030399 | 0.24414

## Valine neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Valine neg | 7.85 | Target Compound | KD100520_100520_A1 | 80343419 |  | 1171336078 | C5H11NO2 | 52412.95
Valine neg | 7.85 | Target Compound | KD100520_100520_A2 | 41057890 |  | 601929417 | C5H11NO2 | 25600.17
Valine neg | 7.85 | Target Compound | KD100520_100520_A3 | 24006241 |  | 361558120 | C5H11NO2 | 26691.93
Valine neg | 7.85 | Target Compound | KD100520_100520_A4 | 12816263 |  | 184929167 | C5H11NO2 | 28876.11
Valine neg | 7.85 | Target Compound | KD100520_100520_A5 | 8590126 |  | 118748897 | C5H11NO2 | 9129.76 | Calibration Amount Diff 40.446% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_A6 | 4170824 |  | 54647977 | C5H11NO2 | 7590.27
Valine neg | 7.85 | Target Compound | KD100520_100520_A7 | 1805310 |  | 24553174 | C5H11NO2 | 1936.58 | Calibration Amount Diff -27.03% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_A8 | 944361 |  | 12012939 | C5H11NO2 | 1334.64 | Calibration Amount Diff -78.453% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_A9 | 512807 |  | 6383856 | C5H11NO2 | 556.53 | Calibration Amount Diff -172.729% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_A10 | 221437 |  | 2686507 | C5H11NO2 | 402.50 | Calibration Amount Diff -399.125% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_A11 | 120478 |  | 1465195 | C5H11NO2 | 245.69 | Calibration Amount Diff -802.235% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_A12 | 42081 |  | 512903 | C5H11NO2 | 94.52 | Calibration Amount Diff -1660.812% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_B1 | 89339706 |  | 1351114530 | C5H11NO2 | 63888.43
Valine neg | 7.85 | Target Compound | KD100520_100520_B2 | 42584040 |  | 659487128 | C5H11NO2 | 37714.45
Valine neg | 7.85 | Target Compound | KD100520_100520_B3 | 25405789 |  | 364355945 | C5H11NO2 | 45066.18
Valine neg | 7.85 | Target Compound | KD100520_100520_B4 | 13929963 |  | 194789167 | C5H11NO2 | 12772.20
Valine neg | 7.85 | Target Compound | KD100520_100520_B5 | 6690593 |  | 89710318 | C5H11NO2 | 7848.12
Valine neg | 7.85 | Target Compound | KD100520_100520_B6 | 3452857 |  | 46381175 | C5H11NO2 | 3758.50
Valine neg | 7.85 | Target Compound | KD100520_100520_B7 | 1755704 |  | 24367118 | C5H11NO2 | 2026.17 | Calibration Amount Diff -29.95% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_B8 | 885238 |  | 11730346 | C5H11NO2 | 1339.88 | Calibration Amount Diff -75.006% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_B9 | 489051 |  | 6094573 | C5H11NO2 | 635.04 | Calibration Amount Diff -173.224% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_B10 | 203726 |  | 2646794 | C5H11NO2 | 296.72 | Calibration Amount Diff -400.786% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_B11 | 90213 |  | 1076796 | C5H11NO2 | 247.38 | Calibration Amount Diff -825.426% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_B12 | 29744 |  | 431682 | C5H11NO2 | 76.75 | Calibration Amount Diff -1667.551% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H11NO2 | N/A
Valine neg | 7.85 | Target Compound | KD100520_100520_C1 | 82711039 |  | 1207513745 | C5H11NO2 | 1130641.64
Valine neg | 7.85 | Target Compound | KD100520_100520_C2 | 42480522 |  | 658735334 | C5H11NO2 | 36473.00
Valine neg | 7.85 | Target Compound | KD100520_100520_C3 | 22478936 |  | 341749099 | C5H11NO2 | 22575.70
Valine neg | 7.85 | Target Compound | KD100520_100520_C4 | 11545912 |  | 169256421 | C5H11NO2 | 18657.11
Valine neg | 7.85 | Target Compound | KD100520_100520_C5 | 6503281 |  | 89928569 | C5H11NO2 | 5834.98
Valine neg | 7.85 | Target Compound | KD100520_100520_C6 | 3033320 |  | 41751035 | C5H11NO2 | 3823.58
Valine neg | 7.85 | Target Compound | KD100520_100520_C7 | 1932551 |  | 24738429 | C5H11NO2 | 2406.52 | Calibration Amount Diff -25.233% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_C8 | 817529 |  | 11144468 | C5H11NO2 | 1181.49 | Calibration Amount Diff -84.985% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_C9 | 466996 |  | 5894213 | C5H11NO2 | 737.15 | Calibration Amount Diff -185.882% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_C10 | 179354 |  | 2117456 | C5H11NO2 | 269.31 | Calibration Amount Diff -422.274% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_C11 | 92931 |  | 1093929 | C5H11NO2 | 189.76 | Calibration Amount Diff -829.335% (absolute) > max calibration amount diff 20%
Valine neg | 7.85 | Target Compound | KD100520_100520_C12 | 38278 |  | 499323 | C5H11NO2 | 165.47 | Calibration Amount Diff -1653.717% (absolute) > max calibration amount diff 20%

## Valine pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Valine pos | 7.85 | Target Compound | KD100520_100520_A1 | 711529953 |  | 10542012293 | C5H11NO2 | 7361.33
Valine pos | 7.85 | Target Compound | KD100520_100520_A2 | 437362172 |  | 6235328590 | C5H11NO2 | 2939.81
Valine pos | 7.85 | Target Compound | KD100520_100520_A3 | 326421198 |  | 4360466498 | C5H11NO2 | 2620.89
Valine pos | 7.85 | Target Compound | KD100520_100520_A4 | 173234784 |  | 2437987028 | C5H11NO2 | 3793.98
Valine pos | 7.85 | Target Compound | KD100520_100520_A5 | 128138479 |  | 1729311640 | C5H11NO2 | 1602.92 | Calibration Amount Diff 39.578% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_A6 | 65045323 |  | 868507231 | C5H11NO2 | 749.85
Valine pos | 7.85 | Target Compound | KD100520_100520_A7 | 30503004 |  | 414275220 | C5H11NO2 | 431.29 | Calibration Amount Diff -23.203% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_A8 | 15164683 |  | 208661362 | C5H11NO2 | 404.86 | Calibration Amount Diff -87.811% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_A9 | 8815528 |  | 122900858 | C5H11NO2 | 182.06 | Calibration Amount Diff -198.022% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_A10 | 4545182 |  | 67614515 | C5H11NO2 | 64.76 | Calibration Amount Diff -462.803% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_A11 | 2898142 |  | 51550938 | C5H11NO2 | 67.46 | Calibration Amount Diff -924.649% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_A12 | 1631034 |  | 28371879 | C5H11NO2 | 50.33 | Calibration Amount Diff -2027.381% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_B1 | 795196497 |  | 11011246304 | C5H11NO2 | 7878.58
Valine pos | 7.85 | Target Compound | KD100520_100520_B2 | 461331833 |  | 6559113429 | C5H11NO2 | 4921.28
Valine pos | 7.85 | Target Compound | KD100520_100520_B3 | 300310744 |  | 4116029938 | C5H11NO2 | 5911.21
Valine pos | 7.85 | Target Compound | KD100520_100520_B4 | 171694001 |  | 2473092350 | C5H11NO2 | 1201.73
Valine pos | 7.85 | Target Compound | KD100520_100520_B5 | 99574614 |  | 1304802175 | C5H11NO2 | 376.71
Valine pos | 7.85 | Target Compound | KD100520_100520_B6 | 54411736 |  | 748244144 | C5H11NO2 | 442.30
Valine pos | 7.85 | Target Compound | KD100520_100520_B7 | 30032418 |  | 405367695 | C5H11NO2 | 619.29 | Calibration Amount Diff -23.338% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_B8 | 15564671 |  | 221947441 | C5H11NO2 | 228.39 | Calibration Amount Diff -73.056% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_B9 | 7906944 |  | 120914881 | C5H11NO2 | 127.90 | Calibration Amount Diff -194.021% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_B10 | 4631588 |  | 94766479 | C5H11NO2 | 97.27 | Calibration Amount Diff -378.199% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_B11 | 2831347 |  | 44187811 | C5H11NO2 | 81.46 | Calibration Amount Diff -945.338% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_B12 | 1715336 |  | 24018852 | C5H11NO2 | 75.30 | Calibration Amount Diff -2052.739% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_blank | N/F |  | N/F | C5H11NO2 | N/A
Valine pos | 7.85 | Target Compound | KD100520_100520_C1 | 792031394 |  | 10647052071 | C5H11NO2 | 6869.47
Valine pos | 7.85 | Target Compound | KD100520_100520_C2 | 458937631 |  | 6525849988 | C5H11NO2 | 4125.23
Valine pos | 7.85 | Target Compound | KD100520_100520_C3 | 287072610 |  | 3908137305 | C5H11NO2 | 4582.12
Valine pos | 7.85 | Target Compound | KD100520_100520_C4 | 165899714 |  | 2259634610 | C5H11NO2 | 2334.14
Valine pos | 7.85 | Target Compound | KD100520_100520_C5 | 89092736 |  | 1275982449 | C5H11NO2 | 777.15
Valine pos | 7.85 | Target Compound | KD100520_100520_C6 | 47020197 |  | 672270294 | C5H11NO2 | 1106.33
Valine pos | 7.85 | Target Compound | KD100520_100520_C7 | 28565064 |  | 391839683 | C5H11NO2 | 507.44 | Calibration Amount Diff -22.171% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_C8 | 13951751 |  | 220760477 | C5H11NO2 | 189.95 | Calibration Amount Diff -73.32% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_C9 | 7143057 |  | 104805427 | C5H11NO2 | 155.78 | Calibration Amount Diff -219.171% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_C10 | 3700725 |  | 55643498 | C5H11NO2 | 61.67 | Calibration Amount Diff -494.893% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_C11 | 2510331 |  | 45702971 | C5H11NO2 | 34.50 | Calibration Amount Diff -952.483% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD100520_100520_C12 | 1643906 |  | 31274052 | C5H11NO2 | 52.42 | Calibration Amount Diff -1952.809% (absolute) > max calibration amount diff 20%

## Valine U-13C neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A1 | 335194 |  | 4040086 | [13]C5H11NO2 | 16075.53 | Rsquared 0.819 < R2Threshold 0.99
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A2 | 361834 |  | 4613668 | [13]C5H11NO2 | 29113.30 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 24.604% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A3 | 338232 |  | 4547209 | [13]C5H11NO2 | 2013.89 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 125.305% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A4 | 390392 |  | 5005354 | [13]C5H11NO2 | 7479.74 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 27.813% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A5 | 388446 |  | 4936849 | [13]C5H11NO2 | 1127.86 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 50.019% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A6 | 425636 |  | 5232697 | [13]C5H11NO2 | 1564.98 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -375.873% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A7 | 398852 |  | 5250619 | [13]C5H11NO2 | 1243.74 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -949.162% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A8 | 416378 |  | 5206239 | [13]C5H11NO2 | 1023.27 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -2683.655% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A9 | 405467 |  | 5164282 | [13]C5H11NO2 | 1242.07 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -4761.574% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A10 | 397057 |  | 5200135 | [13]C5H11NO2 | 1619.29 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -10923.854% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A11 | 390238 |  | 5049876 | [13]C5H11NO2 | 1393.73 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 3827.475% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_A12 | 415056 |  | 5287041 | [13]C5H11NO2 | 1329.08 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -32406.216% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B1 | 336908 |  | 4478198 | [13]C5H11NO2 | INF | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -26.852% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B2 | 350745 |  | 4710002 | [13]C5H11NO2 | 1679.24 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 29.157% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B3 | 394042 |  | 5036052 | [13]C5H11NO2 | 3839.82 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -26.936% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B4 | 390636 |  | 4889329 | [13]C5H11NO2 | 19414.83 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 148.664% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B5 | 402179 |  | 5119167 | [13]C5H11NO2 | 7173.57 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 63.693% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B6 | 377908 |  | 4662386 | [13]C5H11NO2 | 1181.09 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 789.304% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B7 | 400356 |  | 5197472 | [13]C5H11NO2 | 791.46 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -33.572% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B8 | 370197 |  | 4854989 | [13]C5H11NO2 | 1675.23 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -1363.563% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B9 | 389707 |  | 4802002 | [13]C5H11NO2 | 1024.25 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -546.4% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B10 | 354579 |  | 4963538 | [13]C5H11NO2 | 2984.86 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 2107.991% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B11 | 350661 |  | 4757610 | [13]C5H11NO2 | 1118.12 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -14456.394% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_B12 | 367363 |  | 4575933 | [13]C5H11NO2 | 2428.92 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 30126.235% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H11NO2 | N/A
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C1 | 344150 |  | 4065367 | [13]C5H11NO2 | 4623.27 | Rsquared 0.819 < R2Threshold 0.99
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C2 | 329173 |  | 4295260 | [13]C5H11NO2 | 2622.13 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 26.104% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C3 | 334292 |  | 4541798 | [13]C5H11NO2 | 13787.82 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 48.366% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C4 | 347885 |  | 4633638 | [13]C5H11NO2 | 1182.74 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -35.916% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C5 | 382464 |  | 4790570 | [13]C5H11NO2 | 1584.16 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -170.128% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C6 | 377224 |  | 4808536 | [13]C5H11NO2 | 2784.04 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff -505.079% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C7 | 403971 |  | 4949626 | [13]C5H11NO2 | 1074.69 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 842.154% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C8 | 361788 |  | 4859837 | [13]C5H11NO2 | 2433.11 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 647.534% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C9 | 393393 |  | 5108314 | [13]C5H11NO2 | 1178.75 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 200.779% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C10 | 390644 |  | 4957398 | [13]C5H11NO2 | 1163.93 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 1026.556% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C11 | 373925 |  | 4821417 | [13]C5H11NO2 | 1222.62 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 9759.617% (absolute) > max calibration amount diff 20%
Valine U-13C neg | 7.85 | Target Compound | KD100520_100520_C12 | 337554 |  | 4530699 | [13]C5H11NO2 | 1742.93 | Rsquared 0.819 < R2Threshold 0.99. Calibration Amount Diff 17961.297% (absolute) > max calibration amount diff 20%

## Valine U-13C pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A1 | 2986615 |  | 36910534 | [13]C5H11NO2 | INF | Rsquared 0.887 < R2Threshold 0.99
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A2 | 3861009 |  | 49935503 | [13]C5H11NO2 | INF | Rsquared 0.887 < R2Threshold 0.99
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A3 | 5071203 |  | 58892618 | [13]C5H11NO2 | 6634.60 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 75.295% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A4 | 5347999 |  | 71257418 | [13]C5H11NO2 | 16207.92 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 35.336% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A5 | 6101917 |  | 78002803 | [13]C5H11NO2 | 9917.04 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 27.217% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A6 | 6359117 |  | 84444086 | [13]C5H11NO2 | 9793.44 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -299.712% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A7 | 6034723 |  | 82837652 | [13]C5H11NO2 | 3388.07 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -453.212% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A8 | 6168414 |  | 81505800 | [13]C5H11NO2 | 2374.01 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -1621.192% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A9 | 5899502 |  | 79305319 | [13]C5H11NO2 | 3012.79 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -2760.857% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A10 | 5760256 |  | 77465566 | [13]C5H11NO2 | 4517.19 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -1410.554% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A11 | 5421728 |  | 75296119 | [13]C5H11NO2 | 3479.11 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 8224.767% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_A12 | 5505459 |  | 75234861 | [13]C5H11NO2 | 2145.20 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 14080.805% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B1 | 2817695 |  | 34188214 | [13]C5H11NO2 | INF | Rsquared 0.887 < R2Threshold 0.99
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B2 | 3779596 |  | 49267494 | [13]C5H11NO2 | INF | Rsquared 0.887 < R2Threshold 0.99
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B3 | 4578880 |  | 53888840 | [13]C5H11NO2 | 121234.52 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 122.855% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B4 | 5188744 |  | 71377542 | [13]C5H11NO2 | 65118.80 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -55.575% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B5 | 6390770 |  | 83628838 | [13]C5H11NO2 | 29057.15 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -217.078% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B6 | 5952693 |  | 80822478 | [13]C5H11NO2 | 3058.06 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -88.59% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B7 | 5954263 |  | 79210753 | [13]C5H11NO2 | 3616.11 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 67.412% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B8 | 5732329 |  | 78429841 | [13]C5H11NO2 | 4949.07 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -1231.3% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B9 | 5277096 |  | 75155892 | [13]C5H11NO2 | 5681.08 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -1332.552% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B10 | 5786524 |  | 78415251 | [13]C5H11NO2 | 3643.06 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -5102.281% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B11 | 4956484 |  | 69016972 | [13]C5H11NO2 | 2300.48 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 8835.478% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_B12 | 5215385 |  | 68638757 | [13]C5H11NO2 | 2881.12 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 15316.378% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_blank | N/F |  | N/F | [13]C5H11NO2 | N/A
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C1 | 3068695 |  | 34500518 | [13]C5H11NO2 | 3319.35 | Rsquared 0.887 < R2Threshold 0.99
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C2 | 3479100 |  | 45358561 | [13]C5H11NO2 | INF | Rsquared 0.887 < R2Threshold 0.99
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C3 | 4604238 |  | 54421234 | [13]C5H11NO2 | 8633.22 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 44.694% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C4 | 5192911 |  | 67081006 | [13]C5H11NO2 | 4873.12 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -30.535% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C5 | 5355709 |  | 74188408 | [13]C5H11NO2 | 4339.28 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -259.207% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C6 | 5706863 |  | 78633137 | [13]C5H11NO2 | 3421.05 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff -474.539% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C7 | 5571331 |  | 75020306 | [13]C5H11NO2 | 2646.77 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 334.927% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C8 | 5299528 |  | 75042994 | [13]C5H11NO2 | 5507.84 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 525.472% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C9 | 5027105 |  | 72606172 | [13]C5H11NO2 | 3429.41 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 1581.111% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C10 | 5297480 |  | 74178488 | [13]C5H11NO2 | 3857.18 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 1919.888% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C11 | 5488420 |  | 74562134 | [13]C5H11NO2 | 4670.91 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 2484.258% (absolute) > max calibration amount diff 20%
Valine U-13C pos | 7.85 | Target Compound | KD100520_100520_C12 | 4877397 |  | 67595135 | [13]C5H11NO2 | 4894.68 | Rsquared 0.887 < R2Threshold 0.99. Calibration Amount Diff 19590.361% (absolute) > max calibration amount diff 20%

## Valine U-13C, U-15N neg
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A1 | 10361207 |  | 143544837 | [13]C5H11[15]NO2 | INF
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A2 | 10283225 |  | 152676021 | [13]C5H11[15]NO2 | 572300.20
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A3 | 9957840 |  | 148606312 | [13]C5H11[15]NO2 | 155102.95
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A4 | 10623848 |  | 150912774 | [13]C5H11[15]NO2 | 1588359.20
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A5 | 10494265 |  | 146985579 | [13]C5H11[15]NO2 | 26069.51
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A6 | 11250720 |  | 150655830 | [13]C5H11[15]NO2 | 33458.69
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A7 | 10660064 |  | 149895120 | [13]C5H11[15]NO2 | 52601.92
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A8 | 10942342 |  | 146783989 | [13]C5H11[15]NO2 | 10024.39
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A9 | 10488870 |  | 146119023 | [13]C5H11[15]NO2 | 10735.77
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A10 | 10561152 |  | 146358124 | [13]C5H11[15]NO2 | 10888.73
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A11 | 10826814 |  | 148787886 | [13]C5H11[15]NO2 | 33177.48
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_A12 | 10747089 |  | 150253350 | [13]C5H11[15]NO2 | 9725.79
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B1 | 10492837 |  | 151657335 | [13]C5H11[15]NO2 | 927242.87
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B2 | 10588047 |  | 156614909 | [13]C5H11[15]NO2 | 74403.84
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B3 | 10980267 |  | 152508305 | [13]C5H11[15]NO2 | 466275.09
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B4 | 10856620 |  | 151815610 | [13]C5H11[15]NO2 | 319403.47
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B5 | 11157206 |  | 152660997 | [13]C5H11[15]NO2 | 77709.48
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B6 | 10626886 |  | 143833087 | [13]C5H11[15]NO2 | 40649.08
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B7 | 10810935 |  | 152338069 | [13]C5H11[15]NO2 | 6578.17
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B8 | 10077656 |  | 139461731 | [13]C5H11[15]NO2 | 40832.37
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B9 | 10476065 |  | 140020235 | [13]C5H11[15]NO2 | 13235.69
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B10 | 9893650 |  | 146376148 | [13]C5H11[15]NO2 | 25938.43
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B11 | 9245541 |  | 135637590 | [13]C5H11[15]NO2 | 36228.23
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_B12 | 10259917 |  | 137641718 | [13]C5H11[15]NO2 | 26166.37
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_blank | N/F |  | N/F | [13]C5H11[15]NO2 | N/A
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C1 | 10662516 |  | 146568678 | [13]C5H11[15]NO2 | INF
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C2 | 9718478 |  | 142364175 | [13]C5H11[15]NO2 | 808931.15
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C3 | 9564163 |  | 142719404 | [13]C5H11[15]NO2 | 1528433.94
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C4 | 9353829 |  | 137601903 | [13]C5H11[15]NO2 | 45859.18
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C5 | 9847664 |  | 139006001 | [13]C5H11[15]NO2 | 152042.94
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C6 | 9741636 |  | 137426892 | [13]C5H11[15]NO2 | 25556.28
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C7 | 11213811 |  | 148874164 | [13]C5H11[15]NO2 | 24356.68
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C8 | 9963301 |  | 143729472 | [13]C5H11[15]NO2 | 26436.49
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C9 | 11015044 |  | 149763086 | [13]C5H11[15]NO2 | 11147.62
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C10 | 10829886 |  | 145617525 | [13]C5H11[15]NO2 | 23121.78
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C11 | 10302253 |  | 143615962 | [13]C5H11[15]NO2 | 10161.65
Valine U-13C, U-15N neg | 7.85 | Internal Standard | KD100520_100520_C12 | 9259110 |  | 134749633 | [13]C5H11[15]NO2 | 10602.24

## Valine U-13C, U-15N pos
Compound | RT | Type | Filename | Height | Sample ID | Area | Formula | S/N | Flag Details
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A1 | 95047018 |  | 1336866662 | [13]C5H11[15]NO2 | INF
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A2 | 117658121 |  | 1648112308 | [13]C5H11[15]NO2 | INF
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A3 | 142358081 |  | 1849032526 | [13]C5H11[15]NO2 | 834438.61
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A4 | 147502838 |  | 2049897223 | [13]C5H11[15]NO2 | 528344.71
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A5 | 164512708 |  | 2183029804 | [13]C5H11[15]NO2 | 251286.76
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A6 | 169895376 |  | 2265931356 | [13]C5H11[15]NO2 | 277254.48
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A7 | 162171051 |  | 2227492161 | [13]C5H11[15]NO2 | 505486.36
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A8 | 158299423 |  | 2151839093 | [13]C5H11[15]NO2 | 445530.82
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A9 | 153789652 |  | 2102694140 | [13]C5H11[15]NO2 | 162158.38
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A10 | 156352128 |  | 2101120840 | [13]C5H11[15]NO2 | 302157.52
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A11 | 151861108 |  | 2108388835 | [13]C5H11[15]NO2 | 136524.72
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_A12 | 152884790 |  | 2099013027 | [13]C5H11[15]NO2 | 305256.27
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B1 | 95561489 |  | 1271934572 | [13]C5H11[15]NO2 | 1197904.46
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B2 | 115886349 |  | 1600285249 | [13]C5H11[15]NO2 | 2454658.00
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B3 | 133548068 |  | 1763136608 | [13]C5H11[15]NO2 | 229847.24
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B4 | 139015310 |  | 1983211826 | [13]C5H11[15]NO2 | 3410009.79
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B5 | 166532465 |  | 2237065128 | [13]C5H11[15]NO2 | 286816.12
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B6 | 163498688 |  | 2211099771 | [13]C5H11[15]NO2 | 557844.60
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B7 | 159899386 |  | 2181595801 | [13]C5H11[15]NO2 | 228322.51
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B8 | 148230381 |  | 2088802229 | [13]C5H11[15]NO2 | 202294.34
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B9 | 143057340 |  | 2025083731 | [13]C5H11[15]NO2 | 202190.67
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B10 | 151454900 |  | 2082792182 | [13]C5H11[15]NO2 | 316031.12
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B11 | 137633435 |  | 1936060455 | [13]C5H11[15]NO2 | 166380.71
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_B12 | 147665811 |  | 1918476706 | [13]C5H11[15]NO2 | 221487.41
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_blank | 1169 |  | 4318 | [13]C5H11[15]NO2 | INF | 7.626:Peak area 4317.761 is out of bounds (ISTD Minimum Recovery 973237947.779 and ISTD Max Recovery 2919713843.336)
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C1 | 108385621 |  | 1388384072 | [13]C5H11[15]NO2 | INF
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C2 | 108173000 |  | 1466228910 | [13]C5H11[15]NO2 | 407702.49
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C3 | 126531081 |  | 1665369736 | [13]C5H11[15]NO2 | 322070.94
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C4 | 136347733 |  | 1881535905 | [13]C5H11[15]NO2 | 706939.09
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C5 | 136260724 |  | 1969526692 | [13]C5H11[15]NO2 | 2013354.45
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C6 | 147456324 |  | 2077060832 | [13]C5H11[15]NO2 | 403925.62
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C7 | 153334224 |  | 2092243583 | [13]C5H11[15]NO2 | 291714.58
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C8 | 145016787 |  | 2080884657 | [13]C5H11[15]NO2 | 209561.81
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C9 | 137478173 |  | 2023513034 | [13]C5H11[15]NO2 | 179335.06
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C10 | 147662748 |  | 2051132661 | [13]C5H11[15]NO2 | 406898.65
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C11 | 149139308 |  | 2052991660 | [13]C5H11[15]NO2 | 333189.73
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD100520_100520_C12 | 137071198 |  | 1901298267 | [13]C5H11[15]NO2 | 509954.29

## Val_cali
Compound | Filename | Area | Area_IS | Response ratio | Conc.
Valine pos | KD100520_100520_A1 | 10542012293 | 1336866662 | 7.8856 | 500
Valine pos | KD100520_100520_A2 | 6235328590 | 1648112308 | 3.7833 | 250
Valine pos | KD100520_100520_A3 | 4360466498 | 1849032526 | 2.3582 | 125
Valine pos | KD100520_100520_A4 | 2437987028 | 2049897223 | 1.1893 | 62.5
Valine pos | KD100520_100520_A5 | 1729311640 | 2183029804 | 0.79216 | 31.25
Valine pos | KD100520_100520_A6 | 868507231 | 2265931356 | 0.38329 | 15.625
Valine pos | KD100520_100520_A7 | 414275220 | 2227492161 | 0.18598 | 7.8125
Valine pos | KD100520_100520_A8 | 208661362 | 2151839093 | 0.096969 | 3.9062
Valine pos | KD100520_100520_A9 | 122900858 | 2102694140 | 0.058449 | 1.9531
Valine pos | KD100520_100520_A10 | 67614515 | 2101120840 | 0.03218 | 0.97656
Valine pos | KD100520_100520_A11 | 51550938 | 2108388835 | 0.02445 | 0.48828
Valine pos | KD100520_100520_A12 | 28371879 | 2099013027 | 0.013517 | 0.24414
Valine pos | KD100520_100520_B1 | 11011246304 | 1271934572 | 8.6571 | 500
Valine pos | KD100520_100520_B2 | 6559113429 | 1600285249 | 4.0987 | 250
Valine pos | KD100520_100520_B3 | 4116029938 | 1763136608 | 2.3345 | 125
Valine pos | KD100520_100520_B4 | 2473092350 | 1983211826 | 1.247 | 62.5
Valine pos | KD100520_100520_B5 | 1304802175 | 2237065128 | 0.58327 | 31.25
Valine pos | KD100520_100520_B6 | 748244144 | 2211099771 | 0.3384 | 15.625
Valine pos | KD100520_100520_B7 | 405367695 | 2181595801 | 0.18581 | 7.8125
Valine pos | KD100520_100520_B8 | 221947441 | 2088802229 | 0.10626 | 3.9062
Valine pos | KD100520_100520_B9 | 120914881 | 2025083731 | 0.059709 | 1.9531
Valine pos | KD100520_100520_B10 | 94766479 | 2082792182 | 0.0455 | 0.97656
Valine pos | KD100520_100520_B11 | 44187811 | 1936060455 | 0.022824 | 0.48828
Valine pos | KD100520_100520_B12 | 24018852 | 1918476706 | 0.01252 | 0.24414
Valine pos | KD100520_100520_C1 | 10647052071 | 1388384072 | 7.6687 | 500
Valine pos | KD100520_100520_C2 | 6525849988 | 1466228910 | 4.4508 | 250
Valine pos | KD100520_100520_C3 | 3908137305 | 1665369736 | 2.3467 | 125
Valine pos | KD100520_100520_C4 | 2259634610 | 1881535905 | 1.201 | 62.5
Valine pos | KD100520_100520_C5 | 1275982449 | 1969526692 | 0.64786 | 31.25
Valine pos | KD100520_100520_C6 | 672270294 | 2077060832 | 0.32366 | 15.625
Valine pos | KD100520_100520_C7 | 391839683 | 2092243583 | 0.18728 | 7.8125
Valine pos | KD100520_100520_C8 | 220760477 | 2080884657 | 0.10609 | 3.9062
Valine pos | KD100520_100520_C9 | 104805427 | 2023513034 | 0.051794 | 1.9531
Valine pos | KD100520_100520_C10 | 55643498 | 2051132661 | 0.027128 | 0.97656
Valine pos | KD100520_100520_C11 | 45702971 | 2052991660 | 0.022262 | 0.48828
Valine pos | KD100520_100520_C12 | 31274052 | 1901298267 | 0.016449 | 0.24414