## Alanine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Alanine pos | 9.49 | Target Compound | KD033022_033022_A1 | 5317406467 | 500 | A01 | C3H7NO2 | 103.811 | 51222114
Alanine pos | 9.49 | Target Compound | KD033022_033022_A2 | 3431759283 | 250 | A02 | C3H7NO2 | 42.83 | 80126052
Alanine pos | 9.49 | Target Compound | KD033022_033022_A3 | 2159427856 | 125 | A03 | C3H7NO2 | 20.717 | 104235923
Alanine pos | 9.49 | Target Compound | KD033022_033022_A4 | 1434632162 | 62.5 | A04 | C3H7NO2 | 10.357 | 138514864
Alanine pos | 9.49 | Target Compound | KD033022_033022_A5 | 949879639 | 31.25 | A05 | C3H7NO2 | 5.652 | 168060332
Alanine pos | 9.49 | Target Compound | KD033022_033022_A6 | 532611551 | 15.625 | A06 | C3H7NO2 | 2.722 | 195636186
Alanine pos | 9.49 | Target Compound | KD033022_033022_A7 | 317818893 | 7.813 | A07 | C3H7NO2 | 1.367 | 232471224
Alanine pos | 9.49 | Target Compound | KD033022_033022_A8 | 182709175 | 3.906 | A08 | C3H7NO2 | 0.674 | 271130614
Alanine pos | 9.49 | Target Compound | KD033022_033022_A9 | 104797957 | 1.953 | A09 | C3H7NO2 | 0.391 | 268129403 | Calibration Amount Diff 29.2% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_A10 | 54651329 | 0.977 | A10 | C3H7NO2 | 0.185 | 294635247 | Calibration Amount Diff 22.667% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_A11 | 31680668 | 0.488 | A11 | C3H7NO2 | 0.101 | 313447975 | Calibration Amount Diff 33.863% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_A12 | 17496121 | 0.244 | A12 | C3H7NO2 | 0.055 | 320853572 | Calibration Amount Diff 44.469% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_B1 | 5054055256 | 500 | B01 | C3H7NO2 | 102.159 | 49472550
Alanine pos | 9.49 | Target Compound | KD033022_033022_B2 | 3393584693 | 250 | B02 | C3H7NO2 | 41.657 | 81464767
Alanine pos | 9.49 | Target Compound | KD033022_033022_B3 | 2228722949 | 125 | B03 | C3H7NO2 | 20.125 | 110741924
Alanine pos | 9.49 | Target Compound | KD033022_033022_B4 | 1426775548 | 62.5 | B04 | C3H7NO2 | 9.923 | 143781608
Alanine pos | 9.49 | Target Compound | KD033022_033022_B5 | 899275185 | 31.25 | B05 | C3H7NO2 | 5.289 | 170031038
Alanine pos | 9.49 | Target Compound | KD033022_033022_B6 | 584579531 | 15.625 | B06 | C3H7NO2 | 2.815 | 207685200
Alanine pos | 9.49 | Target Compound | KD033022_033022_B7 | 310533610 | 7.813 | B07 | C3H7NO2 | 1.258 | 246775235
Alanine pos | 9.49 | Target Compound | KD033022_033022_B8 | 192424147 | 3.906 | B08 | C3H7NO2 | 0.673 | 285844105
Alanine pos | 9.49 | Target Compound | KD033022_033022_B9 | 110955401 | 1.953 | B09 | C3H7NO2 | 0.361 | 307124299
Alanine pos | 9.49 | Target Compound | KD033022_033022_B10 | 56227674 | 0.977 | B10 | C3H7NO2 | 0.188 | 298986033 | Calibration Amount Diff 24.367% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_B11 | 31364809 | 0.488 | B11 | C3H7NO2 | 0.099 | 316087897 | Calibration Amount Diff 31.422% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_B12 | 18303991 | 0.244 | B12 | C3H7NO2 | 0.056 | 328945732 | Calibration Amount Diff 47.421% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_C1 | 5135390419 | 500 | C01 | C3H7NO2 | 99.565 | 51578133
Alanine pos | 9.49 | Target Compound | KD033022_033022_C2 | 3288530082 | 250 | C02 | C3H7NO2 | 43.622 | 75386731
Alanine pos | 9.49 | Target Compound | KD033022_033022_C3 | 2075256049 | 125 | C03 | C3H7NO2 | 20.483 | 101317144
Alanine pos | 9.49 | Target Compound | KD033022_033022_C4 | 1391753881 | 62.5 | C04 | C3H7NO2 | 10.341 | 134585078
Alanine pos | 9.49 | Target Compound | KD033022_033022_C5 | 835152745 | 31.25 | C05 | C3H7NO2 | 5.258 | 158830179
Alanine pos | 9.49 | Target Compound | KD033022_033022_C6 | 532387446 | 15.625 | C06 | C3H7NO2 | 2.698 | 197313583
Alanine pos | 9.49 | Target Compound | KD033022_033022_C7 | 295470280 | 7.813 | C07 | C3H7NO2 | 1.262 | 234069916
Alanine pos | 9.49 | Target Compound | KD033022_033022_C8 | 174424271 | 3.906 | C08 | C3H7NO2 | 0.682 | 255794475
Alanine pos | 9.49 | Target Compound | KD033022_033022_C9 | 118563875 | 1.953 | C09 | C3H7NO2 | 0.384 | 308708614 | Calibration Amount Diff 26.961% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_C10 | 59487711 | 0.977 | C10 | C3H7NO2 | 0.189 | 314663361 | Calibration Amount Diff 25.022% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_C11 | 32671948 | 0.488 | C11 | C3H7NO2 | 0.105 | 312409977 | Calibration Amount Diff 38.508% (absolute) > max calibration amount diff 20%
Alanine pos | 9.49 | Target Compound | KD033022_033022_C12 | 15900617 | 0.244 | C12 | C3H7NO2 | 0.055 | 289327021 | Calibration Amount Diff 45.601% (absolute) > max calibration amount diff 20%

## Alanine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A1 | 51222114 | 1 | A01 | [13]C3H7[15]NO2 | N/A | N/A | 9.540:Peak area 51222114.260 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A2 | 80126052 | 1 | A02 | [13]C3H7[15]NO2 | N/A | N/A | 9.531:Peak area 80126051.955 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A3 | 104235923 | 1 | A03 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A4 | 138514864 | 1 | A04 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A5 | 168060332 | 1 | A05 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A6 | 195636186 | 1 | A06 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A7 | 232471224 | 1 | A07 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A8 | 271130614 | 1 | A08 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A9 | 268129403 | 1 | A09 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A10 | 294635247 | 1 | A10 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A11 | 313447975 | 1 | A11 | [13]C3H7[15]NO2 | N/A | N/A | 9.493:Peak area 313447975.440 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_A12 | 320853572 | 1 | A12 | [13]C3H7[15]NO2 | N/A | N/A | 9.491:Peak area 320853572.002 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B1 | 49472550 | 1 | B01 | [13]C3H7[15]NO2 | N/A | N/A | 9.519:Peak area 49472550.343 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B2 | 81464767 | 1 | B02 | [13]C3H7[15]NO2 | N/A | N/A | 9.531:Peak area 81464766.651 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B3 | 110741924 | 1 | B03 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B4 | 143781608 | 1 | B04 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B5 | 170031038 | 1 | B05 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B6 | 207685200 | 1 | B06 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B7 | 246775235 | 1 | B07 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B8 | 285844105 | 1 | B08 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B9 | 307124299 | 1 | B09 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B10 | 298986033 | 1 | B10 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B11 | 316087897 | 1 | B11 | [13]C3H7[15]NO2 | N/A | N/A | 9.489:Peak area 316087897.363 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_B12 | 328945732 | 1 | B12 | [13]C3H7[15]NO2 | N/A | N/A | 9.498:Peak area 328945732.219 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C1 | 51578133 | 1 | C01 | [13]C3H7[15]NO2 | N/A | N/A | 9.520:Peak area 51578133.361 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C2 | 75386731 | 1 | C02 | [13]C3H7[15]NO2 | N/A | N/A | 9.508:Peak area 75386731.311 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C3 | 101317144 | 1 | C03 | [13]C3H7[15]NO2 | N/A | N/A | 9.506:Peak area 101317144.451 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C4 | 134585078 | 1 | C04 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C5 | 158830179 | 1 | C05 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C6 | 197313583 | 1 | C06 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C7 | 234069916 | 1 | C07 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C8 | 255794475 | 1 | C08 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C9 | 308708614 | 1 | C09 | [13]C3H7[15]NO2 | N/A | N/A
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C10 | 314663361 | 1 | C10 | [13]C3H7[15]NO2 | N/A | N/A | 9.500:Peak area 314663361.439 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C11 | 312409977 | 1 | C11 | [13]C3H7[15]NO2 | N/A | N/A | 9.507:Peak area 312409976.771 is out of bounds (ISTD Minimum Recovery 103047057.039 and ISTD Max Recovery 309141171.116)
Alanine U-13C, U-15N pos | 9.49 | Internal Standard | KD033022_033022_C12 | 289327021 | 1 | C12 | [13]C3H7[15]NO2 | N/A | N/A

## Arginine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Arginine pos | 14.24 | Target Compound | KD033022_033022_A1 | 68060155567 | 500 | A01 | C6H14N4O2 | 89.158 | 763366192
Arginine pos | 14.24 | Target Compound | KD033022_033022_A2 | 44784970832 | 250 | A02 | C6H14N4O2 | 40.714 | 1099993338
Arginine pos | 14.24 | Target Compound | KD033022_033022_A3 | 31768623627 | 125 | A03 | C6H14N4O2 | 21.013 | 1511869719
Arginine pos | 14.24 | Target Compound | KD033022_033022_A4 | 21293771863 | 62.5 | A04 | C6H14N4O2 | 11.314 | 1882120429
Arginine pos | 14.24 | Target Compound | KD033022_033022_A5 | 14693766028 | 31.25 | A05 | C6H14N4O2 | 6.301 | 2332031519
Arginine pos | 14.24 | Target Compound | KD033022_033022_A6 | 8886517721 | 15.625 | A06 | C6H14N4O2 | 3.235 | 2746860761
Arginine pos | 14.24 | Target Compound | KD033022_033022_A7 | 5162607088 | 7.813 | A07 | C6H14N4O2 | 1.676 | 3080240162 | Calibration Amount Diff 20.861% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_A8 | 2815805121 | 3.906 | A08 | C6H14N4O2 | 0.854 | 3295655660 | Calibration Amount Diff 23.188% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_A9 | 1559689307 | 1.953 | A09 | C6H14N4O2 | 0.489 | 3191078228 | Calibration Amount Diff 40.915% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_A10 | 823847274 | 0.977 | A10 | C6H14N4O2 | 0.242 | 3408098655 | Calibration Amount Diff 39.298% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_A11 | 469240672 | 0.488 | A11 | C6H14N4O2 | 0.129 | 3639887614 | Calibration Amount Diff 48.72% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_A12 | 236154953 | 0.244 | A12 | C6H14N4O2 | 0.066 | 3560508819 | Calibration Amount Diff 53.025% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_B1 | 66025369993 | 500 | B01 | C6H14N4O2 | 83.916 | 786806047
Arginine pos | 14.24 | Target Compound | KD033022_033022_B2 | 45238097786 | 250 | B02 | C6H14N4O2 | 40.011 | 1130639543
Arginine pos | 14.24 | Target Compound | KD033022_033022_B3 | 31556668888 | 125 | B03 | C6H14N4O2 | 21.116 | 1494442262
Arginine pos | 14.24 | Target Compound | KD033022_033022_B4 | 21349616117 | 62.5 | B04 | C6H14N4O2 | 11.315 | 1886815588
Arginine pos | 14.24 | Target Compound | KD033022_033022_B5 | 14688947652 | 31.25 | B05 | C6H14N4O2 | 6.06 | 2424059270
Arginine pos | 14.24 | Target Compound | KD033022_033022_B6 | 9280832058 | 15.625 | B06 | C6H14N4O2 | 3.363 | 2760008676 | Calibration Amount Diff 21.353% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_B7 | 4798722795 | 7.813 | B07 | C6H14N4O2 | 1.551 | 3093032546
Arginine pos | 14.24 | Target Compound | KD033022_033022_B8 | 2900718899 | 3.906 | B08 | C6H14N4O2 | 0.857 | 3386036383 | Calibration Amount Diff 23.516% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_B9 | 1610515418 | 1.953 | B09 | C6H14N4O2 | 0.47 | 3425550843 | Calibration Amount Diff 35.546% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_B10 | 850898812 | 0.977 | B10 | C6H14N4O2 | 0.24 | 3541659188 | Calibration Amount Diff 38.446% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_B11 | 459686962 | 0.488 | B11 | C6H14N4O2 | 0.129 | 3564038512 | Calibration Amount Diff 48.792% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_B12 | 242806592 | 0.244 | B12 | C6H14N4O2 | 0.068 | 3549643032 | Calibration Amount Diff 57.817% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_C1 | 65418427383 | 500 | C01 | C6H14N4O2 | 86.125 | 759575435
Arginine pos | 14.24 | Target Compound | KD033022_033022_C2 | 42986475418 | 250 | C02 | C6H14N4O2 | 40.449 | 1062728376
Arginine pos | 14.24 | Target Compound | KD033022_033022_C3 | 30799140780 | 125 | C03 | C6H14N4O2 | 21.851 | 1409515180
Arginine pos | 14.24 | Target Compound | KD033022_033022_C4 | 21322662600 | 62.5 | C04 | C6H14N4O2 | 11.273 | 1891492538
Arginine pos | 14.24 | Target Compound | KD033022_033022_C5 | 13523229461 | 31.25 | C05 | C6H14N4O2 | 6.241 | 2166792308
Arginine pos | 14.24 | Target Compound | KD033022_033022_C6 | 8776897485 | 15.625 | C06 | C6H14N4O2 | 3.247 | 2702687321
Arginine pos | 14.24 | Target Compound | KD033022_033022_C7 | 4485485050 | 7.813 | C07 | C6H14N4O2 | 1.564 | 2868482677
Arginine pos | 14.24 | Target Compound | KD033022_033022_C8 | 2655675631 | 3.906 | C08 | C6H14N4O2 | 0.853 | 3114305682 | Calibration Amount Diff 22.948% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_C9 | 1680177858 | 1.953 | C09 | C6H14N4O2 | 0.49 | 3430975291 | Calibration Amount Diff 41.187% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_C10 | 830015691 | 0.977 | C10 | C6H14N4O2 | 0.243 | 3411771652 | Calibration Amount Diff 40.19% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_C11 | 438719838 | 0.488 | C11 | C6H14N4O2 | 0.13 | 3365081354 | Calibration Amount Diff 50.402% (absolute) > max calibration amount diff 20%
Arginine pos | 14.24 | Target Compound | KD033022_033022_C12 | 227448353 | 0.244 | C12 | C6H14N4O2 | 0.068 | 3346592152 | Calibration Amount Diff 56.804% (absolute) > max calibration amount diff 20%

## Arginine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A1 | 763366192 | 1 | A01 | [13]C6H14[15]N4O2 | N/A | N/A | 14.300:Peak area 763366192.029 is out of bounds (ISTD Minimum Recovery 1265061707.672 and ISTD Max Recovery 3795185123.015)
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A2 | 1099993338 | 1 | A02 | [13]C6H14[15]N4O2 | N/A | N/A | 14.305:Peak area 1099993338.077 is out of bounds (ISTD Minimum Recovery 1265061707.672 and ISTD Max Recovery 3795185123.015)
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A3 | 1511869719 | 1 | A03 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A4 | 1882120429 | 1 | A04 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A5 | 2332031519 | 1 | A05 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A6 | 2746860761 | 1 | A06 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A7 | 3080240162 | 1 | A07 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A8 | 3295655660 | 1 | A08 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A9 | 3191078228 | 1 | A09 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A10 | 3408098655 | 1 | A10 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A11 | 3639887614 | 1 | A11 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_A12 | 3560508819 | 1 | A12 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B1 | 786806047 | 1 | B01 | [13]C6H14[15]N4O2 | N/A | N/A | 14.281:Peak area 786806046.617 is out of bounds (ISTD Minimum Recovery 1265061707.672 and ISTD Max Recovery 3795185123.015)
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B2 | 1130639543 | 1 | B02 | [13]C6H14[15]N4O2 | N/A | N/A | 14.335:Peak area 1130639543.387 is out of bounds (ISTD Minimum Recovery 1265061707.672 and ISTD Max Recovery 3795185123.015)
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B3 | 1494442262 | 1 | B03 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B4 | 1886815588 | 1 | B04 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B5 | 2424059270 | 1 | B05 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B6 | 2760008676 | 1 | B06 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B7 | 3093032546 | 1 | B07 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B8 | 3386036383 | 1 | B08 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B9 | 3425550843 | 1 | B09 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B10 | 3541659188 | 1 | B10 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B11 | 3564038512 | 1 | B11 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_B12 | 3549643032 | 1 | B12 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C1 | 759575435 | 1 | C01 | [13]C6H14[15]N4O2 | N/A | N/A | 14.313:Peak area 759575434.991 is out of bounds (ISTD Minimum Recovery 1265061707.672 and ISTD Max Recovery 3795185123.015)
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C2 | 1062728376 | 1 | C02 | [13]C6H14[15]N4O2 | N/A | N/A | 14.314:Peak area 1062728375.966 is out of bounds (ISTD Minimum Recovery 1265061707.672 and ISTD Max Recovery 3795185123.015)
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C3 | 1409515180 | 1 | C03 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C4 | 1891492538 | 1 | C04 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C5 | 2166792308 | 1 | C05 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C6 | 2702687321 | 1 | C06 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C7 | 2868482677 | 1 | C07 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C8 | 3114305682 | 1 | C08 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C9 | 3430975291 | 1 | C09 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C10 | 3411771652 | 1 | C10 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C11 | 3365081354 | 1 | C11 | [13]C6H14[15]N4O2 | N/A | N/A
Arginine U-13C, U-15N pos | 14.24 | Internal Standard | KD033022_033022_C12 | 3346592152 | 1 | C12 | [13]C6H14[15]N4O2 | N/A | N/A

## Asparagine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | RR | Theoretical Amt | power pred | %diff
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A1 | 6630301238 | 500 | A01 | C4H8N2O3 | 196.172 | 33798409
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A2 | 3812570840 | 250 | A02 | C4H8N2O3 | 67.967 | 56094804 | 67.9666 | 250 | 264.4649 | 5.7859
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A3 | 2284951695 | 125 | A03 | C4H8N2O3 | 32.493 | 70322377 | 32.4925 | 125 | 124.7721 | 0.18234
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A4 | 1411952066 | 62.5 | A04 | C4H8N2O3 | 16.733 | 84382157 | 16.7328 | 62.5 | 63.4955 | 1.5928
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A5 | 960415427 | 31.25 | A05 | C4H8N2O3 | 8.854 | 108477019 | 8.8536 | 31.25 | 33.2158 | 6.2906
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A6 | 573650501 | 15.625 | A06 | C4H8N2O3 | 4.395 | 130535534 | 4.3946 | 15.625 | 16.2815 | 4.2019
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A7 | 367695509 | 7.813 | A07 | C4H8N2O3 | 2.19 | 167872636 | 2.1903 | 7.813 | 8.0144 | 2.5777
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A8 | 240803397 | 3.906 | A08 | C4H8N2O3 | 1.07 | 225083253 | 1.0698 | 3.906 | 3.8646 | 1.0587
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A9 | 147431006 | 1.953 | A09 | C4H8N2O3 | 0.606 | 243339141 | 0.60587 | 1.953 | 2.1664 | 10.9286
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A10 | 85815250 | 0.977 | A10 | C4H8N2O3 | 0.283 | 303369098 | 0.28287 | 0.977 | 0.99779 | 2.128
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A11 | 48497296 | 0.488 | A11 | C4H8N2O3 | 0.15 | 324199432 | 0.14959 | 0.488 | 0.52167 | 6.9
Asparagine pos | 9.77 | Target Compound | KD033022_033022_A12 | 26298334 | 0.244 | A12 | C4H8N2O3 | 0.071 | 369857419 | 0.071104 | 0.244 | 0.24468 | 0.27962
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B1 | 6057905857 | 500 | B01 | C4H8N2O3 | 165.8 | 36537388
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B2 | 3656729832 | 250 | B02 | C4H8N2O3 | 67.683 | 54027290 | 67.683 | 250 | 263.3418 | 5.3367
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B3 | 2234928594 | 125 | B03 | C4H8N2O3 | 31.983 | 69877894 | 31.9833 | 125 | 122.7821 | 1.7743
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B4 | 1400043975 | 62.5 | B04 | C4H8N2O3 | 16.388 | 85429529 | 16.3883 | 62.5 | 62.1649 | 0.53614
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B5 | 894546124 | 31.25 | B05 | C4H8N2O3 | 8.318 | 107539853 | 8.3183 | 31.25 | 31.1725 | 0.24793
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B6 | 605752419 | 15.625 | B06 | C4H8N2O3 | 4.319 | 140261738 | 4.3187 | 15.625 | 15.9955 | 2.3711
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B7 | 352942435 | 7.813 | B07 | C4H8N2O3 | 1.904 | 185325224 | 1.9044 | 7.813 | 6.9509 | 11.0336
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B8 | 254223901 | 3.906 | B08 | C4H8N2O3 | 1.045 | 243210360 | 1.0453 | 3.906 | 3.7744 | 3.37
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B9 | 154633137 | 1.953 | B09 | C4H8N2O3 | 0.548 | 282037106 | 0.54827 | 1.953 | 1.957 | 0.20427
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B10 | 86859091 | 0.977 | B10 | C4H8N2O3 | 0.273 | 318508509 | 0.27271 | 0.977 | 0.96129 | 1.6077
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B11 | 51427995 | 0.488 | B11 | C4H8N2O3 | 0.141 | 365763444 | 0.1406 | 0.488 | 0.48979 | 0.36675
Asparagine pos | 9.77 | Target Compound | KD033022_033022_B12 | 27703673 | 0.244 | B12 | C4H8N2O3 | 0.071 | 388654199 | 0.071281 | 0.244 | 0.2453 | 0.5338
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C1 | 5521156781 | 500 | C01 | C4H8N2O3 | 166.644 | 33131409
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C2 | 3293379656 | 250 | C02 | C4H8N2O3 | 64.83 | 50800175 | 64.8301 | 250 | 252.0472 | 0.81888
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C3 | 2094702469 | 125 | C03 | C4H8N2O3 | 31.996 | 65468519 | 31.9956 | 125 | 122.8298 | 1.7361
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C4 | 1292197359 | 62.5 | C04 | C4H8N2O3 | 15.612 | 82769292 | 15.612 | 62.5 | 59.169 | 5.3297
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C5 | 854980239 | 31.25 | C05 | C4H8N2O3 | 8.083 | 105779020 | 8.0827 | 31.25 | 30.2741 | 3.1228
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C6 | 583996191 | 15.625 | C06 | C4H8N2O3 | 4.114 | 141954398 | 4.114 | 15.625 | 15.2239 | 2.5673
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C7 | 369403656 | 7.813 | C07 | C4H8N2O3 | 1.902 | 194226840 | 1.9019 | 7.813 | 6.9415 | 11.1539
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C8 | 232634693 | 3.906 | C08 | C4H8N2O3 | 1.012 | 229835210 | 1.0122 | 3.906 | 3.6527 | 6.4841
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C9 | 164027468 | 1.953 | C09 | C4H8N2O3 | 0.563 | 291455960 | 0.56279 | 1.953 | 2.0097 | 2.9051
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C10 | 94219951 | 0.977 | C10 | C4H8N2O3 | 0.278 | 338912350 | 0.27801 | 0.977 | 0.98032 | 0.33956
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C11 | 51587196 | 0.488 | C11 | C4H8N2O3 | 0.142 | 363094298 | 0.14208 | 0.488 | 0.49501 | 1.4364
Asparagine pos | 9.77 | Target Compound | KD033022_033022_C12 | 25635141 | 0.244 | C12 | C4H8N2O3 | 0.07 | 366489321 | 0.069948 | 0.244 | 0.24063 | 1.3798

## Asparagine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A1 | 33798409 | 1 | A01 | [13]C4H8[15]N2O3 | N/A | N/A | 9.817:Peak area 33798409.405 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A2 | 56094804 | 1 | A02 | [13]C4H8[15]N2O3 | N/A | N/A | 9.807:Peak area 56094804.313 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A3 | 70322377 | 1 | A03 | [13]C4H8[15]N2O3 | N/A | N/A | 9.807:Peak area 70322377.082 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A4 | 84382157 | 1 | A04 | [13]C4H8[15]N2O3 | N/A | N/A | 9.812:Peak area 84382157.285 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A5 | 108477019 | 1 | A05 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A6 | 130535534 | 1 | A06 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A7 | 167872636 | 1 | A07 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A8 | 225083253 | 1 | A08 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A9 | 243339141 | 1 | A09 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A10 | 303369098 | 1 | A10 | [13]C4H8[15]N2O3 | N/A | N/A | 9.822:Peak area 303369097.925 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A11 | 324199432 | 1 | A11 | [13]C4H8[15]N2O3 | N/A | N/A | 9.831:Peak area 324199432.227 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_A12 | 369857419 | 1 | A12 | [13]C4H8[15]N2O3 | N/A | N/A | 9.859:Peak area 369857419.111 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B1 | 36537388 | 1 | B01 | [13]C4H8[15]N2O3 | N/A | N/A | 9.851:Peak area 36537388.114 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B2 | 54027290 | 1 | B02 | [13]C4H8[15]N2O3 | N/A | N/A | 9.834:Peak area 54027290.219 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B3 | 69877894 | 1 | B03 | [13]C4H8[15]N2O3 | N/A | N/A | 9.831:Peak area 69877894.110 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B4 | 85429529 | 1 | B04 | [13]C4H8[15]N2O3 | N/A | N/A | 9.832:Peak area 85429529.034 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B5 | 107539853 | 1 | B05 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B6 | 140261738 | 1 | B06 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B7 | 185325224 | 1 | B07 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B8 | 243210360 | 1 | B08 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B9 | 282037106 | 1 | B09 | [13]C4H8[15]N2O3 | N/A | N/A | 9.837:Peak area 282037105.778 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B10 | 318508509 | 1 | B10 | [13]C4H8[15]N2O3 | N/A | N/A | 9.851:Peak area 318508509.156 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B11 | 365763444 | 1 | B11 | [13]C4H8[15]N2O3 | N/A | N/A | 9.856:Peak area 365763443.774 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_B12 | 388654199 | 1 | B12 | [13]C4H8[15]N2O3 | N/A | N/A | 9.864:Peak area 388654199.398 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C1 | 33131409 | 1 | C01 | [13]C4H8[15]N2O3 | N/A | N/A | 9.824:Peak area 33131409.249 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C2 | 50800175 | 1 | C02 | [13]C4H8[15]N2O3 | N/A | N/A | 9.839:Peak area 50800174.970 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C3 | 65468519 | 1 | C03 | [13]C4H8[15]N2O3 | N/A | N/A | 9.838:Peak area 65468519.485 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C4 | 82769292 | 1 | C04 | [13]C4H8[15]N2O3 | N/A | N/A | 9.815:Peak area 82769291.699 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C5 | 105779020 | 1 | C05 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C6 | 141954398 | 1 | C06 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C7 | 194226840 | 1 | C07 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C8 | 229835210 | 1 | C08 | [13]C4H8[15]N2O3 | N/A | N/A
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C9 | 291455960 | 1 | C09 | [13]C4H8[15]N2O3 | N/A | N/A | 9.821:Peak area 291455959.656 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C10 | 338912350 | 1 | C10 | [13]C4H8[15]N2O3 | N/A | N/A | 9.838:Peak area 338912349.581 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C11 | 363094298 | 1 | C11 | [13]C4H8[15]N2O3 | N/A | N/A | 9.874:Peak area 363094297.834 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)
Asparagine U-13C, U-15N pos | 9.77 | Internal Standard | KD033022_033022_C12 | 366489321 | 1 | C12 | [13]C4H8[15]N2O3 | N/A | N/A | 9.861:Peak area 366489320.815 is out of bounds (ISTD Minimum Recovery 92478063.990 and ISTD Max Recovery 277434191.970)

## Aspartate neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | RR | Theoretical Amt | Power pred | %diff
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A1 | 9505475013 | 500 | A01 | C4H7NO4 | 91.288 | 104126550 | 91.2877 | 500 | 567.6922 | 13.5384
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A2 | 5375568234 | 250 | A02 | C4H7NO4 | 40.298 | 133394152 | 40.2984 | 250 | 237.5262 | 4.9895
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A3 | 3885416389 | 125 | A03 | C4H7NO4 | 21.199 | 183284491 | 21.1988 | 125 | 119.7982 | 4.1614
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A4 | 3188682403 | 62.5 | A04 | C4H7NO4 | 11.231 | 283922238 | 11.2308 | 62.5 | 60.8789 | 2.5938
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A5 | 2368070884 | 31.25 | A05 | C4H7NO4 | 6.213 | 381124511 | 6.2134 | 31.25 | 32.399 | 3.6769
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A6 | 1553684153 | 15.625 | A06 | C4H7NO4 | 3.163 | 491201220 | 3.163 | 15.625 | 15.7793 | 0.98756
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A7 | 983503456 | 7.813 | A07 | C4H7NO4 | 1.598 | 615402697 | 1.5981 | 7.813 | 7.6237 | 2.4224
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A8 | 542206946 | 3.906 | A08 | C4H7NO4 | 0.817 | 663701678 | 0.81694 | 3.906 | 3.7294 | 4.5206
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A9 | 307817589 | 1.953 | A09 | C4H7NO4 | 0.465 | 662671558 | 0.46451 | 1.953 | 2.0435 | 4.6337
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A10 | 165473899 | 0.977 | A10 | C4H7NO4 | 0.229 | 723929472 | 0.22858 | 0.977 | 0.9599 | 1.7502
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A11 | 91934267 | 0.488 | A11 | C4H7NO4 | 0.122 | 751326187 | 0.12236 | 0.488 | 0.49324 | 1.0728
Aspartate neg | 10.26 | Target Compound | KD033022_033022_A12 | 50948060 | 0.244 | A12 | C4H7NO4 | 0.065 | 783189475 | 0.065052 | 0.244 | 0.25158 | 3.1078
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B1 | 9257377773 | 500 | B01 | C4H7NO4 | 86.487 | 107037424 | 86.4873 | 500 | 535.9388 | 7.1878
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B2 | 5676418961 | 250 | B02 | C4H7NO4 | 40.713 | 139424184 | 40.7133 | 250 | 240.133 | 3.9468
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B3 | 4298601583 | 125 | B03 | C4H7NO4 | 21.146 | 203280829 | 21.1461 | 125 | 119.4809 | 4.4153
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B4 | 3130887300 | 62.5 | B04 | C4H7NO4 | 11.265 | 277920637 | 11.2654 | 62.5 | 61.0786 | 2.2743
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B5 | 2284976292 | 31.25 | B05 | C4H7NO4 | 5.984 | 381849401 | 5.984 | 31.25 | 31.126 | 0.39694
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B6 | 1652125828 | 15.625 | B06 | C4H7NO4 | 3.272 | 505002763 | 3.2715 | 15.625 | 16.3566 | 4.6825
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B7 | 984336296 | 7.813 | B07 | C4H7NO4 | 1.487 | 661822723 | 1.4873 | 7.813 | 7.0617 | 9.6164
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B8 | 625389427 | 3.906 | B08 | C4H7NO4 | 0.83 | 753578003 | 0.82989 | 3.906 | 3.7925 | 2.907
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B9 | 335613436 | 1.953 | B09 | C4H7NO4 | 0.446 | 751941929 | 0.44633 | 1.953 | 1.9584 | 0.27549
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B10 | 186652958 | 0.977 | B10 | C4H7NO4 | 0.23 | 812870812 | 0.22962 | 0.977 | 0.96458 | 1.2717
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B11 | 100620553 | 0.488 | B11 | C4H7NO4 | 0.122 | 824376630 | 0.12206 | 0.488 | 0.49192 | 0.8034
Aspartate neg | 10.26 | Target Compound | KD033022_033022_B12 | 56479086 | 0.244 | B12 | C4H7NO4 | 0.066 | 854706387 | 0.06608 | 0.244 | 0.25582 | 4.845
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C1 | 9308849904 | 500 | C01 | C4H7NO4 | 89.335 | 104202136 | 89.3345 | 500 | 554.759 | 10.9518
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C2 | 5767960227 | 250 | C02 | C4H7NO4 | 40.734 | 141599321 | 40.7344 | 250 | 240.2655 | 3.8938
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C3 | 4220942507 | 125 | C03 | C4H7NO4 | 21.751 | 194058996 | 21.7508 | 125 | 123.1249 | 1.5001
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C4 | 3297951884 | 62.5 | C04 | C4H7NO4 | 11.168 | 295309099 | 11.1678 | 62.5 | 60.5148 | 3.1763
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C5 | 2382607067 | 31.25 | C05 | C4H7NO4 | 6.096 | 390869748 | 6.0957 | 31.25 | 31.7453 | 1.5851
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C6 | 1715980069 | 15.625 | C06 | C4H7NO4 | 3.168 | 541623092 | 3.1682 | 15.625 | 15.8069 | 1.1641
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C7 | 991633725 | 7.813 | C07 | C4H7NO4 | 1.504 | 659165756 | 1.5044 | 7.813 | 7.148 | 8.511
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C8 | 591254631 | 3.906 | C08 | C4H7NO4 | 0.824 | 717456406 | 0.8241 | 3.906 | 3.7642 | 3.6293
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C9 | 375013162 | 1.953 | C09 | C4H7NO4 | 0.464 | 807365196 | 0.46449 | 1.953 | 2.0434 | 4.6289
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C10 | 196706954 | 0.977 | C10 | C4H7NO4 | 0.231 | 849932404 | 0.23144 | 0.977 | 0.97271 | 0.43928
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C11 | 98909074 | 0.488 | C11 | C4H7NO4 | 0.124 | 797406113 | 0.12404 | 0.488 | 0.50044 | 2.5485
Aspartate neg | 10.26 | Target Compound | KD033022_033022_C12 | 49532914 | 0.244 | C12 | C4H7NO4 | 0.066 | 749596958 | 0.066079 | 0.244 | 0.25582 | 4.8438

## Aspartate U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A1 | 104126550 | 1 | A01 | [13]C4H7[15]NO4 | N/A | N/A | 10.357:Peak area 104126550.119 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A2 | 133394152 | 1 | A02 | [13]C4H7[15]NO4 | N/A | N/A | 10.318:Peak area 133394152.170 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A3 | 183284491 | 1 | A03 | [13]C4H7[15]NO4 | N/A | N/A | 10.290:Peak area 183284490.506 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A4 | 283922238 | 1 | A04 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A5 | 381124511 | 1 | A05 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A6 | 491201220 | 1 | A06 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A7 | 615402697 | 1 | A07 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A8 | 663701678 | 1 | A08 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A9 | 662671558 | 1 | A09 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A10 | 723929472 | 1 | A10 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A11 | 751326187 | 1 | A11 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_A12 | 783189475 | 1 | A12 | [13]C4H7[15]NO4 | N/A | N/A | 10.356:Peak area 783189475.257 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B1 | 107037424 | 1 | B01 | [13]C4H7[15]NO4 | N/A | N/A | 10.363:Peak area 107037423.992 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B2 | 139424184 | 1 | B02 | [13]C4H7[15]NO4 | N/A | N/A | 10.317:Peak area 139424184.469 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B3 | 203280829 | 1 | B03 | [13]C4H7[15]NO4 | N/A | N/A | 10.315:Peak area 203280829.383 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B4 | 277920637 | 1 | B04 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B5 | 381849401 | 1 | B05 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B6 | 505002763 | 1 | B06 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B7 | 661822723 | 1 | B07 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B8 | 753578003 | 1 | B08 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B9 | 751941929 | 1 | B09 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B10 | 812870812 | 1 | B10 | [13]C4H7[15]NO4 | N/A | N/A | 10.348:Peak area 812870812.254 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B11 | 824376630 | 1 | B11 | [13]C4H7[15]NO4 | N/A | N/A | 10.352:Peak area 824376630.388 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_B12 | 854706387 | 1 | B12 | [13]C4H7[15]NO4 | N/A | N/A | 10.362:Peak area 854706387.364 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C1 | 104202136 | 1 | C01 | [13]C4H7[15]NO4 | N/A | N/A | 10.364:Peak area 104202135.896 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C2 | 141599321 | 1 | C02 | [13]C4H7[15]NO4 | N/A | N/A | 10.322:Peak area 141599320.596 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C3 | 194058996 | 1 | C03 | [13]C4H7[15]NO4 | N/A | N/A | 10.321:Peak area 194058995.622 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C4 | 295309099 | 1 | C04 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C5 | 390869748 | 1 | C05 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C6 | 541623092 | 1 | C06 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C7 | 659165756 | 1 | C07 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C8 | 717456406 | 1 | C08 | [13]C4H7[15]NO4 | N/A | N/A
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C9 | 807365196 | 1 | C09 | [13]C4H7[15]NO4 | N/A | N/A | 10.345:Peak area 807365196.121 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C10 | 849932404 | 1 | C10 | [13]C4H7[15]NO4 | N/A | N/A | 10.335:Peak area 849932403.811 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C11 | 797406113 | 1 | C11 | [13]C4H7[15]NO4 | N/A | N/A | 10.372:Peak area 797406112.601 is out of bounds (ISTD Minimum Recovery 254162099.668 and ISTD Max Recovery 762486299.004)
Aspartate U-13C, U-15N neg | 10.26 | Internal Standard | KD033022_033022_C12 | 749596958 | 1 | C12 | [13]C4H7[15]NO4 | N/A | N/A

## Cystine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Cystine pos | 10.6 | Target Compound | KD033022_033022_A1 | 835875962 | 500 | A01 | C6H12N2O4S2 | 515.897 | 1620237 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_A2 | 3556291831 | 250 | A02 | C6H12N2O4S2 | 114.765 | 30987597 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_A3 | 2347139675 | 125 | A03 | C6H12N2O4S2 | 54.197 | 43307530 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_A4 | 1431964348 | 62.5 | A04 | C6H12N2O4S2 | 26.108 | 54848499 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_A5 | 888446478 | 31.25 | A05 | C6H12N2O4S2 | 14.303 | 62117713 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 34.338% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_A6 | 453824751 | 15.625 | A06 | C6H12N2O4S2 | 7.078 | 64121888 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 42.062% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_A7 | 257641530 | 7.813 | A07 | C6H12N2O4S2 | 3.536 | 72864785 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 47.39% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_A8 | 139215537 | 3.906 | A08 | C6H12N2O4S2 | 1.764 | 78929912 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 50.114% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_A9 | 65954644 | 1.953 | A09 | C6H12N2O4S2 | 0.969 | 68089751 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 66.477% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_A10 | 39556732 | 0.977 | A10 | C6H12N2O4S2 | 0.458 | 86412897 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 58.268% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_A11 | 19093411 | 0.488 | A11 | C6H12N2O4S2 | 0.222 | 85901051 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 54.311% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_A12 | 8064209 | 0.244 | A12 | C6H12N2O4S2 | 0.094 | 86120921 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 30.228% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_B1 | 898963938 | 500 | B01 | C6H12N2O4S2 | 371.194 | 2421816 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_B2 | 3215130838 | 250 | B02 | C6H12N2O4S2 | 118.453 | 27142567 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_B3 | 2272011038 | 125 | B03 | C6H12N2O4S2 | 52.773 | 43052648 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_B4 | 1392725392 | 62.5 | B04 | C6H12N2O4S2 | 26.581 | 52396198 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_B5 | 828852785 | 31.25 | B05 | C6H12N2O4S2 | 13.842 | 59880577 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 30.515% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_B6 | 495535429 | 15.625 | B06 | C6H12N2O4S2 | 7.393 | 67027365 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 47.925% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_B7 | 245239352 | 7.813 | B07 | C6H12N2O4S2 | 3.245 | 75581709 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 35.698% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_B8 | 142933110 | 3.906 | B08 | C6H12N2O4S2 | 1.76 | 81220273 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 49.783% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_B9 | 85537408 | 1.953 | B09 | C6H12N2O4S2 | 0.933 | 91676242 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 60.428% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_B10 | 44714432 | 0.977 | B10 | C6H12N2O4S2 | 0.459 | 97358324 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 58.788% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_B11 | 23910520 | 0.488 | B11 | C6H12N2O4S2 | 0.233 | 102765602 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 61.509% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_B12 | 10221884 | 0.244 | B12 | C6H12N2O4S2 | 0.106 | 96832247 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 46.79% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_C1 | 707945301 | 500 | C01 | C6H12N2O4S2 | 413.85 | 1710634 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_C2 | 2167932042 | 250 | C02 | C6H12N2O4S2 | 127.236 | 17038613 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_C3 | 2431418184 | 125 | C03 | C6H12N2O4S2 | 55.642 | 43697710 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_C4 | 1497371367 | 62.5 | C04 | C6H12N2O4S2 | 26.498 | 56508924 | Rsquared 0.977 < R2Threshold 0.99
Cystine pos | 10.6 | Target Compound | KD033022_033022_C5 | 924281387 | 31.25 | C05 | C6H12N2O4S2 | 13.821 | 66876652 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 30.339% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_C6 | 534630392 | 15.625 | C06 | C6H12N2O4S2 | 6.992 | 76465701 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 40.462% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_C7 | 288189096 | 7.813 | C07 | C6H12N2O4S2 | 3.195 | 90207079 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 33.685% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_C8 | 146363089 | 3.906 | C08 | C6H12N2O4S2 | 1.745 | 83889145 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 48.525% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_C9 | 93363862 | 1.953 | C09 | C6H12N2O4S2 | 0.97 | 96238436 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 66.729% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_C10 | 48340269 | 0.977 | C10 | C6H12N2O4S2 | 0.462 | 104726876 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 59.581% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_C11 | 20785037 | 0.488 | C11 | C6H12N2O4S2 | 0.226 | 91839633 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 57.112% (absolute) > max calibration amount diff 20%
Cystine pos | 10.6 | Target Compound | KD033022_033022_C12 | 11296664 | 0.244 | C12 | C6H12N2O4S2 | 0.108 | 104970168 | Rsquared 0.977 < R2Threshold 0.99. Calibration Amount Diff 49.643% (absolute) > max calibration amount diff 20%

## Cystine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A1 | 1620237 | 1 | A01 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.704:Peak area 1620236.747 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A2 | 30987597 | 1 | A02 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.691:Peak area 30987597.014 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A3 | 43307530 | 1 | A03 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A4 | 54848499 | 1 | A04 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A5 | 62117713 | 1 | A05 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A6 | 64121888 | 1 | A06 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A7 | 72864785 | 1 | A07 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A8 | 78929912 | 1 | A08 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A9 | 68089751 | 1 | A09 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A10 | 86412897 | 1 | A10 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A11 | 85901051 | 1 | A11 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_A12 | 86120921 | 1 | A12 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B1 | 2421816 | 1 | B01 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.711:Peak area 2421815.505 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B2 | 27142567 | 1 | B02 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.718:Peak area 27142567.083 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B3 | 43052648 | 1 | B03 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B4 | 52396198 | 1 | B04 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B5 | 59880577 | 1 | B05 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B6 | 67027365 | 1 | B06 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B7 | 75581709 | 1 | B07 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B8 | 81220273 | 1 | B08 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B9 | 91676242 | 1 | B09 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B10 | 97358324 | 1 | B10 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B11 | 102765602 | 1 | B11 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.685:Peak area 102765601.609 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_B12 | 96832247 | 1 | B12 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C1 | 1710634 | 1 | C01 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.740:Peak area 1710634.153 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C2 | 17038613 | 1 | C02 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.695:Peak area 17038612.959 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C3 | 43697710 | 1 | C03 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C4 | 56508924 | 1 | C04 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C5 | 66876652 | 1 | C05 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C6 | 76465701 | 1 | C06 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C7 | 90207079 | 1 | C07 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C8 | 83889145 | 1 | C08 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C9 | 96238436 | 1 | C09 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C10 | 104726876 | 1 | C10 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.666:Peak area 104726876.350 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C11 | 91839633 | 1 | C11 | [13]C6H12[15]N2O4S2 | N/A | N/A
Cystine U-13C, U-15N pos | 10.6 | Internal Standard | KD033022_033022_C12 | 104970168 | 1 | C12 | [13]C6H12[15]N2O4S2 | N/A | N/A | 10.695:Peak area 104970168.091 is out of bounds (ISTD Minimum Recovery 32872887.781 and ISTD Max Recovery 98618663.342)

## Glutamate neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | RR | Theoretical Amt | power pred | %diff
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A1 | 8407582638 | 500 | A01 | C5H9NO4 | 93.337 | 90077742 | 93.337 | 500 | 574.2605 | 14.8521
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A2 | 5501732155 | 250 | A02 | C5H9NO4 | 40.436 | 136060746 | 40.4359 | 250 | 237.3778 | 5.0489
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A3 | 3508930423 | 125 | A03 | C5H9NO4 | 20.738 | 169206914 | 20.7375 | 125 | 117.2626 | 6.19
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A4 | 2618500497 | 62.5 | A04 | C5H9NO4 | 11.149 | 234865269 | 11.1489 | 62.5 | 60.8857 | 2.5829
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A5 | 1848563746 | 31.25 | A05 | C5H9NO4 | 6.15 | 300587128 | 6.1498 | 31.25 | 32.4825 | 3.9441
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A6 | 1165877847 | 15.625 | A06 | C5H9NO4 | 3.096 | 376623972 | 3.0956 | 15.625 | 15.7328 | 0.68974
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A7 | 778514344 | 7.813 | A07 | C5H9NO4 | 1.571 | 495431013 | 1.5714 | 7.813 | 7.6882 | 1.5977
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A8 | 450395350 | 3.906 | A08 | C5H9NO4 | 0.804 | 559846486 | 0.8045 | 3.906 | 3.791 | 2.9446
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A9 | 265213051 | 1.953 | A09 | C5H9NO4 | 0.449 | 591300904 | 0.44852 | 1.953 | 2.0454 | 4.731
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A10 | 144991249 | 0.977 | A10 | C5H9NO4 | 0.221 | 654851431 | 0.22141 | 0.977 | 0.97049 | 0.6667
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A11 | 84098213 | 0.488 | A11 | C5H9NO4 | 0.12 | 703332613 | 0.11957 | 0.488 | 0.5063 | 3.7492
Glutamate neg | 10.05 | Target Compound | KD033022_033022_A12 | 42534923 | 0.244 | A12 | C5H9NO4 | 0.06 | 705769398 | 0.060267 | 0.244 | 0.24557 | 0.64175
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B1 | 8377924193 | 500 | B01 | C5H9NO4 | 90.663 | 92407128 | 90.6632 | 500 | 556.9011 | 11.3802
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B2 | 5248445432 | 250 | B02 | C5H9NO4 | 40.818 | 128581866 | 40.8179 | 250 | 239.7472 | 4.1011
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B3 | 3619826616 | 125 | B03 | C5H9NO4 | 20.765 | 174323655 | 20.765 | 125 | 117.4266 | 6.0587
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B4 | 2536341060 | 62.5 | B04 | C5H9NO4 | 11.046 | 229626222 | 11.0455 | 62.5 | 60.2893 | 3.5371
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B5 | 1757718016 | 31.25 | B05 | C5H9NO4 | 5.9 | 297900413 | 5.9004 | 31.25 | 31.0924 | 0.50418
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B6 | 1266966905 | 15.625 | B06 | C5H9NO4 | 3.257 | 388975417 | 3.2572 | 15.625 | 16.6013 | 6.2486
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B7 | 770365289 | 7.813 | B07 | C5H9NO4 | 1.461 | 527227815 | 1.4612 | 7.813 | 7.1198 | 8.8728
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B8 | 500822062 | 3.906 | B08 | C5H9NO4 | 0.809 | 618853138 | 0.80927 | 3.906 | 3.8148 | 2.3359
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B9 | 283029435 | 1.953 | B09 | C5H9NO4 | 0.429 | 660300682 | 0.42864 | 1.953 | 1.9497 | 0.16708
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B10 | 159686935 | 0.977 | B10 | C5H9NO4 | 0.222 | 719411342 | 0.22197 | 0.977 | 0.97307 | 0.4023
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B11 | 85733720 | 0.488 | B11 | C5H9NO4 | 0.117 | 729674397 | 0.1175 | 0.488 | 0.49702 | 1.8485
Glutamate neg | 10.05 | Target Compound | KD033022_033022_B12 | 45669546 | 0.244 | B12 | C5H9NO4 | 0.06 | 759213321 | 0.060154 | 0.244 | 0.24508 | 0.44127
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C1 | 8406386373 | 500 | C01 | C5H9NO4 | 89.632 | 93788228 | 89.6316 | 500 | 550.211 | 10.0422
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C2 | 5227162666 | 250 | C02 | C5H9NO4 | 40.349 | 129547480 | 40.3494 | 250 | 236.8418 | 5.2633
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C3 | 3660744094 | 125 | C03 | C5H9NO4 | 21.461 | 170575729 | 21.4611 | 125 | 121.588 | 2.7296
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C4 | 2677911732 | 62.5 | C04 | C5H9NO4 | 11.043 | 242508066 | 11.0426 | 62.5 | 60.2723 | 3.5643
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C5 | 1876589385 | 31.25 | C05 | C5H9NO4 | 5.988 | 313381638 | 5.9882 | 31.25 | 31.5815 | 1.0607
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C6 | 1299817728 | 15.625 | C06 | C5H9NO4 | 3.065 | 424089764 | 3.065 | 15.625 | 15.5683 | 0.3626
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C7 | 779434444 | 7.813 | C07 | C5H9NO4 | 1.467 | 531340936 | 1.4669 | 7.813 | 7.1494 | 8.4935
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C8 | 465724520 | 3.906 | C08 | C5H9NO4 | 0.808 | 576731938 | 0.80752 | 3.906 | 3.806 | 2.5591
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C9 | 314505040 | 1.953 | C09 | C5H9NO4 | 0.452 | 695055991 | 0.45249 | 1.953 | 2.0645 | 5.7088
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C10 | 165184187 | 0.977 | C10 | C5H9NO4 | 0.225 | 733283802 | 0.22527 | 0.977 | 0.98834 | 1.1609
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C11 | 85447263 | 0.488 | C11 | C5H9NO4 | 0.117 | 730523366 | 0.11697 | 0.488 | 0.49466 | 1.3646
Glutamate neg | 10.05 | Target Compound | KD033022_033022_C12 | 42195793 | 0.244 | C12 | C5H9NO4 | 0.063 | 673619680 | 0.06264 | 0.244 | 0.25579 | 4.8312

## Glutamate U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A1 | 90077742 | 1 | A01 | [13]C5H9[15]NO4 | N/A | N/A | 10.163:Peak area 90077741.979 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A2 | 136060746 | 1 | A02 | [13]C5H9[15]NO4 | N/A | N/A | 10.124:Peak area 136060745.542 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A3 | 169206914 | 1 | A03 | [13]C5H9[15]NO4 | N/A | N/A | 10.124:Peak area 169206913.872 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A4 | 234865269 | 1 | A04 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A5 | 300587128 | 1 | A05 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A6 | 376623972 | 1 | A06 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A7 | 495431013 | 1 | A07 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A8 | 559846486 | 1 | A08 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A9 | 591300904 | 1 | A09 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A10 | 654851431 | 1 | A10 | [13]C5H9[15]NO4 | N/A | N/A | 10.090:Peak area 654851431.259 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A11 | 703332613 | 1 | A11 | [13]C5H9[15]NO4 | N/A | N/A | 10.100:Peak area 703332613.142 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_A12 | 705769398 | 1 | A12 | [13]C5H9[15]NO4 | N/A | N/A | 10.129:Peak area 705769398.053 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B1 | 92407128 | 1 | B01 | [13]C5H9[15]NO4 | N/A | N/A | 10.169:Peak area 92407127.509 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B2 | 128581866 | 1 | B02 | [13]C5H9[15]NO4 | N/A | N/A | 10.152:Peak area 128581865.646 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B3 | 174323655 | 1 | B03 | [13]C5H9[15]NO4 | N/A | N/A | 10.149:Peak area 174323655.257 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B4 | 229626222 | 1 | B04 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B5 | 297900413 | 1 | B05 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B6 | 388975417 | 1 | B06 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B7 | 527227815 | 1 | B07 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B8 | 618853138 | 1 | B08 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B9 | 660300682 | 1 | B09 | [13]C5H9[15]NO4 | N/A | N/A | 10.105:Peak area 660300681.506 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B10 | 719411342 | 1 | B10 | [13]C5H9[15]NO4 | N/A | N/A | 10.119:Peak area 719411341.819 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B11 | 729674397 | 1 | B11 | [13]C5H9[15]NO4 | N/A | N/A | 10.125:Peak area 729674397.125 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_B12 | 759213321 | 1 | B12 | [13]C5H9[15]NO4 | N/A | N/A | 10.133:Peak area 759213320.884 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C1 | 93788228 | 1 | C01 | [13]C5H9[15]NO4 | N/A | N/A | 10.171:Peak area 93788228.377 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C2 | 129547480 | 1 | C02 | [13]C5H9[15]NO4 | N/A | N/A | 10.156:Peak area 129547480.156 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C3 | 170575729 | 1 | C03 | [13]C5H9[15]NO4 | N/A | N/A | 10.155:Peak area 170575729.371 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C4 | 242508066 | 1 | C04 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C5 | 313381638 | 1 | C05 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C6 | 424089764 | 1 | C06 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C7 | 531340936 | 1 | C07 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C8 | 576731938 | 1 | C08 | [13]C5H9[15]NO4 | N/A | N/A
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C9 | 695055991 | 1 | C09 | [13]C5H9[15]NO4 | N/A | N/A | 10.117:Peak area 695055991.386 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C10 | 733283802 | 1 | C10 | [13]C5H9[15]NO4 | N/A | N/A | 10.106:Peak area 733283801.783 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C11 | 730523366 | 1 | C11 | [13]C5H9[15]NO4 | N/A | N/A | 10.144:Peak area 730523366.237 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)
Glutamate U-13C, U-15N neg | 10.05 | Internal Standard | KD033022_033022_C12 | 673619680 | 1 | C12 | [13]C5H9[15]NO4 | N/A | N/A | 10.131:Peak area 673619679.955 is out of bounds (ISTD Minimum Recovery 217484661.521 and ISTD Max Recovery 652453984.563)

## Glutamine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A1 | 42681684105 | 500 | A01 | C5H10N2O3 | 169.383 | 251982913 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A2 | 32134289145 | 250 | A02 | C5H10N2O3 | 69.02 | 465581917 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A3 | 23236361528 | 125 | A03 | C5H10N2O3 | 36.234 | 641282928 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A4 | 16909682150 | 62.5 | A04 | C5H10N2O3 | 18.788 | 900018895 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A5 | 12648120513 | 31.25 | A05 | C5H10N2O3 | 10.448 | 1210535528 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 25.5% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A6 | 8111521517 | 15.625 | A06 | C5H10N2O3 | 5.27 | 1539329967 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 27.126% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A7 | 5079245028 | 7.813 | A07 | C5H10N2O3 | 2.696 | 1884330062 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 30.326% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A8 | 3179435836 | 3.906 | A08 | C5H10N2O3 | 1.355 | 2347090996 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 31.153% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A9 | 1882831276 | 1.953 | A09 | C5H10N2O3 | 0.772 | 2438035548 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 49.613% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A10 | 1022261953 | 0.977 | A10 | C5H10N2O3 | 0.375 | 2726617005 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 45.241% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A11 | 558997675 | 0.488 | A11 | C5H10N2O3 | 0.198 | 2816299824 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 53.964% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_A12 | 302952452 | 0.244 | A12 | C5H10N2O3 | 0.099 | 3072018159 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 53.005% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B1 | 33794688980 | 500 | B01 | C5H10N2O3 | 114.908 | 294102736 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff -20.033% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B2 | 24790581867 | 250 | B02 | C5H10N2O3 | 51.127 | 484882695 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff -25.619% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B3 | 18721461825 | 125 | B03 | C5H10N2O3 | 26.079 | 717887364 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff -22.659% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B4 | 13670220032 | 62.5 | B04 | C5H10N2O3 | 13.612 | 1004284167 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B5 | 9546345153 | 31.25 | B05 | C5H10N2O3 | 7.251 | 1316614747 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B6 | 6672562183 | 15.625 | B06 | C5H10N2O3 | 3.933 | 1696588725 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B7 | 3775335131 | 7.813 | B07 | C5H10N2O3 | 1.792 | 2106405906 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B8 | 2443113798 | 3.906 | B08 | C5H10N2O3 | 0.985 | 2480384103 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B9 | 1482809625 | 1.953 | B09 | C5H10N2O3 | 0.536 | 2764074307 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B10 | 757341815 | 0.977 | B10 | C5H10N2O3 | 0.269 | 2814837499 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B11 | 431156313 | 0.488 | B11 | C5H10N2O3 | 0.14 | 3071945771 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_B12 | 222174670 | 0.244 | B12 | C5H10N2O3 | 0.072 | 3099201833 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C1 | 41800050927 | 500 | C01 | C5H10N2O3 | 163.522 | 255623499 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C2 | 31271098671 | 250 | C02 | C5H10N2O3 | 72.596 | 430753961 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C3 | 23642711713 | 125 | C03 | C5H10N2O3 | 37.998 | 622205913 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C4 | 17635736902 | 62.5 | C04 | C5H10N2O3 | 19.506 | 904122387 | Rsquared 0.972 < R2Threshold 0.99
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C5 | 12585593049 | 31.25 | C05 | C5H10N2O3 | 10.53 | 1195264728 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 26.467% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C6 | 8373638800 | 15.625 | C06 | C5H10N2O3 | 5.442 | 1538657869 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 31.273% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C7 | 5096163434 | 7.813 | C07 | C5H10N2O3 | 2.563 | 1988069252 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 23.951% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C8 | 3079959900 | 3.906 | C08 | C5H10N2O3 | 1.389 | 2217965962 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 34.442% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C9 | 2115631462 | 1.953 | C09 | C5H10N2O3 | 0.787 | 2688474654 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 52.45% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C10 | 1108325498 | 0.977 | C10 | C5H10N2O3 | 0.39 | 2844763564 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 50.927% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C11 | 600833106 | 0.488 | C11 | C5H10N2O3 | 0.203 | 2953873290 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 57.779% (absolute) > max calibration amount diff 20%
Glutamine pos | 9.64 | Target Compound | KD033022_033022_C12 | 304182101 | 0.244 | C12 | C5H10N2O3 | 0.103 | 2966206714 | Rsquared 0.972 < R2Threshold 0.99. Calibration Amount Diff 59.106% (absolute) > max calibration amount diff 20%

## Glutamine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A1 | 251982913 | 1 | A01 | [13]C5H10[15]N2O3 | N/A | N/A | 9.734:Peak area 251982912.593 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A2 | 465581917 | 1 | A02 | [13]C5H10[15]N2O3 | N/A | N/A | 9.725:Peak area 465581916.957 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A3 | 641282928 | 1 | A03 | [13]C5H10[15]N2O3 | N/A | N/A | 9.696:Peak area 641282928.463 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A4 | 900018895 | 1 | A04 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A5 | 1210535528 | 1 | A05 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A6 | 1539329967 | 1 | A06 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A7 | 1884330062 | 1 | A07 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A8 | 2347090996 | 1 | A08 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A9 | 2438035548 | 1 | A09 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A10 | 2726617005 | 1 | A10 | [13]C5H10[15]N2O3 | N/A | N/A | 9.709:Peak area 2726617004.715 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A11 | 2816299824 | 1 | A11 | [13]C5H10[15]N2O3 | N/A | N/A | 9.719:Peak area 2816299823.949 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_A12 | 3072018159 | 1 | A12 | [13]C5H10[15]N2O3 | N/A | N/A | 9.718:Peak area 3072018158.954 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B1 | 294102736 | 1 | B01 | [13]C5H10[15]N2O3 | N/A | N/A | 9.740:Peak area 294102736.351 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B2 | 484882695 | 1 | B02 | [13]C5H10[15]N2O3 | N/A | N/A | 9.724:Peak area 484882695.052 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B3 | 717887364 | 1 | B03 | [13]C5H10[15]N2O3 | N/A | N/A | 9.721:Peak area 717887364.267 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B4 | 1004284167 | 1 | B04 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B5 | 1316614747 | 1 | B05 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B6 | 1696588725 | 1 | B06 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B7 | 2106405906 | 1 | B07 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B8 | 2480384103 | 1 | B08 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B9 | 2764074307 | 1 | B09 | [13]C5H10[15]N2O3 | N/A | N/A | 9.725:Peak area 2764074307.006 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B10 | 2814837499 | 1 | B10 | [13]C5H10[15]N2O3 | N/A | N/A | 9.710:Peak area 2814837499.067 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B11 | 3071945771 | 1 | B11 | [13]C5H10[15]N2O3 | N/A | N/A | 9.715:Peak area 3071945770.605 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_B12 | 3099201833 | 1 | B12 | [13]C5H10[15]N2O3 | N/A | N/A | 9.724:Peak area 3099201833.016 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C1 | 255623499 | 1 | C01 | [13]C5H10[15]N2O3 | N/A | N/A | 9.714:Peak area 255623499.092 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C2 | 430753961 | 1 | C02 | [13]C5H10[15]N2O3 | N/A | N/A | 9.729:Peak area 430753960.943 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C3 | 622205913 | 1 | C03 | [13]C5H10[15]N2O3 | N/A | N/A | 9.727:Peak area 622205913.125 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C4 | 904122387 | 1 | C04 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C5 | 1195264728 | 1 | C05 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C6 | 1538657869 | 1 | C06 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C7 | 1988069252 | 1 | C07 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C8 | 2217965962 | 1 | C08 | [13]C5H10[15]N2O3 | N/A | N/A
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C9 | 2688474654 | 1 | C09 | [13]C5H10[15]N2O3 | N/A | N/A | 9.709:Peak area 2688474653.725 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C10 | 2844763564 | 1 | C10 | [13]C5H10[15]N2O3 | N/A | N/A | 9.725:Peak area 2844763563.655 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C11 | 2953873290 | 1 | C11 | [13]C5H10[15]N2O3 | N/A | N/A | 9.733:Peak area 2953873290.404 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)
Glutamine U-13C, U-15N pos | 9.64 | Internal Standard | KD033022_033022_C12 | 2966206714 | 1 | C12 | [13]C5H10[15]N2O3 | N/A | N/A | 9.719:Peak area 2966206714.140 is out of bounds (ISTD Minimum Recovery 871532158.165 and ISTD Max Recovery 2614596474.495)

## Glycine neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Glycine neg | 10.11 | Target Compound | KD033022_033022_A1 | 794228190 | 500 | A01 | C2H5NO2 | 142.458 | 5575180
Glycine neg | 10.11 | Target Compound | KD033022_033022_A2 | 380949299 | 250 | A02 | C2H5NO2 | 64.797 | 5879083
Glycine neg | 10.11 | Target Compound | KD033022_033022_A3 | 201995122 | 125 | A03 | C2H5NO2 | 28.636 | 7053887
Glycine neg | 10.11 | Target Compound | KD033022_033022_A4 | 103317005 | 62.5 | A04 | C2H5NO2 | 15.128 | 6829306
Glycine neg | 10.11 | Target Compound | KD033022_033022_A5 | 52896034 | 31.25 | A05 | C2H5NO2 | 7.159 | 7388282
Glycine neg | 10.11 | Target Compound | KD033022_033022_A6 | 27525034 | 15.625 | A06 | C2H5NO2 | 3.394 | 8109812
Glycine neg | 10.11 | Target Compound | KD033022_033022_A7 | 11948630 | 7.813 | A07 | C2H5NO2 | 1.557 | 7676542
Glycine neg | 10.11 | Target Compound | KD033022_033022_A8 | 5660492 | 3.906 | A08 | C2H5NO2 | 0.718 | 7882809
Glycine neg | 10.11 | Target Compound | KD033022_033022_A9 | 3114560 | 1.953 | A09 | C2H5NO2 | 0.43 | 7235280
Glycine neg | 10.11 | Target Compound | KD033022_033022_A10 | 1321034 | 0.977 | A10 | C2H5NO2 | 0.172 | 7658739 | Calibration Amount Diff -22.668% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD033022_033022_A11 | 1003382 | 0.488 | A11 | C2H5NO2 | 0.126 | 7978561
Glycine neg | 10.11 | Target Compound | KD033022_033022_A12 | 366189 | 0.244 | A12 | C2H5NO2 | 0.05 | 7323014
Glycine neg | 10.11 | Target Compound | KD033022_033022_B1 | 733380656 | 500 | B01 | C2H5NO2 | 120.4 | 6091194
Glycine neg | 10.11 | Target Compound | KD033022_033022_B2 | 345115687 | 250 | B02 | C2H5NO2 | 57.404 | 6012046
Glycine neg | 10.11 | Target Compound | KD033022_033022_B3 | 178519857 | 125 | B03 | C2H5NO2 | 31.258 | 5711231
Glycine neg | 10.11 | Target Compound | KD033022_033022_B4 | 91923641 | 62.5 | B04 | C2H5NO2 | 14.282 | 6436185
Glycine neg | 10.11 | Target Compound | KD033022_033022_B5 | 45401304 | 31.25 | B05 | C2H5NO2 | 7.891 | 5753331
Glycine neg | 10.11 | Target Compound | KD033022_033022_B6 | 26064334 | 15.625 | B06 | C2H5NO2 | 3.884 | 6710470
Glycine neg | 10.11 | Target Compound | KD033022_033022_B7 | 10764394 | 7.813 | B07 | C2H5NO2 | 1.579 | 6819202
Glycine neg | 10.11 | Target Compound | KD033022_033022_B8 | 5113171 | 3.906 | B08 | C2H5NO2 | 0.701 | 7295919 | Calibration Amount Diff -21.486% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD033022_033022_B9 | 2773568 | 1.953 | B09 | C2H5NO2 | 0.386 | 7194328
Glycine neg | 10.11 | Target Compound | KD033022_033022_B10 | 1289749 | 0.977 | B10 | C2H5NO2 | 0.188 | 6874239
Glycine neg | 10.11 | Target Compound | KD033022_033022_B11 | 1048328 | 0.488 | B11 | C2H5NO2 | 0.143 | 7311811 | Calibration Amount Diff 28.697% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD033022_033022_B12 | N/F | 0.244 | B12 | C2H5NO2 | N/F | 7249842
Glycine neg | 10.11 | Target Compound | KD033022_033022_C1 | 703915455 | 500 | C01 | C2H5NO2 | 154.724 | 4549487
Glycine neg | 10.11 | Target Compound | KD033022_033022_C2 | 315124994 | 250 | C02 | C2H5NO2 | 61.205 | 5148653
Glycine neg | 10.11 | Target Compound | KD033022_033022_C3 | 168237021 | 125 | C03 | C2H5NO2 | 33.916 | 4960398
Glycine neg | 10.11 | Target Compound | KD033022_033022_C4 | 85521746 | 62.5 | C04 | C2H5NO2 | 15.872 | 5388232
Glycine neg | 10.11 | Target Compound | KD033022_033022_C5 | 45179658 | 31.25 | C05 | C2H5NO2 | 7.202 | 6273317
Glycine neg | 10.11 | Target Compound | KD033022_033022_C6 | 22389444 | 15.625 | C06 | C2H5NO2 | 3.536 | 6331759
Glycine neg | 10.11 | Target Compound | KD033022_033022_C7 | 9180235 | 7.813 | C07 | C2H5NO2 | 1.436 | 6392208
Glycine neg | 10.11 | Target Compound | KD033022_033022_C8 | 4943646 | 3.906 | C08 | C2H5NO2 | 0.779 | 6342399
Glycine neg | 10.11 | Target Compound | KD033022_033022_C9 | 2657447 | 1.953 | C09 | C2H5NO2 | 0.373 | 7121510
Glycine neg | 10.11 | Target Compound | KD033022_033022_C10 | 1128676 | 0.977 | C10 | C2H5NO2 | 0.157 | 7183164 | Calibration Amount Diff -29.552% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD033022_033022_C11 | 1151325 | 0.488 | C11 | C2H5NO2 | 0.153 | 7529759 | Calibration Amount Diff 37.247% (absolute) > max calibration amount diff 20%
Glycine neg | 10.11 | Target Compound | KD033022_033022_C12 | 570325 | 0.244 | C12 | C2H5NO2 | 0.091 | 6273853 | Calibration Amount Diff 63.214% (absolute) > max calibration amount diff 20%

## Glycine U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A1 | 5575180 | 1 | A01 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A2 | 5879083 | 1 | A02 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A3 | 7053887 | 1 | A03 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A4 | 6829306 | 1 | A04 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A5 | 7388282 | 1 | A05 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A6 | 8109812 | 1 | A06 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A7 | 7676542 | 1 | A07 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A8 | 7882809 | 1 | A08 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A9 | 7235280 | 1 | A09 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A10 | 7658739 | 1 | A10 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A11 | 7978561 | 1 | A11 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_A12 | 7323014 | 1 | A12 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B1 | 6091194 | 1 | B01 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B2 | 6012046 | 1 | B02 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B3 | 5711231 | 1 | B03 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B4 | 6436185 | 1 | B04 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B5 | 5753331 | 1 | B05 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B6 | 6710470 | 1 | B06 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B7 | 6819202 | 1 | B07 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B8 | 7295919 | 1 | B08 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B9 | 7194328 | 1 | B09 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B10 | 6874239 | 1 | B10 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B11 | 7311811 | 1 | B11 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_B12 | 7249842 | 1 | B12 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C1 | 4549487 | 1 | C01 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C2 | 5148653 | 1 | C02 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C3 | 4960398 | 1 | C03 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C4 | 5388232 | 1 | C04 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C5 | 6273317 | 1 | C05 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C6 | 6331759 | 1 | C06 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C7 | 6392208 | 1 | C07 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C8 | 6342399 | 1 | C08 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C9 | 7121510 | 1 | C09 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C10 | 7183164 | 1 | C10 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C11 | 7529759 | 1 | C11 | [13]C2H5[15]NO2 | N/A | N/A
Glycine U-13C, U-15N neg | 10.11 | Internal Standard | KD033022_033022_C12 | 6273853 | 1 | C12 | [13]C2H5[15]NO2 | N/A | N/A

## Isoleucine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A1 | 61295265990 | 500 | A01 | C6H13NO2 | 91.454 | 670233825
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A2 | 38460284602 | 250 | A02 | C6H13NO2 | 37.394 | 1028523018
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A3 | 26058674653 | 125 | A03 | C6H13NO2 | 20.407 | 1276966003
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A4 | 16075556608 | 62.5 | A04 | C6H13NO2 | 10.339 | 1554854272
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A5 | 9542782358 | 31.25 | A05 | C6H13NO2 | 5.688 | 1677636771
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A6 | 4991901987 | 15.625 | A06 | C6H13NO2 | 2.917 | 1711090947
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A7 | 2732719617 | 7.813 | A07 | C6H13NO2 | 1.493 | 1830643798
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A8 | 1469501897 | 3.906 | A08 | C6H13NO2 | 0.751 | 1956991312 | Calibration Amount Diff 20.293% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A9 | 812210560 | 1.953 | A09 | C6H13NO2 | 0.425 | 1910144076 | Calibration Amount Diff 36.272% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A10 | 402691125 | 0.977 | A10 | C6H13NO2 | 0.208 | 1932430355 | Calibration Amount Diff 33.523% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A11 | 226131990 | 0.488 | A11 | C6H13NO2 | 0.111 | 2043279292 | Calibration Amount Diff 41.982% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_A12 | 117559787 | 0.244 | A12 | C6H13NO2 | 0.056 | 2089005350 | Calibration Amount Diff 44.4% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B1 | 60312226907 | 500 | B01 | C6H13NO2 | 83.381 | 723336367
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B2 | 40365308641 | 250 | B02 | C6H13NO2 | 37.196 | 1085214141
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B3 | 26675831053 | 125 | B03 | C6H13NO2 | 20.457 | 1304003358
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B4 | 15604865118 | 62.5 | B04 | C6H13NO2 | 10.418 | 1497900795
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B5 | 8997154157 | 31.25 | B05 | C6H13NO2 | 5.488 | 1639454130
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B6 | 5322298064 | 15.625 | B06 | C6H13NO2 | 3.145 | 1692334158 | Calibration Amount Diff 25.702% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B7 | 2606060929 | 7.813 | B07 | C6H13NO2 | 1.434 | 1817579819
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B8 | 1471245858 | 3.906 | B08 | C6H13NO2 | 0.759 | 1939480760 | Calibration Amount Diff 21.522% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B9 | 843386831 | 1.953 | B09 | C6H13NO2 | 0.431 | 1957062268 | Calibration Amount Diff 38.11% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B10 | 421097807 | 0.977 | B10 | C6H13NO2 | 0.209 | 2015122941 | Calibration Amount Diff 33.897% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B11 | 230992622 | 0.488 | B11 | C6H13NO2 | 0.109 | 2118159989 | Calibration Amount Diff 39.907% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_B12 | 122092632 | 0.244 | B12 | C6H13NO2 | 0.055 | 2209195357 | Calibration Amount Diff 41.809% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C1 | 59488487113 | 500 | C01 | C6H13NO2 | 85.701 | 694139476
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C2 | 38009595528 | 250 | C02 | C6H13NO2 | 37.739 | 1007177147
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C3 | 25443728686 | 125 | C03 | C6H13NO2 | 19.956 | 1274984760
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C4 | 14855538973 | 62.5 | C04 | C6H13NO2 | 10.314 | 1440333699
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C5 | 8598683359 | 31.25 | C05 | C6H13NO2 | 5.897 | 1458237534
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C6 | 4846973450 | 15.625 | C06 | C6H13NO2 | 2.908 | 1666588704
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C7 | 2472972218 | 7.813 | C07 | C6H13NO2 | 1.397 | 1770740242
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C8 | 1374336401 | 3.906 | C08 | C6H13NO2 | 0.75 | 1832341729 | Calibration Amount Diff 20.156% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C9 | 847988129 | 1.953 | C09 | C6H13NO2 | 0.427 | 1987362364 | Calibration Amount Diff 36.747% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C10 | 430417189 | 0.977 | C10 | C6H13NO2 | 0.212 | 2034179329 | Calibration Amount Diff 35.578% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C11 | 234482696 | 0.488 | C11 | C6H13NO2 | 0.111 | 2105983418 | Calibration Amount Diff 42.841% (absolute) > max calibration amount diff 20%
Isoleucine pos | 7.05 | Target Compound | KD033022_033022_C12 | 112413833 | 0.244 | C12 | C6H13NO2 | 0.056 | 1996550938 | Calibration Amount Diff 44.473% (absolute) > max calibration amount diff 20%

## Isoleucine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A1 | 670233825 | 1 | A01 | [13]C6H13[15]NO2 | N/A | N/A | 7.041:Peak area 670233825.010 is out of bounds (ISTD Minimum Recovery 818739756.123 and ISTD Max Recovery 2456219268.369)
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A2 | 1028523018 | 1 | A02 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A3 | 1276966003 | 1 | A03 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A4 | 1554854272 | 1 | A04 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A5 | 1677636771 | 1 | A05 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A6 | 1711090947 | 1 | A06 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A7 | 1830643798 | 1 | A07 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A8 | 1956991312 | 1 | A08 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A9 | 1910144076 | 1 | A09 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A10 | 1932430355 | 1 | A10 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A11 | 2043279292 | 1 | A11 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_A12 | 2089005350 | 1 | A12 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B1 | 723336367 | 1 | B01 | [13]C6H13[15]NO2 | N/A | N/A | 7.047:Peak area 723336367.357 is out of bounds (ISTD Minimum Recovery 818739756.123 and ISTD Max Recovery 2456219268.369)
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B2 | 1085214141 | 1 | B02 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B3 | 1304003358 | 1 | B03 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B4 | 1497900795 | 1 | B04 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B5 | 1639454130 | 1 | B05 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B6 | 1692334158 | 1 | B06 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B7 | 1817579819 | 1 | B07 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B8 | 1939480760 | 1 | B08 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B9 | 1957062268 | 1 | B09 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B10 | 2015122941 | 1 | B10 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B11 | 2118159989 | 1 | B11 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_B12 | 2209195357 | 1 | B12 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C1 | 694139476 | 1 | C01 | [13]C6H13[15]NO2 | N/A | N/A | 7.048:Peak area 694139476.088 is out of bounds (ISTD Minimum Recovery 818739756.123 and ISTD Max Recovery 2456219268.369)
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C2 | 1007177147 | 1 | C02 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C3 | 1274984760 | 1 | C03 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C4 | 1440333699 | 1 | C04 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C5 | 1458237534 | 1 | C05 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C6 | 1666588704 | 1 | C06 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C7 | 1770740242 | 1 | C07 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C8 | 1832341729 | 1 | C08 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C9 | 1987362364 | 1 | C09 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C10 | 2034179329 | 1 | C10 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C11 | 2105983418 | 1 | C11 | [13]C6H13[15]NO2 | N/A | N/A
Isoleucine U-13C, U-15N pos | 7.05 | Internal Standard | KD033022_033022_C12 | 1996550938 | 1 | C12 | [13]C6H13[15]NO2 | N/A | N/A

## Leucine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Leucine pos | 6.74 | Target Compound | KD033022_033022_A1 | 60144419733 | 500 | A01 | C6H13NO2 | 81.724 | 735944481
Leucine pos | 6.74 | Target Compound | KD033022_033022_A2 | 40469433850 | 250 | A02 | C6H13NO2 | 36.917 | 1096220175
Leucine pos | 6.74 | Target Compound | KD033022_033022_A3 | 24686034680 | 125 | A03 | C6H13NO2 | 18.779 | 1314581659
Leucine pos | 6.74 | Target Compound | KD033022_033022_A4 | 14526203215 | 62.5 | A04 | C6H13NO2 | 10.465 | 1388113227
Leucine pos | 6.74 | Target Compound | KD033022_033022_A5 | 8591238178 | 31.25 | A05 | C6H13NO2 | 5.854 | 1467701261
Leucine pos | 6.74 | Target Compound | KD033022_033022_A6 | 4337036186 | 15.625 | A06 | C6H13NO2 | 2.983 | 1453679002
Leucine pos | 6.74 | Target Compound | KD033022_033022_A7 | 2416286699 | 7.813 | A07 | C6H13NO2 | 1.525 | 1583943317 | Calibration Amount Diff 21.119% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_A8 | 1306983199 | 3.906 | A08 | C6H13NO2 | 0.774 | 1688693616 | Calibration Amount Diff 22.862% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_A9 | 717370126 | 1.953 | A09 | C6H13NO2 | 0.445 | 1613298357 | Calibration Amount Diff 41.148% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_A10 | 365722237 | 0.977 | A10 | C6H13NO2 | 0.218 | 1678533283 | Calibration Amount Diff 38.236% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_A11 | 212084460 | 0.488 | A11 | C6H13NO2 | 0.115 | 1844638274 | Calibration Amount Diff 46.031% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_A12 | 110317918 | 0.244 | A12 | C6H13NO2 | 0.057 | 1921443684 | Calibration Amount Diff 45.842% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_B1 | 59464817002 | 500 | B01 | C6H13NO2 | 75.854 | 783935636
Leucine pos | 6.74 | Target Compound | KD033022_033022_B2 | 38805911061 | 250 | B02 | C6H13NO2 | 35.198 | 1102516489
Leucine pos | 6.74 | Target Compound | KD033022_033022_B3 | 23710984406 | 125 | B03 | C6H13NO2 | 18.746 | 1264876185
Leucine pos | 6.74 | Target Compound | KD033022_033022_B4 | 14472235071 | 62.5 | B04 | C6H13NO2 | 10.516 | 1376209727
Leucine pos | 6.74 | Target Compound | KD033022_033022_B5 | 8197787788 | 31.25 | B05 | C6H13NO2 | 5.575 | 1470536045
Leucine pos | 6.74 | Target Compound | KD033022_033022_B6 | 4670742779 | 15.625 | B06 | C6H13NO2 | 2.937 | 1590258854
Leucine pos | 6.74 | Target Compound | KD033022_033022_B7 | 2337258642 | 7.813 | B07 | C6H13NO2 | 1.356 | 1723798232
Leucine pos | 6.74 | Target Compound | KD033022_033022_B8 | 1352317789 | 3.906 | B08 | C6H13NO2 | 0.788 | 1716065359 | Calibration Amount Diff 25.097% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_B9 | 761059520 | 1.953 | B09 | C6H13NO2 | 0.405 | 1878110824 | Calibration Amount Diff 28.628% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_B10 | 383202296 | 0.977 | B10 | C6H13NO2 | 0.215 | 1779257161 | Calibration Amount Diff 36.643% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_B11 | 210659309 | 0.488 | B11 | C6H13NO2 | 0.114 | 1847580474 | Calibration Amount Diff 44.819% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_B12 | 110410134 | 0.244 | B12 | C6H13NO2 | 0.058 | 1893072236 | Calibration Amount Diff 48.151% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_C1 | 57972378933 | 500 | C01 | C6H13NO2 | 77.559 | 747462588
Leucine pos | 6.74 | Target Compound | KD033022_033022_C2 | 37149284112 | 250 | C02 | C6H13NO2 | 37.895 | 980320808
Leucine pos | 6.74 | Target Compound | KD033022_033022_C3 | 22697521629 | 125 | C03 | C6H13NO2 | 20.061 | 1131441145
Leucine pos | 6.74 | Target Compound | KD033022_033022_C4 | 14195004666 | 62.5 | C04 | C6H13NO2 | 10.327 | 1374575976
Leucine pos | 6.74 | Target Compound | KD033022_033022_C5 | 7379788502 | 31.25 | C05 | C6H13NO2 | 5.477 | 1347407209
Leucine pos | 6.74 | Target Compound | KD033022_033022_C6 | 4494807838 | 15.625 | C06 | C6H13NO2 | 2.986 | 1505240044
Leucine pos | 6.74 | Target Compound | KD033022_033022_C7 | 2146519699 | 7.813 | C07 | C6H13NO2 | 1.432 | 1498924977
Leucine pos | 6.74 | Target Compound | KD033022_033022_C8 | 1291467421 | 3.906 | C08 | C6H13NO2 | 0.785 | 1646031944 | Calibration Amount Diff 24.551% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_C9 | 795006129 | 1.953 | C09 | C6H13NO2 | 0.441 | 1803685427 | Calibration Amount Diff 39.912% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_C10 | 414315813 | 0.977 | C10 | C6H13NO2 | 0.218 | 1898896974 | Calibration Amount Diff 38.43% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_C11 | 208017367 | 0.488 | C11 | C6H13NO2 | 0.115 | 1809417319 | Calibration Amount Diff 46.019% (absolute) > max calibration amount diff 20%
Leucine pos | 6.74 | Target Compound | KD033022_033022_C12 | 99921670 | 0.244 | C12 | C6H13NO2 | 0.059 | 1703693854 | Calibration Amount Diff 48.981% (absolute) > max calibration amount diff 20%

## Leucine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A1 | 735944481 | 1 | A01 | [13]C6H13[15]NO2 | N/A | N/A | 6.736:Peak area 735944481.454 is out of bounds (ISTD Minimum Recovery 745279247.531 and ISTD Max Recovery 2235837742.594)
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A2 | 1096220175 | 1 | A02 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A3 | 1314581659 | 1 | A03 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A4 | 1388113227 | 1 | A04 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A5 | 1467701261 | 1 | A05 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A6 | 1453679002 | 1 | A06 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A7 | 1583943317 | 1 | A07 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A8 | 1688693616 | 1 | A08 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A9 | 1613298357 | 1 | A09 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A10 | 1678533283 | 1 | A10 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A11 | 1844638274 | 1 | A11 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_A12 | 1921443684 | 1 | A12 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B1 | 783935636 | 1 | B01 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B2 | 1102516489 | 1 | B02 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B3 | 1264876185 | 1 | B03 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B4 | 1376209727 | 1 | B04 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B5 | 1470536045 | 1 | B05 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B6 | 1590258854 | 1 | B06 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B7 | 1723798232 | 1 | B07 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B8 | 1716065359 | 1 | B08 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B9 | 1878110824 | 1 | B09 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B10 | 1779257161 | 1 | B10 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B11 | 1847580474 | 1 | B11 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_B12 | 1893072236 | 1 | B12 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C1 | 747462588 | 1 | C01 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C2 | 980320808 | 1 | C02 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C3 | 1131441145 | 1 | C03 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C4 | 1374575976 | 1 | C04 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C5 | 1347407209 | 1 | C05 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C6 | 1505240044 | 1 | C06 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C7 | 1498924977 | 1 | C07 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C8 | 1646031944 | 1 | C08 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C9 | 1803685427 | 1 | C09 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C10 | 1898896974 | 1 | C10 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C11 | 1809417319 | 1 | C11 | [13]C6H13[15]NO2 | N/A | N/A
Leucine U-13C, U-15N pos | 6.74 | Internal Standard | KD033022_033022_C12 | 1703693854 | 1 | C12 | [13]C6H13[15]NO2 | N/A | N/A

## Lysine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Lysine pos | 13.78 | Target Compound | KD033022_033022_A1 | 13500574255 | 500 | A01 | C6H14N2O2 | 95.218 | 141786615
Lysine pos | 13.78 | Target Compound | KD033022_033022_A2 | 8653168484 | 250 | A02 | C6H14N2O2 | 42.007 | 205993968
Lysine pos | 13.78 | Target Compound | KD033022_033022_A3 | 5736897368 | 125 | A03 | C6H14N2O2 | 21.637 | 265148954
Lysine pos | 13.78 | Target Compound | KD033022_033022_A4 | 3319810875 | 62.5 | A04 | C6H14N2O2 | 11.576 | 286787378
Lysine pos | 13.78 | Target Compound | KD033022_033022_A5 | 1893894016 | 31.25 | A05 | C6H14N2O2 | 6.385 | 296610919
Lysine pos | 13.78 | Target Compound | KD033022_033022_A6 | 991140850 | 15.625 | A06 | C6H14N2O2 | 3.222 | 307625232
Lysine pos | 13.78 | Target Compound | KD033022_033022_A7 | 526408523 | 7.813 | A07 | C6H14N2O2 | 1.635 | 322039695
Lysine pos | 13.78 | Target Compound | KD033022_033022_A8 | 271277261 | 3.906 | A08 | C6H14N2O2 | 0.826 | 328619412
Lysine pos | 13.78 | Target Compound | KD033022_033022_A9 | 154808676 | 1.953 | A09 | C6H14N2O2 | 0.471 | 328730514 | Calibration Amount Diff 36.039% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_A10 | 76207410 | 0.977 | A10 | C6H14N2O2 | 0.227 | 336132298 | Calibration Amount Diff 30.923% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_A11 | 41505280 | 0.488 | A11 | C6H14N2O2 | 0.123 | 336273758 | Calibration Amount Diff 42.699% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_A12 | 20787682 | 0.244 | A12 | C6H14N2O2 | 0.062 | 337073032 | Calibration Amount Diff 42.602% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_B1 | 13519023828 | 500 | B01 | C6H14N2O2 | 88.136 | 153388043
Lysine pos | 13.78 | Target Compound | KD033022_033022_B2 | 9146175013 | 250 | B02 | C6H14N2O2 | 40.524 | 225696857
Lysine pos | 13.78 | Target Compound | KD033022_033022_B3 | 5792615358 | 125 | B03 | C6H14N2O2 | 21.429 | 270322948
Lysine pos | 13.78 | Target Compound | KD033022_033022_B4 | 3340527525 | 62.5 | B04 | C6H14N2O2 | 11.453 | 291671009
Lysine pos | 13.78 | Target Compound | KD033022_033022_B5 | 1846114585 | 31.25 | B05 | C6H14N2O2 | 6.11 | 302156152
Lysine pos | 13.78 | Target Compound | KD033022_033022_B6 | 1097315776 | 15.625 | B06 | C6H14N2O2 | 3.342 | 328372205 | Calibration Amount Diff 20.611% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_B7 | 507109679 | 7.813 | B07 | C6H14N2O2 | 1.52 | 333521506
Lysine pos | 13.78 | Target Compound | KD033022_033022_B8 | 282385494 | 3.906 | B08 | C6H14N2O2 | 0.838 | 336981282 | Calibration Amount Diff 21.03% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_B9 | 154669407 | 1.953 | B09 | C6H14N2O2 | 0.448 | 345508053 | Calibration Amount Diff 29.317% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_B10 | 82808628 | 0.977 | B10 | C6H14N2O2 | 0.231 | 359140234 | Calibration Amount Diff 33.15% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_B11 | 43313481 | 0.488 | B11 | C6H14N2O2 | 0.12 | 361233402 | Calibration Amount Diff 38.626% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_B12 | 22721678 | 0.244 | B12 | C6H14N2O2 | 0.064 | 355608945 | Calibration Amount Diff 47.745% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_C1 | 13797708783 | 500 | C01 | C6H14N2O2 | 90.692 | 152138172
Lysine pos | 13.78 | Target Compound | KD033022_033022_C2 | 8607595488 | 250 | C02 | C6H14N2O2 | 41.246 | 208690403
Lysine pos | 13.78 | Target Compound | KD033022_033022_C3 | 5601675094 | 125 | C03 | C6H14N2O2 | 22.202 | 252310392
Lysine pos | 13.78 | Target Compound | KD033022_033022_C4 | 3295317438 | 62.5 | C04 | C6H14N2O2 | 11.358 | 290139418
Lysine pos | 13.78 | Target Compound | KD033022_033022_C5 | 1819076128 | 31.25 | C05 | C6H14N2O2 | 6.243 | 291360449
Lysine pos | 13.78 | Target Compound | KD033022_033022_C6 | 1008338626 | 15.625 | C06 | C6H14N2O2 | 3.232 | 311956402
Lysine pos | 13.78 | Target Compound | KD033022_033022_C7 | 490710394 | 7.813 | C07 | C6H14N2O2 | 1.546 | 317372082
Lysine pos | 13.78 | Target Compound | KD033022_033022_C8 | 278189272 | 3.906 | C08 | C6H14N2O2 | 0.832 | 334462028 | Calibration Amount Diff 20.129% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_C9 | 171638767 | 1.953 | C09 | C6H14N2O2 | 0.476 | 360488772 | Calibration Amount Diff 37.54% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_C10 | 82282326 | 0.977 | C10 | C6H14N2O2 | 0.231 | 356333798 | Calibration Amount Diff 33.346% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_C11 | 41928404 | 0.488 | C11 | C6H14N2O2 | 0.123 | 340209704 | Calibration Amount Diff 42.486% (absolute) > max calibration amount diff 20%
Lysine pos | 13.78 | Target Compound | KD033022_033022_C12 | 21097260 | 0.244 | C12 | C6H14N2O2 | 0.063 | 334991567 | Calibration Amount Diff 45.625% (absolute) > max calibration amount diff 20%

## Lysine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A1 | 141786615 | 1 | A01 | [13]C6H14[15]N2O2 | N/A | N/A | 13.772:Peak area 141786614.697 is out of bounds (ISTD Minimum Recovery 148706605.529 and ISTD Max Recovery 446119816.586)
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A2 | 205993968 | 1 | A02 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A3 | 265148954 | 1 | A03 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A4 | 286787378 | 1 | A04 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A5 | 296610919 | 1 | A05 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A6 | 307625232 | 1 | A06 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A7 | 322039695 | 1 | A07 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A8 | 328619412 | 1 | A08 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A9 | 328730514 | 1 | A09 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A10 | 336132298 | 1 | A10 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A11 | 336273758 | 1 | A11 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_A12 | 337073032 | 1 | A12 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B1 | 153388043 | 1 | B01 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B2 | 225696857 | 1 | B02 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B3 | 270322948 | 1 | B03 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B4 | 291671009 | 1 | B04 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B5 | 302156152 | 1 | B05 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B6 | 328372205 | 1 | B06 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B7 | 333521506 | 1 | B07 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B8 | 336981282 | 1 | B08 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B9 | 345508053 | 1 | B09 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B10 | 359140234 | 1 | B10 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B11 | 361233402 | 1 | B11 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_B12 | 355608945 | 1 | B12 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C1 | 152138172 | 1 | C01 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C2 | 208690403 | 1 | C02 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C3 | 252310392 | 1 | C03 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C4 | 290139418 | 1 | C04 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C5 | 291360449 | 1 | C05 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C6 | 311956402 | 1 | C06 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C7 | 317372082 | 1 | C07 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C8 | 334462028 | 1 | C08 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C9 | 360488772 | 1 | C09 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C10 | 356333798 | 1 | C10 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C11 | 340209704 | 1 | C11 | [13]C6H14[15]N2O2 | N/A | N/A
Lysine U-13C, U-15N pos | 13.78 | Internal Standard | KD033022_033022_C12 | 334991567 | 1 | C12 | [13]C6H14[15]N2O2 | N/A | N/A

## Methionine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Methionine pos | 7.14 | Target Compound | KD033022_033022_A1 | 64316783366 | 500 | A01 | C5H11NO2S | 96.644 | 665499527
Methionine pos | 7.14 | Target Compound | KD033022_033022_A2 | 43653100644 | 250 | A02 | C5H11NO2S | 41.308 | 1056775645
Methionine pos | 7.14 | Target Compound | KD033022_033022_A3 | 30025377985 | 125 | A03 | C5H11NO2S | 21.448 | 1399933484
Methionine pos | 7.14 | Target Compound | KD033022_033022_A4 | 19628042378 | 62.5 | A04 | C5H11NO2S | 11.48 | 1709708919
Methionine pos | 7.14 | Target Compound | KD033022_033022_A5 | 12288456098 | 31.25 | A05 | C5H11NO2S | 6.352 | 1934576446
Methionine pos | 7.14 | Target Compound | KD033022_033022_A6 | 6696703412 | 15.625 | A06 | C5H11NO2S | 3.192 | 2097964808
Methionine pos | 7.14 | Target Compound | KD033022_033022_A7 | 3767075414 | 7.813 | A07 | C5H11NO2S | 1.619 | 2327359394
Methionine pos | 7.14 | Target Compound | KD033022_033022_A8 | 1940718073 | 3.906 | A08 | C5H11NO2S | 0.816 | 2378566483
Methionine pos | 7.14 | Target Compound | KD033022_033022_A9 | 1139553758 | 1.953 | A09 | C5H11NO2S | 0.466 | 2447251233 | Calibration Amount Diff 35.256% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_A10 | 570890805 | 0.977 | A10 | C5H11NO2S | 0.226 | 2523279544 | Calibration Amount Diff 31.386% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_A11 | 311010063 | 0.488 | A11 | C5H11NO2S | 0.118 | 2633227294 | Calibration Amount Diff 37.325% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_A12 | 160946038 | 0.244 | A12 | C5H11NO2S | 0.059 | 2733631846 | Calibration Amount Diff 36.913% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_B1 | 62616082599 | 500 | B01 | C5H11NO2S | 91.597 | 683602466
Methionine pos | 7.14 | Target Compound | KD033022_033022_B2 | 44828713118 | 250 | B02 | C5H11NO2S | 41.624 | 1076986517
Methionine pos | 7.14 | Target Compound | KD033022_033022_B3 | 30261744209 | 125 | B03 | C5H11NO2S | 21.726 | 1392899370
Methionine pos | 7.14 | Target Compound | KD033022_033022_B4 | 19650648010 | 62.5 | B04 | C5H11NO2S | 11.4 | 1723722889
Methionine pos | 7.14 | Target Compound | KD033022_033022_B5 | 11777035614 | 31.25 | B05 | C5H11NO2S | 6.061 | 1943232770
Methionine pos | 7.14 | Target Compound | KD033022_033022_B6 | 6986256373 | 15.625 | B06 | C5H11NO2S | 3.329 | 2098452578 | Calibration Amount Diff 20.694% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_B7 | 3448295765 | 7.813 | B07 | C5H11NO2S | 1.498 | 2302263157
Methionine pos | 7.14 | Target Compound | KD033022_033022_B8 | 2069502925 | 3.906 | B08 | C5H11NO2S | 0.822 | 2516845766
Methionine pos | 7.14 | Target Compound | KD033022_033022_B9 | 1150339042 | 1.953 | B09 | C5H11NO2S | 0.446 | 2580765273 | Calibration Amount Diff 29.474% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_B10 | 570293049 | 0.977 | B10 | C5H11NO2S | 0.224 | 2547701243 | Calibration Amount Diff 29.991% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_B11 | 312307519 | 0.488 | B11 | C5H11NO2S | 0.118 | 2655800637 | Calibration Amount Diff 36.725% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_B12 | 164148214 | 0.244 | B12 | C5H11NO2S | 0.059 | 2773659985 | Calibration Amount Diff 37.622% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_C1 | 61545868476 | 500 | C01 | C5H11NO2S | 93.078 | 661229789
Methionine pos | 7.14 | Target Compound | KD033022_033022_C2 | 41620749672 | 250 | C02 | C5H11NO2S | 42.534 | 978526807
Methionine pos | 7.14 | Target Compound | KD033022_033022_C3 | 28948906710 | 125 | C03 | C5H11NO2S | 22.341 | 1295772085
Methionine pos | 7.14 | Target Compound | KD033022_033022_C4 | 18789048337 | 62.5 | C04 | C5H11NO2S | 11.518 | 1631299610
Methionine pos | 7.14 | Target Compound | KD033022_033022_C5 | 11146476188 | 31.25 | C05 | C5H11NO2S | 6.203 | 1797064086
Methionine pos | 7.14 | Target Compound | KD033022_033022_C6 | 6559212130 | 15.625 | C06 | C5H11NO2S | 3.181 | 2061750343
Methionine pos | 7.14 | Target Compound | KD033022_033022_C7 | 3274409604 | 7.813 | C07 | C5H11NO2S | 1.524 | 2148753291
Methionine pos | 7.14 | Target Compound | KD033022_033022_C8 | 1894303436 | 3.906 | C08 | C5H11NO2S | 0.826 | 2292359918
Methionine pos | 7.14 | Target Compound | KD033022_033022_C9 | 1191155706 | 1.953 | C09 | C5H11NO2S | 0.463 | 2570307329 | Calibration Amount Diff 34.613% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_C10 | 617929546 | 0.977 | C10 | C5H11NO2S | 0.229 | 2694414710 | Calibration Amount Diff 33.179% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_C11 | 312176042 | 0.488 | C11 | C5H11NO2S | 0.118 | 2641366592 | Calibration Amount Diff 37.415% (absolute) > max calibration amount diff 20%
Methionine pos | 7.14 | Target Compound | KD033022_033022_C12 | 148161652 | 0.244 | C12 | C5H11NO2S | 0.059 | 2506564148 | Calibration Amount Diff 37.456% (absolute) > max calibration amount diff 20%

## Methionine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A1 | 665499527 | 1 | A01 | [13]C5H11[15]NO2S | N/A | N/A | 7.207:Peak area 665499527.335 is out of bounds (ISTD Minimum Recovery 992821055.303 and ISTD Max Recovery 2978463165.908)
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A2 | 1056775645 | 1 | A02 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A3 | 1399933484 | 1 | A03 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A4 | 1709708919 | 1 | A04 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A5 | 1934576446 | 1 | A05 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A6 | 2097964808 | 1 | A06 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A7 | 2327359394 | 1 | A07 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A8 | 2378566483 | 1 | A08 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A9 | 2447251233 | 1 | A09 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A10 | 2523279544 | 1 | A10 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A11 | 2633227294 | 1 | A11 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_A12 | 2733631846 | 1 | A12 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B1 | 683602466 | 1 | B01 | [13]C5H11[15]NO2S | N/A | N/A | 7.214:Peak area 683602466.131 is out of bounds (ISTD Minimum Recovery 992821055.303 and ISTD Max Recovery 2978463165.908)
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B2 | 1076986517 | 1 | B02 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B3 | 1392899370 | 1 | B03 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B4 | 1723722889 | 1 | B04 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B5 | 1943232770 | 1 | B05 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B6 | 2098452578 | 1 | B06 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B7 | 2302263157 | 1 | B07 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B8 | 2516845766 | 1 | B08 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B9 | 2580765273 | 1 | B09 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B10 | 2547701243 | 1 | B10 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B11 | 2655800637 | 1 | B11 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_B12 | 2773659985 | 1 | B12 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C1 | 661229789 | 1 | C01 | [13]C5H11[15]NO2S | N/A | N/A | 7.214:Peak area 661229789.227 is out of bounds (ISTD Minimum Recovery 992821055.303 and ISTD Max Recovery 2978463165.908)
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C2 | 978526807 | 1 | C02 | [13]C5H11[15]NO2S | N/A | N/A | 7.185:Peak area 978526807.106 is out of bounds (ISTD Minimum Recovery 992821055.303 and ISTD Max Recovery 2978463165.908)
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C3 | 1295772085 | 1 | C03 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C4 | 1631299610 | 1 | C04 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C5 | 1797064086 | 1 | C05 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C6 | 2061750343 | 1 | C06 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C7 | 2148753291 | 1 | C07 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C8 | 2292359918 | 1 | C08 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C9 | 2570307329 | 1 | C09 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C10 | 2694414710 | 1 | C10 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C11 | 2641366592 | 1 | C11 | [13]C5H11[15]NO2S | N/A | N/A
Methionine U-13C, U-15N pos | 7.14 | Internal Standard | KD033022_033022_C12 | 2506564148 | 1 | C12 | [13]C5H11[15]NO2S | N/A | N/A

## Phenylalanine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A1 | 97867768790 | 500 | A01 | C9H11NO2 | 91.728 | 1066933124
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A2 | 61223817644 | 250 | A02 | C9H11NO2 | 40.164 | 1524344491
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A3 | 40394259709 | 125 | A03 | C9H11NO2 | 20.872 | 1935293511
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A4 | 23594617589 | 62.5 | A04 | C9H11NO2 | 11.193 | 2107905495
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A5 | 14197876729 | 31.25 | A05 | C9H11NO2 | 6.138 | 2313162945
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A6 | 7507889982 | 15.625 | A06 | C9H11NO2 | 3.135 | 2395148407
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A7 | 3934500560 | 7.813 | A07 | C9H11NO2 | 1.592 | 2471871911
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A8 | 2031810294 | 3.906 | A08 | C9H11NO2 | 0.804 | 2528297898
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A9 | 1114273481 | 1.953 | A09 | C9H11NO2 | 0.458 | 2434422585 | Calibration Amount Diff 36.221% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A10 | 581734150 | 0.977 | A10 | C9H11NO2 | 0.225 | 2588158180 | Calibration Amount Diff 33.72% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A11 | 307376952 | 0.488 | A11 | C9H11NO2 | 0.119 | 2592728079 | Calibration Amount Diff 41.206% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_A12 | 159087671 | 0.244 | A12 | C9H11NO2 | 0.059 | 2686875899 | Calibration Amount Diff 41.045% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B1 | 96438639003 | 500 | B01 | C9H11NO2 | 84.991 | 1134686617
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B2 | 62572361013 | 250 | B02 | C9H11NO2 | 39.46 | 1585727657
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B3 | 39810285399 | 125 | B03 | C9H11NO2 | 21.134 | 1883736444
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B4 | 24190236064 | 62.5 | B04 | C9H11NO2 | 11.103 | 2178777709
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B5 | 14017521915 | 31.25 | B05 | C9H11NO2 | 5.92 | 2367690326
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B6 | 7925728066 | 15.625 | B06 | C9H11NO2 | 3.232 | 2452438673 | Calibration Amount Diff 20.201% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B7 | 3770187594 | 7.813 | B07 | C9H11NO2 | 1.477 | 2552502719
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B8 | 2075774114 | 3.906 | B08 | C9H11NO2 | 0.815 | 2547891717 | Calibration Amount Diff 21.23% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B9 | 1184916839 | 1.953 | B09 | C9H11NO2 | 0.439 | 2699735802 | Calibration Amount Diff 30.622% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B10 | 583353768 | 0.977 | B10 | C9H11NO2 | 0.223 | 2614848118 | Calibration Amount Diff 32.723% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B11 | 315136362 | 0.488 | B11 | C9H11NO2 | 0.118 | 2681125926 | Calibration Amount Diff 39.997% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_B12 | 162422298 | 0.244 | B12 | C9H11NO2 | 0.06 | 2726998287 | Calibration Amount Diff 41.883% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C1 | 96787793366 | 500 | C01 | C9H11NO2 | 87.431 | 1107021226
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C2 | 59376645467 | 250 | C02 | C9H11NO2 | 39.719 | 1494904832
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C3 | 37770970135 | 125 | C03 | C9H11NO2 | 21.509 | 1756082248
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C4 | 22734395243 | 62.5 | C04 | C9H11NO2 | 11.07 | 2053765167
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C5 | 13180427706 | 31.25 | C05 | C9H11NO2 | 6.077 | 2168754466
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C6 | 7384318252 | 15.625 | C06 | C9H11NO2 | 3.11 | 2374246923
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C7 | 3550272542 | 7.813 | C07 | C9H11NO2 | 1.49 | 2382793082
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C8 | 1978856717 | 3.906 | C08 | C9H11NO2 | 0.809 | 2444679694 | Calibration Amount Diff 20.449% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C9 | 1252470166 | 1.953 | C09 | C9H11NO2 | 0.457 | 2739061205 | Calibration Amount Diff 36.086% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C10 | 626821289 | 0.977 | C10 | C9H11NO2 | 0.228 | 2754075807 | Calibration Amount Diff 35.403% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C11 | 313138972 | 0.488 | C11 | C9H11NO2 | 0.117 | 2671091068 | Calibration Amount Diff 39.633% (absolute) > max calibration amount diff 20%
Phenylalanine pos | 6.19 | Target Compound | KD033022_033022_C12 | 151506166 | 0.244 | C12 | C9H11NO2 | 0.06 | 2524596943 | Calibration Amount Diff 42.958% (absolute) > max calibration amount diff 20%

## Phenylalanine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A1 | 1066933124 | 1 | A01 | [13]C9H11[15]NO2 | N/A | N/A | 6.261:Peak area 1066933123.516 is out of bounds (ISTD Minimum Recovery 1118644099.731 and ISTD Max Recovery 3355932299.193)
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A2 | 1524344491 | 1 | A02 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A3 | 1935293511 | 1 | A03 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A4 | 2107905495 | 1 | A04 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A5 | 2313162945 | 1 | A05 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A6 | 2395148407 | 1 | A06 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A7 | 2471871911 | 1 | A07 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A8 | 2528297898 | 1 | A08 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A9 | 2434422585 | 1 | A09 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A10 | 2588158180 | 1 | A10 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A11 | 2592728079 | 1 | A11 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_A12 | 2686875899 | 1 | A12 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B1 | 1134686617 | 1 | B01 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B2 | 1585727657 | 1 | B02 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B3 | 1883736444 | 1 | B03 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B4 | 2178777709 | 1 | B04 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B5 | 2367690326 | 1 | B05 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B6 | 2452438673 | 1 | B06 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B7 | 2552502719 | 1 | B07 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B8 | 2547891717 | 1 | B08 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B9 | 2699735802 | 1 | B09 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B10 | 2614848118 | 1 | B10 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B11 | 2681125926 | 1 | B11 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_B12 | 2726998287 | 1 | B12 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C1 | 1107021226 | 1 | C01 | [13]C9H11[15]NO2 | N/A | N/A | 6.240:Peak area 1107021225.586 is out of bounds (ISTD Minimum Recovery 1118644099.731 and ISTD Max Recovery 3355932299.193)
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C2 | 1494904832 | 1 | C02 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C3 | 1756082248 | 1 | C03 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C4 | 2053765167 | 1 | C04 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C5 | 2168754466 | 1 | C05 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C6 | 2374246923 | 1 | C06 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C7 | 2382793082 | 1 | C07 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C8 | 2444679694 | 1 | C08 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C9 | 2739061205 | 1 | C09 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C10 | 2754075807 | 1 | C10 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C11 | 2671091068 | 1 | C11 | [13]C9H11[15]NO2 | N/A | N/A
Phenylalanine U-13C, U-15N pos | 6.19 | Internal Standard | KD033022_033022_C12 | 2524596943 | 1 | C12 | [13]C9H11[15]NO2 | N/A | N/A

## Proline pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | RR | Theoretical Amt | power pred | %diff
Proline pos | 8.15 | Target Compound | KD033022_033022_A1 | 105830151069 | 500 | A01 | C5H9NO2 | 85.159 | 1242732852 | 85.1592 | 500 | 561.9377 | 12.3875
Proline pos | 8.15 | Target Compound | KD033022_033022_A2 | 81972073862 | 250 | A02 | C5H9NO2 | 38.321 | 2139101252 | 38.3208 | 250 | 239.6425 | 4.143
Proline pos | 8.15 | Target Compound | KD033022_033022_A3 | 58724665409 | 125 | A03 | C5H9NO2 | 19.803 | 2965490433 | 19.8027 | 125 | 118.4589 | 5.2329
Proline pos | 8.15 | Target Compound | KD033022_033022_A4 | 40900709685 | 62.5 | A04 | C5H9NO2 | 10.753 | 3803501397 | 10.7534 | 62.5 | 61.7382 | 1.2189
Proline pos | 8.15 | Target Compound | KD033022_033022_A5 | 29097280816 | 31.25 | A05 | C5H9NO2 | 5.907 | 4926074358 | 5.9068 | 31.25 | 32.5729 | 4.2332
Proline pos | 8.15 | Target Compound | KD033022_033022_A6 | 17589912501 | 15.625 | A06 | C5H9NO2 | 2.977 | 5907778992 | 2.9774 | 15.625 | 15.6795 | 0.34869
Proline pos | 8.15 | Target Compound | KD033022_033022_A7 | 10468748959 | 7.813 | A07 | C5H9NO2 | 1.524 | 6867412775 | 1.5244 | 7.813 | 7.6743 | 1.7755
Proline pos | 8.15 | Target Compound | KD033022_033022_A8 | 5719658121 | 3.906 | A08 | C5H9NO2 | 0.784 | 7295434518 | 0.78401 | 3.906 | 3.7742 | 3.3732
Proline pos | 8.15 | Target Compound | KD033022_033022_A9 | 3357710952 | 1.953 | A09 | C5H9NO2 | 0.448 | 7500604498 | 0.44766 | 1.953 | 2.0753 | 6.2638
Proline pos | 8.15 | Target Compound | KD033022_033022_A10 | 1730312178 | 0.977 | A10 | C5H9NO2 | 0.222 | 7789091331 | 0.22215 | 0.977 | 0.98245 | 0.55744
Proline pos | 8.15 | Target Compound | KD033022_033022_A11 | 981628598 | 0.488 | A11 | C5H9NO2 | 0.119 | 8262470437 | 0.11881 | 0.488 | 0.50376 | 3.23
Proline pos | 8.15 | Target Compound | KD033022_033022_A12 | 492483253 | 0.244 | A12 | C5H9NO2 | 0.059 | 8297480502 | 0.059353 | 0.244 | 0.24019 | 1.5602
Proline pos | 8.15 | Target Compound | KD033022_033022_B1 | 106921714713 | 500 | B01 | C5H9NO2 | 81.926 | 1305101761 | 81.926 | 500 | 539.1968 | 7.8394
Proline pos | 8.15 | Target Compound | KD033022_033022_B2 | 82702000283 | 250 | B02 | C5H9NO2 | 37.966 | 2178291440 | 37.9665 | 250 | 237.2782 | 5.0887
Proline pos | 8.15 | Target Compound | KD033022_033022_B3 | 60133331382 | 125 | B03 | C5H9NO2 | 20.079 | 2994772544 | 20.0794 | 125 | 120.2266 | 3.8187
Proline pos | 8.15 | Target Compound | KD033022_033022_B4 | 41955824239 | 62.5 | B04 | C5H9NO2 | 10.487 | 4000767373 | 10.4869 | 62.5 | 60.1066 | 3.8294
Proline pos | 8.15 | Target Compound | KD033022_033022_B5 | 28550995736 | 31.25 | B05 | C5H9NO2 | 5.64 | 5061843352 | 5.6404 | 31.25 | 31.0077 | 0.77546
Proline pos | 8.15 | Target Compound | KD033022_033022_B6 | 18065076765 | 15.625 | B06 | C5H9NO2 | 3.079 | 5866942870 | 3.0791 | 15.625 | 16.2518 | 4.0115
Proline pos | 8.15 | Target Compound | KD033022_033022_B7 | 9846431878 | 7.813 | B07 | C5H9NO2 | 1.408 | 6992483144 | 1.4081 | 7.813 | 7.0512 | 9.7499
Proline pos | 8.15 | Target Compound | KD033022_033022_B8 | 5851714138 | 3.906 | B08 | C5H9NO2 | 0.785 | 7456566145 | 0.78477 | 3.906 | 3.7782 | 3.2722
Proline pos | 8.15 | Target Compound | KD033022_033022_B9 | 3511810253 | 1.953 | B09 | C5H9NO2 | 0.428 | 8209978758 | 0.42775 | 1.953 | 1.977 | 1.2274
Proline pos | 8.15 | Target Compound | KD033022_033022_B10 | 1715794747 | 0.977 | B10 | C5H9NO2 | 0.22 | 7782367010 | 0.22047 | 0.977 | 0.97455 | 0.25084
Proline pos | 8.15 | Target Compound | KD033022_033022_B11 | 948446696 | 0.488 | B11 | C5H9NO2 | 0.117 | 8078721043 | 0.1174 | 0.488 | 0.49741 | 1.9275
Proline pos | 8.15 | Target Compound | KD033022_033022_B12 | 486913221 | 0.244 | B12 | C5H9NO2 | 0.061 | 8029141632 | 0.060643 | 0.244 | 0.24577 | 0.7247
Proline pos | 8.15 | Target Compound | KD033022_033022_C1 | 104949593085 | 500 | C01 | C5H9NO2 | 83.241 | 1260796977 | 83.2407 | 500 | 548.4367 | 9.6873
Proline pos | 8.15 | Target Compound | KD033022_033022_C2 | 78667601433 | 250 | C02 | C5H9NO2 | 38.024 | 2068880863 | 38.0242 | 250 | 237.6636 | 4.9346
Proline pos | 8.15 | Target Compound | KD033022_033022_C3 | 56357151017 | 125 | C03 | C5H9NO2 | 20.484 | 2751241058 | 20.4843 | 125 | 122.8153 | 1.7477
Proline pos | 8.15 | Target Compound | KD033022_033022_C4 | 38970085470 | 62.5 | C04 | C5H9NO2 | 10.6 | 3676576244 | 10.5996 | 62.5 | 60.7958 | 2.7268
Proline pos | 8.15 | Target Compound | KD033022_033022_C5 | 25559949886 | 31.25 | C05 | C5H9NO2 | 5.774 | 4426536932 | 5.7743 | 31.25 | 31.7934 | 1.739
Proline pos | 8.15 | Target Compound | KD033022_033022_C6 | 16717662343 | 15.625 | C06 | C5H9NO2 | 3.005 | 5564206533 | 3.0045 | 15.625 | 15.8318 | 1.3233
Proline pos | 8.15 | Target Compound | KD033022_033022_C7 | 8990016187 | 7.813 | C07 | C5H9NO2 | 1.436 | 6258824962 | 1.4364 | 7.813 | 7.2022 | 7.8176
Proline pos | 8.15 | Target Compound | KD033022_033022_C8 | 5401439770 | 3.906 | C08 | C5H9NO2 | 0.788 | 6852898301 | 0.7882 | 3.906 | 3.7958 | 2.8216
Proline pos | 8.15 | Target Compound | KD033022_033022_C9 | 3357617687 | 1.953 | C09 | C5H9NO2 | 0.447 | 7511707323 | 0.44698 | 1.953 | 2.072 | 6.093
Proline pos | 8.15 | Target Compound | KD033022_033022_C10 | 1708402224 | 0.977 | C10 | C5H9NO2 | 0.226 | 7568491398 | 0.22573 | 0.977 | 0.99935 | 2.2879
Proline pos | 8.15 | Target Compound | KD033022_033022_C11 | 899086282 | 0.488 | C11 | C5H9NO2 | 0.119 | 7537254714 | 0.11929 | 0.488 | 0.50593 | 3.6751
Proline pos | 8.15 | Target Compound | KD033022_033022_C12 | 452617037 | 0.244 | C12 | C5H9NO2 | 0.06 | 7481297535 | 0.0605 | 0.244 | 0.24515 | 0.47045

## Proline U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A1 | 1242732852 | 1 | A01 | [13]C5H9[15]NO2 | N/A | N/A | 8.155:Peak area 1242732851.687 is out of bounds (ISTD Minimum Recovery 2748789767.497 and ISTD Max Recovery 8246369302.492)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A2 | 2139101252 | 1 | A02 | [13]C5H9[15]NO2 | N/A | N/A | 8.125:Peak area 2139101251.560 is out of bounds (ISTD Minimum Recovery 2748789767.497 and ISTD Max Recovery 8246369302.492)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A3 | 2965490433 | 1 | A03 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A4 | 3803501397 | 1 | A04 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A5 | 4926074358 | 1 | A05 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A6 | 5907778992 | 1 | A06 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A7 | 6867412775 | 1 | A07 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A8 | 7295434518 | 1 | A08 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A9 | 7500604498 | 1 | A09 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A10 | 7789091331 | 1 | A10 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A11 | 8262470437 | 1 | A11 | [13]C5H9[15]NO2 | N/A | N/A | 8.162:Peak area 8262470437.161 is out of bounds (ISTD Minimum Recovery 2748789767.497 and ISTD Max Recovery 8246369302.492)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_A12 | 8297480502 | 1 | A12 | [13]C5H9[15]NO2 | N/A | N/A | 8.160:Peak area 8297480502.010 is out of bounds (ISTD Minimum Recovery 2748789767.497 and ISTD Max Recovery 8246369302.492)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B1 | 1305101761 | 1 | B01 | [13]C5H9[15]NO2 | N/A | N/A | 8.133:Peak area 1305101760.826 is out of bounds (ISTD Minimum Recovery 2748789767.497 and ISTD Max Recovery 8246369302.492)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B2 | 2178291440 | 1 | B02 | [13]C5H9[15]NO2 | N/A | N/A | 8.152:Peak area 2178291440.284 is out of bounds (ISTD Minimum Recovery 2748789767.497 and ISTD Max Recovery 8246369302.492)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B3 | 2994772544 | 1 | B03 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B4 | 4000767373 | 1 | B04 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B5 | 5061843352 | 1 | B05 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B6 | 5866942870 | 1 | B06 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B7 | 6992483144 | 1 | B07 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B8 | 7456566145 | 1 | B08 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B9 | 8209978758 | 1 | B09 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B10 | 7782367010 | 1 | B10 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B11 | 8078721043 | 1 | B11 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_B12 | 8029141632 | 1 | B12 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C1 | 1260796977 | 1 | C01 | [13]C5H9[15]NO2 | N/A | N/A | 8.162:Peak area 1260796976.855 is out of bounds (ISTD Minimum Recovery 2748789767.497 and ISTD Max Recovery 8246369302.492)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C2 | 2068880863 | 1 | C02 | [13]C5H9[15]NO2 | N/A | N/A | 8.156:Peak area 2068880862.965 is out of bounds (ISTD Minimum Recovery 2748789767.497 and ISTD Max Recovery 8246369302.492)
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C3 | 2751241058 | 1 | C03 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C4 | 3676576244 | 1 | C04 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C5 | 4426536932 | 1 | C05 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C6 | 5564206533 | 1 | C06 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C7 | 6258824962 | 1 | C07 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C8 | 6852898301 | 1 | C08 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C9 | 7511707323 | 1 | C09 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C10 | 7568491398 | 1 | C10 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C11 | 7537254714 | 1 | C11 | [13]C5H9[15]NO2 | N/A | N/A
Proline U-13C, U-15N pos | 8.15 | Internal Standard | KD033022_033022_C12 | 7481297535 | 1 | C12 | [13]C5H9[15]NO2 | N/A | N/A

## Serine neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Serine neg | 10.04 | Target Compound | KD033022_033022_A1 | 2139337813 | 500 | A01 | C3H7NO3 | 117.061 | 18275465
Serine neg | 10.04 | Target Compound | KD033022_033022_A2 | 1139542703 | 250 | A02 | C3H7NO3 | 46.623 | 24441817
Serine neg | 10.04 | Target Compound | KD033022_033022_A3 | 639141597 | 125 | A03 | C3H7NO3 | 23.421 | 27289227
Serine neg | 10.04 | Target Compound | KD033022_033022_A4 | 387346920 | 62.5 | A04 | C3H7NO3 | 12.189 | 31777615
Serine neg | 10.04 | Target Compound | KD033022_033022_A5 | 242969588 | 31.25 | A05 | C3H7NO3 | 6.642 | 36580588
Serine neg | 10.04 | Target Compound | KD033022_033022_A6 | 147942409 | 15.625 | A06 | C3H7NO3 | 3.182 | 46494187
Serine neg | 10.04 | Target Compound | KD033022_033022_A7 | 94809908 | 7.813 | A07 | C3H7NO3 | 1.623 | 58405804
Serine neg | 10.04 | Target Compound | KD033022_033022_A8 | 54170812 | 3.906 | A08 | C3H7NO3 | 0.786 | 68879315
Serine neg | 10.04 | Target Compound | KD033022_033022_A9 | 33741632 | 1.953 | A09 | C3H7NO3 | 0.437 | 77201989 | Calibration Amount Diff 20.072% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD033022_033022_A10 | 17891724 | 0.977 | A10 | C3H7NO3 | 0.216 | 82715737
Serine neg | 10.04 | Target Compound | KD033022_033022_A11 | 9057537 | 0.488 | A11 | C3H7NO3 | 0.105 | 86583246
Serine neg | 10.04 | Target Compound | KD033022_033022_A12 | 4049751 | 0.244 | A12 | C3H7NO3 | 0.046 | 87183547
Serine neg | 10.04 | Target Compound | KD033022_033022_B1 | 2037290702 | 500 | B01 | C3H7NO3 | 114.449 | 17800815
Serine neg | 10.04 | Target Compound | KD033022_033022_B2 | 1047164605 | 250 | B02 | C3H7NO3 | 52.702 | 19869720
Serine neg | 10.04 | Target Compound | KD033022_033022_B3 | 601720300 | 125 | B03 | C3H7NO3 | 23.62 | 25475560
Serine neg | 10.04 | Target Compound | KD033022_033022_B4 | 358898616 | 62.5 | B04 | C3H7NO3 | 12.405 | 28932808
Serine neg | 10.04 | Target Compound | KD033022_033022_B5 | 220716593 | 31.25 | B05 | C3H7NO3 | 6.288 | 35101099
Serine neg | 10.04 | Target Compound | KD033022_033022_B6 | 149848596 | 15.625 | B06 | C3H7NO3 | 3.47 | 43187487
Serine neg | 10.04 | Target Compound | KD033022_033022_B7 | 87613883 | 7.813 | B07 | C3H7NO3 | 1.442 | 60777825
Serine neg | 10.04 | Target Compound | KD033022_033022_B8 | 56457717 | 3.906 | B08 | C3H7NO3 | 0.805 | 70144969
Serine neg | 10.04 | Target Compound | KD033022_033022_B9 | 34016597 | 1.953 | B09 | C3H7NO3 | 0.443 | 76765160 | Calibration Amount Diff 21.737% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD033022_033022_B10 | 17296787 | 0.977 | B10 | C3H7NO3 | 0.207 | 83470378
Serine neg | 10.04 | Target Compound | KD033022_033022_B11 | 9979254 | 0.488 | B11 | C3H7NO3 | 0.116 | 85687406 | Calibration Amount Diff 28.17% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD033022_033022_B12 | 4801117 | 0.244 | B12 | C3H7NO3 | 0.053 | 91196249
Serine neg | 10.04 | Target Compound | KD033022_033022_C1 | 1990505068 | 500 | C01 | C3H7NO3 | 129.118 | 15416191
Serine neg | 10.04 | Target Compound | KD033022_033022_C2 | 978573334 | 250 | C02 | C3H7NO3 | 53.394 | 18327370
Serine neg | 10.04 | Target Compound | KD033022_033022_C3 | 587514443 | 125 | C03 | C3H7NO3 | 26.437 | 22222792
Serine neg | 10.04 | Target Compound | KD033022_033022_C4 | 355565085 | 62.5 | C04 | C3H7NO3 | 12.467 | 28519430
Serine neg | 10.04 | Target Compound | KD033022_033022_C5 | 221911400 | 31.25 | C05 | C3H7NO3 | 6.409 | 34623388
Serine neg | 10.04 | Target Compound | KD033022_033022_C6 | 143734864 | 15.625 | C06 | C3H7NO3 | 3.238 | 44393757
Serine neg | 10.04 | Target Compound | KD033022_033022_C7 | 84992111 | 7.813 | C07 | C3H7NO3 | 1.459 | 58261066
Serine neg | 10.04 | Target Compound | KD033022_033022_C8 | 53262125 | 3.906 | C08 | C3H7NO3 | 0.795 | 67000200
Serine neg | 10.04 | Target Compound | KD033022_033022_C9 | 35125683 | 1.953 | C09 | C3H7NO3 | 0.461 | 76240950 | Calibration Amount Diff 26.564% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD033022_033022_C10 | 17817221 | 0.977 | C10 | C3H7NO3 | 0.216 | 82635974
Serine neg | 10.04 | Target Compound | KD033022_033022_C11 | 9950011 | 0.488 | C11 | C3H7NO3 | 0.116 | 85877633 | Calibration Amount Diff 27.512% (absolute) > max calibration amount diff 20%
Serine neg | 10.04 | Target Compound | KD033022_033022_C12 | 4054178 | 0.244 | C12 | C3H7NO3 | 0.049 | 83567195

## Serine U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A1 | 18275465 | 1 | A01 | [13]C3H7[15]NO3 | N/A | N/A | 10.135:Peak area 18275465.445 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A2 | 24441817 | 1 | A02 | [13]C3H7[15]NO3 | N/A | N/A | 10.124:Peak area 24441817.080 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A3 | 27289227 | 1 | A03 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A4 | 31777615 | 1 | A04 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A5 | 36580588 | 1 | A05 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A6 | 46494187 | 1 | A06 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A7 | 58405804 | 1 | A07 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A8 | 68879315 | 1 | A08 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A9 | 77201989 | 1 | A09 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A10 | 82715737 | 1 | A10 | [13]C3H7[15]NO3 | N/A | N/A | 10.118:Peak area 82715736.961 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A11 | 86583246 | 1 | A11 | [13]C3H7[15]NO3 | N/A | N/A | 10.128:Peak area 86583245.804 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_A12 | 87183547 | 1 | A12 | [13]C3H7[15]NO3 | N/A | N/A | 10.157:Peak area 87183547.259 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B1 | 17800815 | 1 | B01 | [13]C3H7[15]NO3 | N/A | N/A | 10.169:Peak area 17800814.790 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B2 | 19869720 | 1 | B02 | [13]C3H7[15]NO3 | N/A | N/A | 10.152:Peak area 19869719.830 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B3 | 25475560 | 1 | B03 | [13]C3H7[15]NO3 | N/A | N/A | 10.149:Peak area 25475559.779 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B4 | 28932808 | 1 | B04 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B5 | 35101099 | 1 | B05 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B6 | 43187487 | 1 | B06 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B7 | 60777825 | 1 | B07 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B8 | 70144969 | 1 | B08 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B9 | 76765160 | 1 | B09 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B10 | 83470378 | 1 | B10 | [13]C3H7[15]NO3 | N/A | N/A | 10.148:Peak area 83470378.019 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B11 | 85687406 | 1 | B11 | [13]C3H7[15]NO3 | N/A | N/A | 10.153:Peak area 85687405.938 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_B12 | 91196249 | 1 | B12 | [13]C3H7[15]NO3 | N/A | N/A | 10.162:Peak area 91196248.587 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C1 | 15416191 | 1 | C01 | [13]C3H7[15]NO3 | N/A | N/A | 10.143:Peak area 15416190.524 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C2 | 18327370 | 1 | C02 | [13]C3H7[15]NO3 | N/A | N/A | 10.128:Peak area 18327370.365 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C3 | 22222792 | 1 | C03 | [13]C3H7[15]NO3 | N/A | N/A | 10.155:Peak area 22222791.830 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C4 | 28519430 | 1 | C04 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C5 | 34623388 | 1 | C05 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C6 | 44393757 | 1 | C06 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C7 | 58261066 | 1 | C07 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C8 | 67000200 | 1 | C08 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C9 | 76240950 | 1 | C09 | [13]C3H7[15]NO3 | N/A | N/A
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C10 | 82635974 | 1 | C10 | [13]C3H7[15]NO3 | N/A | N/A | 10.135:Peak area 82635973.571 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C11 | 85877633 | 1 | C11 | [13]C3H7[15]NO3 | N/A | N/A | 10.172:Peak area 85877633.133 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)
Serine U-13C, U-15N neg | 10.04 | Internal Standard | KD033022_033022_C12 | 83567195 | 1 | C12 | [13]C3H7[15]NO3 | N/A | N/A | 10.160:Peak area 83567194.656 is out of bounds (ISTD Minimum Recovery 26407277.191 and ISTD Max Recovery 79221831.574)

## Threonine neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Threonine neg | 9.15 | Target Compound | KD033022_033022_A1 | 2937933466 | 500 | A01 | C4H9NO3 | 135.125 | 21742277
Threonine neg | 9.15 | Target Compound | KD033022_033022_A2 | 1480585356 | 250 | A02 | C4H9NO3 | 47.35 | 31269152
Threonine neg | 9.15 | Target Compound | KD033022_033022_A3 | 896335707 | 125 | A03 | C4H9NO3 | 21.901 | 40926346
Threonine neg | 9.15 | Target Compound | KD033022_033022_A4 | 582303010 | 62.5 | A04 | C4H9NO3 | 11.821 | 49261808
Threonine neg | 9.15 | Target Compound | KD033022_033022_A5 | 401813233 | 31.25 | A05 | C4H9NO3 | 6.12 | 65650878
Threonine neg | 9.15 | Target Compound | KD033022_033022_A6 | 274429011 | 15.625 | A06 | C4H9NO3 | 3.234 | 84857716 | Calibration Amount Diff 23.665% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_A7 | 167452634 | 7.813 | A07 | C4H9NO3 | 1.596 | 104920874 | Calibration Amount Diff 23.289% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_A8 | 99778632 | 3.906 | A08 | C4H9NO3 | 0.797 | 125198088 | Calibration Amount Diff 23.768% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_A9 | 56351828 | 1.953 | A09 | C4H9NO3 | 0.457 | 123178098 | Calibration Amount Diff 42.4% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_A10 | 29832512 | 0.977 | A10 | C4H9NO3 | 0.217 | 137511246 | Calibration Amount Diff 35.196% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_A11 | 16869044 | 0.488 | A11 | C4H9NO3 | 0.124 | 136280180 | Calibration Amount Diff 54.528% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_A12 | 7969495 | 0.244 | A12 | C4H9NO3 | 0.057 | 140696164 | Calibration Amount Diff 41.486% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_B1 | 2641748819 | 500 | B01 | C4H9NO3 | 127.315 | 20749656
Threonine neg | 9.15 | Target Compound | KD033022_033022_B2 | 1407585375 | 250 | B02 | C4H9NO3 | 47.203 | 29819638
Threonine neg | 9.15 | Target Compound | KD033022_033022_B3 | 846658518 | 125 | B03 | C4H9NO3 | 22.665 | 37355906
Threonine neg | 9.15 | Target Compound | KD033022_033022_B4 | 531135817 | 62.5 | B04 | C4H9NO3 | 11.894 | 44656552
Threonine neg | 9.15 | Target Compound | KD033022_033022_B5 | 372996294 | 31.25 | B05 | C4H9NO3 | 6.152 | 60628624
Threonine neg | 9.15 | Target Compound | KD033022_033022_B6 | 273914204 | 15.625 | B06 | C4H9NO3 | 3.275 | 83628846 | Calibration Amount Diff 25.215% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_B7 | 159836846 | 7.813 | B07 | C4H9NO3 | 1.492 | 107154181
Threonine neg | 9.15 | Target Compound | KD033022_033022_B8 | 99605645 | 3.906 | B08 | C4H9NO3 | 0.814 | 122308768 | Calibration Amount Diff 26.458% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_B9 | 57114241 | 1.953 | B09 | C4H9NO3 | 0.433 | 131927372 | Calibration Amount Diff 34.777% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_B10 | 30860963 | 0.977 | B10 | C4H9NO3 | 0.224 | 137506953 | Calibration Amount Diff 39.855% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_B11 | 17504378 | 0.488 | B11 | C4H9NO3 | 0.12 | 145888405 | Calibration Amount Diff 49.791% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_B12 | 8818202 | 0.244 | B12 | C4H9NO3 | 0.059 | 148544352 | Calibration Amount Diff 48.28% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_C1 | 2539380955 | 500 | C01 | C4H9NO3 | 120.71 | 21036998
Threonine neg | 9.15 | Target Compound | KD033022_033022_C2 | 1243038383 | 250 | C02 | C4H9NO3 | 46.676 | 26631076
Threonine neg | 9.15 | Target Compound | KD033022_033022_C3 | 788974495 | 125 | C03 | C4H9NO3 | 23.438 | 33662558
Threonine neg | 9.15 | Target Compound | KD033022_033022_C4 | 517803094 | 62.5 | C04 | C4H9NO3 | 11.702 | 44247955
Threonine neg | 9.15 | Target Compound | KD033022_033022_C5 | 360494595 | 31.25 | C05 | C4H9NO3 | 6.062 | 59470323
Threonine neg | 9.15 | Target Compound | KD033022_033022_C6 | 260617361 | 15.625 | C06 | C4H9NO3 | 3.215 | 81055863 | Calibration Amount Diff 22.963% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_C7 | 155330208 | 7.813 | C07 | C4H9NO3 | 1.492 | 104094651
Threonine neg | 9.15 | Target Compound | KD033022_033022_C8 | 98813370 | 3.906 | C08 | C4H9NO3 | 0.812 | 121695846 | Calibration Amount Diff 26.086% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_C9 | 64015059 | 1.953 | C09 | C4H9NO3 | 0.461 | 138732864 | Calibration Amount Diff 43.624% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_C10 | 32187025 | 0.977 | C10 | C4H9NO3 | 0.232 | 138824750 | Calibration Amount Diff 44.473% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_C11 | 17413935 | 0.488 | C11 | C4H9NO3 | 0.119 | 146282706 | Calibration Amount Diff 48.616% (absolute) > max calibration amount diff 20%
Threonine neg | 9.15 | Target Compound | KD033022_033022_C12 | 7569495 | 0.244 | C12 | C4H9NO3 | 0.055 | 138746296 | Calibration Amount Diff 36.275% (absolute) > max calibration amount diff 20%

## Threonine U-13C, U-15N neg
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A1 | 21742277 | 1 | A01 | [13]C4H9[15]NO3 | N/A | N/A | 9.222:Peak area 21742276.951 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A2 | 31269152 | 1 | A02 | [13]C4H9[15]NO3 | N/A | N/A | 9.269:Peak area 31269151.615 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A3 | 40926346 | 1 | A03 | [13]C4H9[15]NO3 | N/A | N/A | 9.241:Peak area 40926346.118 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A4 | 49261808 | 1 | A04 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A5 | 65650878 | 1 | A05 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A6 | 84857716 | 1 | A06 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A7 | 104920874 | 1 | A07 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A8 | 125198088 | 1 | A08 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A9 | 123178098 | 1 | A09 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A10 | 137511246 | 1 | A10 | [13]C4H9[15]NO3 | N/A | N/A | 9.215:Peak area 137511245.616 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A11 | 136280180 | 1 | A11 | [13]C4H9[15]NO3 | N/A | N/A | 9.224:Peak area 136280179.861 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_A12 | 140696164 | 1 | A12 | [13]C4H9[15]NO3 | N/A | N/A | 9.250:Peak area 140696164.156 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B1 | 20749656 | 1 | B01 | [13]C4H9[15]NO3 | N/A | N/A | 9.228:Peak area 20749656.444 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B2 | 29819638 | 1 | B02 | [13]C4H9[15]NO3 | N/A | N/A | 9.296:Peak area 29819638.065 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B3 | 37355906 | 1 | B03 | [13]C4H9[15]NO3 | N/A | N/A | 9.265:Peak area 37355905.742 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B4 | 44656552 | 1 | B04 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B5 | 60628624 | 1 | B05 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B6 | 83628846 | 1 | B06 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B7 | 107154181 | 1 | B07 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B8 | 122308768 | 1 | B08 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B9 | 131927372 | 1 | B09 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B10 | 137506953 | 1 | B10 | [13]C4H9[15]NO3 | N/A | N/A | 9.244:Peak area 137506953.171 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B11 | 145888405 | 1 | B11 | [13]C4H9[15]NO3 | N/A | N/A | 9.248:Peak area 145888404.552 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_B12 | 148544352 | 1 | B12 | [13]C4H9[15]NO3 | N/A | N/A | 9.228:Peak area 148544352.276 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C1 | 21036998 | 1 | C01 | [13]C4H9[15]NO3 | N/A | N/A | 9.312:Peak area 21036998.312 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C2 | 26631076 | 1 | C02 | [13]C4H9[15]NO3 | N/A | N/A | 9.273:Peak area 26631076.223 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C3 | 33662558 | 1 | C03 | [13]C4H9[15]NO3 | N/A | N/A | 9.271:Peak area 33662557.845 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C4 | 44247955 | 1 | C04 | [13]C4H9[15]NO3 | N/A | N/A | 9.247:Peak area 44247955.196 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C5 | 59470323 | 1 | C05 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C6 | 81055863 | 1 | C06 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C7 | 104094651 | 1 | C07 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C8 | 121695846 | 1 | C08 | [13]C4H9[15]NO3 | N/A | N/A
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C9 | 138732864 | 1 | C09 | [13]C4H9[15]NO3 | N/A | N/A | 9.216:Peak area 138732863.586 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C10 | 138824750 | 1 | C10 | [13]C4H9[15]NO3 | N/A | N/A | 9.230:Peak area 138824749.649 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C11 | 146282706 | 1 | C11 | [13]C4H9[15]NO3 | N/A | N/A | 9.237:Peak area 146282705.806 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)
Threonine U-13C, U-15N neg | 9.15 | Internal Standard | KD033022_033022_C12 | 138746296 | 1 | C12 | [13]C4H9[15]NO3 | N/A | N/A | 9.250:Peak area 138746296.058 is out of bounds (ISTD Minimum Recovery 44251999.520 and ISTD Max Recovery 132755998.560)

## Tryptophan pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A1 | 55435740507 | 500 | A01 | C11H12N2O2 | 82.511 | 671862408
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A2 | 32262163617 | 250 | A02 | C11H12N2O2 | 35.147 | 917912344
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A3 | 19734879514 | 125 | A03 | C11H12N2O2 | 17.939 | 1100111650
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A4 | 11273110467 | 62.5 | A04 | C11H12N2O2 | 9.561 | 1179088096
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A5 | 6838254346 | 31.25 | A05 | C11H12N2O2 | 5.22 | 1309997121
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A6 | 3544784018 | 15.625 | A06 | C11H12N2O2 | 2.632 | 1346846755
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A7 | 1954555421 | 7.813 | A07 | C11H12N2O2 | 1.32 | 1480370837
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A8 | 997504250 | 3.906 | A08 | C11H12N2O2 | 0.668 | 1493358877
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A9 | 572722260 | 1.953 | A09 | C11H12N2O2 | 0.379 | 1511299337 | Calibration Amount Diff 32.83% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A10 | 298435881 | 0.977 | A10 | C11H12N2O2 | 0.183 | 1629514526 | Calibration Amount Diff 28.352% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A11 | 154030644 | 0.488 | A11 | C11H12N2O2 | 0.094 | 1641908093 | Calibration Amount Diff 31.641% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_A12 | 80171886 | 0.244 | A12 | C11H12N2O2 | 0.046 | 1737889223 | Calibration Amount Diff 29.475% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B1 | 51778636251 | 500 | B01 | C11H12N2O2 | 78.994 | 655474563
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B2 | 30894472927 | 250 | B02 | C11H12N2O2 | 35.351 | 873929504
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B3 | 19987014947 | 125 | B03 | C11H12N2O2 | 18.044 | 1107702420
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B4 | 11601196050 | 62.5 | B04 | C11H12N2O2 | 9.538 | 1216350203
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B5 | 6620853799 | 31.25 | B05 | C11H12N2O2 | 5.032 | 1315637224
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B6 | 3665334715 | 15.625 | B06 | C11H12N2O2 | 2.732 | 1341666515
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B7 | 1812525090 | 7.813 | B07 | C11H12N2O2 | 1.227 | 1477770669
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B8 | 1079215751 | 3.906 | B08 | C11H12N2O2 | 0.676 | 1596729992
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B9 | 599665136 | 1.953 | B09 | C11H12N2O2 | 0.361 | 1660455870 | Calibration Amount Diff 26.588% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B10 | 288389168 | 0.977 | B10 | C11H12N2O2 | 0.182 | 1583383872 | Calibration Amount Diff 27.645% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B11 | 154347804 | 0.488 | B11 | C11H12N2O2 | 0.092 | 1669924026 | Calibration Amount Diff 29.699% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_B12 | 77296646 | 0.244 | B12 | C11H12N2O2 | 0.046 | 1676097726 | Calibration Amount Diff 29.434% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C1 | 49744832558 | 500 | C01 | C11H12N2O2 | 80.324 | 619299322
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C2 | 29915661829 | 250 | C02 | C11H12N2O2 | 36.078 | 829189600
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C3 | 18645464079 | 125 | C03 | C11H12N2O2 | 18.816 | 990952442
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C4 | 10915712264 | 62.5 | C04 | C11H12N2O2 | 9.629 | 1133651378
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C5 | 6086189145 | 31.25 | C05 | C11H12N2O2 | 5.167 | 1177999366
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C6 | 3476543548 | 15.625 | C06 | C11H12N2O2 | 2.614 | 1329958929
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C7 | 1746599283 | 7.813 | C07 | C11H12N2O2 | 1.242 | 1405997967
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C8 | 983579303 | 3.906 | C08 | C11H12N2O2 | 0.67 | 1468597313
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C9 | 626634434 | 1.953 | C09 | C11H12N2O2 | 0.378 | 1658766457 | Calibration Amount Diff 32.413% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C10 | 313420241 | 0.977 | C10 | C11H12N2O2 | 0.187 | 1675250715 | Calibration Amount Diff 31.116% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C11 | 157122206 | 0.488 | C11 | C11H12N2O2 | 0.095 | 1657832599 | Calibration Amount Diff 32.993% (absolute) > max calibration amount diff 20%
Tryptophan pos | 7.3 | Target Compound | KD033022_033022_C12 | 75302318 | 0.244 | C12 | C11H12N2O2 | 0.047 | 1590650461 | Calibration Amount Diff 32.868% (absolute) > max calibration amount diff 20%

## Tryptophan U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A1 | 671862408 | 1 | A01 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A2 | 917912344 | 1 | A02 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A3 | 1100111650 | 1 | A03 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A4 | 1179088096 | 1 | A04 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A5 | 1309997121 | 1 | A05 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A6 | 1346846755 | 1 | A06 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A7 | 1480370837 | 1 | A07 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A8 | 1493358877 | 1 | A08 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A9 | 1511299337 | 1 | A09 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A10 | 1629514526 | 1 | A10 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A11 | 1641908093 | 1 | A11 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_A12 | 1737889223 | 1 | A12 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B1 | 655474563 | 1 | B01 | [13]C11H12[15]N2O2 | N/A | N/A | 7.381:Peak area 655474563.188 is out of bounds (ISTD Minimum Recovery 662964283.304 and ISTD Max Recovery 1988892849.911)
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B2 | 873929504 | 1 | B02 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B3 | 1107702420 | 1 | B03 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B4 | 1216350203 | 1 | B04 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B5 | 1315637224 | 1 | B05 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B6 | 1341666515 | 1 | B06 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B7 | 1477770669 | 1 | B07 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B8 | 1596729992 | 1 | B08 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B9 | 1660455870 | 1 | B09 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B10 | 1583383872 | 1 | B10 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B11 | 1669924026 | 1 | B11 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_B12 | 1676097726 | 1 | B12 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C1 | 619299322 | 1 | C01 | [13]C11H12[15]N2O2 | N/A | N/A | 7.353:Peak area 619299322.187 is out of bounds (ISTD Minimum Recovery 662964283.304 and ISTD Max Recovery 1988892849.911)
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C2 | 829189600 | 1 | C02 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C3 | 990952442 | 1 | C03 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C4 | 1133651378 | 1 | C04 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C5 | 1177999366 | 1 | C05 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C6 | 1329958929 | 1 | C06 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C7 | 1405997967 | 1 | C07 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C8 | 1468597313 | 1 | C08 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C9 | 1658766457 | 1 | C09 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C10 | 1675250715 | 1 | C10 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C11 | 1657832599 | 1 | C11 | [13]C11H12[15]N2O2 | N/A | N/A
Tryptophan U-13C, U-15N pos | 7.3 | Internal Standard | KD033022_033022_C12 | 1590650461 | 1 | C12 | [13]C11H12[15]N2O2 | N/A | N/A

## Tyrosine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A1 | 20379544180 | 500 | A01 | C9H11NO3 | 108.902 | 187136207
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A2 | 13480526965 | 250 | A02 | C9H11NO3 | 42.786 | 315067078
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A3 | 8420977233 | 125 | A03 | C9H11NO3 | 21.253 | 396216071
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A4 | 4890683555 | 62.5 | A04 | C9H11NO3 | 11.391 | 429356582
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A5 | 3287758011 | 31.25 | A05 | C9H11NO3 | 6.105 | 538503397
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A6 | 2152491273 | 15.625 | A06 | C9H11NO3 | 3.056 | 704272207
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A7 | 1432521048 | 7.813 | A07 | C9H11NO3 | 1.556 | 920694866 | Calibration Amount Diff 20.601% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A8 | 808224879 | 3.906 | A08 | C9H11NO3 | 0.781 | 1035431011 | Calibration Amount Diff 21.332% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A9 | 458775341 | 1.953 | A09 | C9H11NO3 | 0.445 | 1031732717 | Calibration Amount Diff 38.392% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A10 | 254789600 | 0.977 | A10 | C9H11NO3 | 0.216 | 1180196434 | Calibration Amount Diff 34.414% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A11 | 139306261 | 0.488 | A11 | C9H11NO3 | 0.113 | 1228490204 | Calibration Amount Diff 41.397% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_A12 | 74240390 | 0.244 | A12 | C9H11NO3 | 0.056 | 1329224205 | Calibration Amount Diff 39.315% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B1 | 19717497412 | 500 | B01 | C9H11NO3 | 104.495 | 188693795
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B2 | 13854305408 | 250 | B02 | C9H11NO3 | 41.864 | 330938501
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B3 | 8444961237 | 125 | B03 | C9H11NO3 | 21.446 | 393782655
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B4 | 4839804405 | 62.5 | B04 | C9H11NO3 | 11.415 | 423975989
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B5 | 3224305172 | 31.25 | B05 | C9H11NO3 | 5.963 | 540717479
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B6 | 2274472590 | 15.625 | B06 | C9H11NO3 | 3.185 | 714164681 | Calibration Amount Diff 22.783% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B7 | 1363975445 | 7.813 | B07 | C9H11NO3 | 1.438 | 948790149
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B8 | 818568295 | 3.906 | B08 | C9H11NO3 | 0.788 | 1038843891 | Calibration Amount Diff 22.478% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B9 | 508722481 | 1.953 | B09 | C9H11NO3 | 0.426 | 1195291294 | Calibration Amount Diff 32.469% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B10 | 259724208 | 0.977 | B10 | C9H11NO3 | 0.215 | 1208848584 | Calibration Amount Diff 33.77% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B11 | 147433774 | 0.488 | B11 | C9H11NO3 | 0.112 | 1311321739 | Calibration Amount Diff 40.194% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_B12 | 75741226 | 0.244 | B12 | C9H11NO3 | 0.056 | 1344614533 | Calibration Amount Diff 40.504% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C1 | 19385048695 | 500 | C01 | C9H11NO3 | 106.925 | 181296134
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C2 | 12702675772 | 250 | C02 | C9H11NO3 | 42.507 | 298838356
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C3 | 7843545541 | 125 | C03 | C9H11NO3 | 22.176 | 353687370
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C4 | 4778634745 | 62.5 | C04 | C9H11NO3 | 11.099 | 430558628
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C5 | 3106330511 | 31.25 | C05 | C9H11NO3 | 5.987 | 518804652
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C6 | 2223353033 | 15.625 | C06 | C9H11NO3 | 3.084 | 720985144
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C7 | 1346760680 | 7.813 | C07 | C9H11NO3 | 1.445 | 931842806
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C8 | 826180159 | 3.906 | C08 | C9H11NO3 | 0.79 | 1046261619 | Calibration Amount Diff 22.74% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C9 | 537083569 | 1.953 | C09 | C9H11NO3 | 0.446 | 1205512074 | Calibration Amount Diff 38.659% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C10 | 286121108 | 0.977 | C10 | C9H11NO3 | 0.219 | 1304178916 | Calibration Amount Diff 36.592% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C11 | 152000650 | 0.488 | C11 | C9H11NO3 | 0.116 | 1311482892 | Calibration Amount Diff 44.518% (absolute) > max calibration amount diff 20%
Tyrosine pos | 8.28 | Target Compound | KD033022_033022_C12 | 71147685 | 0.244 | C12 | C9H11NO3 | 0.057 | 1245736431 | Calibration Amount Diff 42.458% (absolute) > max calibration amount diff 20%

## Tyrosine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A1 | 187136207 | 1 | A01 | [13]C9H11[15]NO3 | N/A | N/A | 8.432:Peak area 187136207.169 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A2 | 315067078 | 1 | A02 | [13]C9H11[15]NO3 | N/A | N/A | 8.373:Peak area 315067078.055 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A3 | 396216071 | 1 | A03 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A4 | 429356582 | 1 | A04 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A5 | 538503397 | 1 | A05 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A6 | 704272207 | 1 | A06 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A7 | 920694866 | 1 | A07 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A8 | 1035431011 | 1 | A08 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A9 | 1031732717 | 1 | A09 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A10 | 1180196434 | 1 | A10 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A11 | 1228490204 | 1 | A11 | [13]C9H11[15]NO3 | N/A | N/A | 8.330:Peak area 1228490204.176 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_A12 | 1329224205 | 1 | A12 | [13]C9H11[15]NO3 | N/A | N/A | 8.328:Peak area 1329224204.558 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B1 | 188693795 | 1 | B01 | [13]C9H11[15]NO3 | N/A | N/A | 8.411:Peak area 188693795.111 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B2 | 330938501 | 1 | B02 | [13]C9H11[15]NO3 | N/A | N/A | 8.373:Peak area 330938500.559 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B3 | 393782655 | 1 | B03 | [13]C9H11[15]NO3 | N/A | N/A | 8.368:Peak area 393782654.561 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B4 | 423975989 | 1 | B04 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B5 | 540717479 | 1 | B05 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B6 | 714164681 | 1 | B06 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B7 | 948790149 | 1 | B07 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B8 | 1038843891 | 1 | B08 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B9 | 1195291294 | 1 | B09 | [13]C9H11[15]NO3 | N/A | N/A | 8.340:Peak area 1195291293.606 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B10 | 1208848584 | 1 | B10 | [13]C9H11[15]NO3 | N/A | N/A | 8.323:Peak area 1208848583.913 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B11 | 1311321739 | 1 | B11 | [13]C9H11[15]NO3 | N/A | N/A | 8.327:Peak area 1311321738.770 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_B12 | 1344614533 | 1 | B12 | [13]C9H11[15]NO3 | N/A | N/A | 8.334:Peak area 1344614533.170 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C1 | 181296134 | 1 | C01 | [13]C9H11[15]NO3 | N/A | N/A | 8.439:Peak area 181296133.870 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C2 | 298838356 | 1 | C02 | [13]C9H11[15]NO3 | N/A | N/A | 8.405:Peak area 298838356.423 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C3 | 353687370 | 1 | C03 | [13]C9H11[15]NO3 | N/A | N/A | 8.374:Peak area 353687370.134 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C4 | 430558628 | 1 | C04 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C5 | 518804652 | 1 | C05 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C6 | 720985144 | 1 | C06 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C7 | 931842806 | 1 | C07 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C8 | 1046261619 | 1 | C08 | [13]C9H11[15]NO3 | N/A | N/A
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C9 | 1205512074 | 1 | C09 | [13]C9H11[15]NO3 | N/A | N/A | 8.325:Peak area 1205512073.873 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C10 | 1304178916 | 1 | C10 | [13]C9H11[15]NO3 | N/A | N/A | 8.336:Peak area 1304178916.415 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C11 | 1311482892 | 1 | C11 | [13]C9H11[15]NO3 | N/A | N/A | 8.315:Peak area 1311482891.925 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)
Tyrosine U-13C, U-15N pos | 8.28 | Internal Standard | KD033022_033022_C12 | 1245736431 | 1 | C12 | [13]C9H11[15]NO3 | N/A | N/A | 8.326:Peak area 1245736431.322 is out of bounds (ISTD Minimum Recovery 395631795.685 and ISTD Max Recovery 1186895387.054)

## Valine pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Valine pos | 7.85 | Target Compound | KD033022_033022_A1 | 38882720237 | 500 | A01 | C5H11NO2 | 97.956 | 396939654
Valine pos | 7.85 | Target Compound | KD033022_033022_A2 | 25266780194 | 250 | A02 | C5H11NO2 | 43.773 | 577217810
Valine pos | 7.85 | Target Compound | KD033022_033022_A3 | 15617010787 | 125 | A03 | C5H11NO2 | 22.262 | 701517657
Valine pos | 7.85 | Target Compound | KD033022_033022_A4 | 10348035984 | 62.5 | A04 | C5H11NO2 | 11.508 | 899239321
Valine pos | 7.85 | Target Compound | KD033022_033022_A5 | 6630471663 | 31.25 | A05 | C5H11NO2 | 6.233 | 1063777453
Valine pos | 7.85 | Target Compound | KD033022_033022_A6 | 3647852355 | 15.625 | A06 | C5H11NO2 | 3.179 | 1147426052
Valine pos | 7.85 | Target Compound | KD033022_033022_A7 | 2003304525 | 7.813 | A07 | C5H11NO2 | 1.628 | 1230637358
Valine pos | 7.85 | Target Compound | KD033022_033022_A8 | 1047948077 | 3.906 | A08 | C5H11NO2 | 0.796 | 1316629753
Valine pos | 7.85 | Target Compound | KD033022_033022_A9 | 595435244 | 1.953 | A09 | C5H11NO2 | 0.473 | 1259777866 | Calibration Amount Diff 34.396% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_A10 | 296686127 | 0.977 | A10 | C5H11NO2 | 0.223 | 1329583142 | Calibration Amount Diff 26.852% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_A11 | 163894875 | 0.488 | A11 | C5H11NO2 | 0.114 | 1439388551 | Calibration Amount Diff 29.6% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_A12 | 108248633 | 0.244 | A12 | C5H11NO2 | 0.072 | 1503531313 | Calibration Amount Diff 63.896% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_B1 | 36896222592 | 500 | B01 | C5H11NO2 | 94.132 | 391963837
Valine pos | 7.85 | Target Compound | KD033022_033022_B2 | 24830553610 | 250 | B02 | C5H11NO2 | 42.669 | 581934836
Valine pos | 7.85 | Target Compound | KD033022_033022_B3 | 15230601193 | 125 | B03 | C5H11NO2 | 22.58 | 674524317
Valine pos | 7.85 | Target Compound | KD033022_033022_B4 | 10469974912 | 62.5 | B04 | C5H11NO2 | 11.691 | 895528874
Valine pos | 7.85 | Target Compound | KD033022_033022_B5 | 6618430693 | 31.25 | B05 | C5H11NO2 | 6.029 | 1097762054
Valine pos | 7.85 | Target Compound | KD033022_033022_B6 | 3861127106 | 15.625 | B06 | C5H11NO2 | 3.237 | 1192940382
Valine pos | 7.85 | Target Compound | KD033022_033022_B7 | 1872897233 | 7.813 | B07 | C5H11NO2 | 1.473 | 1271457277
Valine pos | 7.85 | Target Compound | KD033022_033022_B8 | 1166373443 | 3.906 | B08 | C5H11NO2 | 0.847 | 1377248169 | Calibration Amount Diff 20.378% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_B9 | 651403304 | 1.953 | B09 | C5H11NO2 | 0.453 | 1436926483 | Calibration Amount Diff 28.904% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_B10 | 341035506 | 0.977 | B10 | C5H11NO2 | 0.242 | 1408045891 | Calibration Amount Diff 37.687% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_B11 | 160584690 | 0.488 | B11 | C5H11NO2 | 0.111 | 1440642055 | Calibration Amount Diff 26.872% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_B12 | 78796141 | 0.244 | B12 | C5H11NO2 | 0.057 | 1392203141 | Calibration Amount Diff 28.844% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_C1 | 35934562672 | 500 | C01 | C5H11NO2 | 96.174 | 373639210
Valine pos | 7.85 | Target Compound | KD033022_033022_C2 | 22287274783 | 250 | C02 | C5H11NO2 | 43.311 | 514585747
Valine pos | 7.85 | Target Compound | KD033022_033022_C3 | 14186200894 | 125 | C03 | C5H11NO2 | 23.247 | 610234571
Valine pos | 7.85 | Target Compound | KD033022_033022_C4 | 9799672341 | 62.5 | C04 | C5H11NO2 | 11.626 | 842915123
Valine pos | 7.85 | Target Compound | KD033022_033022_C5 | 5970833401 | 31.25 | C05 | C5H11NO2 | 6.104 | 978111553
Valine pos | 7.85 | Target Compound | KD033022_033022_C6 | 3561506872 | 15.625 | C06 | C5H11NO2 | 3.13 | 1137806428
Valine pos | 7.85 | Target Compound | KD033022_033022_C7 | 1821319089 | 7.813 | C07 | C5H11NO2 | 1.533 | 1187750040
Valine pos | 7.85 | Target Compound | KD033022_033022_C8 | 1040547803 | 3.906 | C08 | C5H11NO2 | 0.828 | 1257024768
Valine pos | 7.85 | Target Compound | KD033022_033022_C9 | 688419810 | 1.953 | C09 | C5H11NO2 | 0.5 | 1376956661 | Calibration Amount Diff 42.158% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_C10 | 314240485 | 0.977 | C10 | C5H11NO2 | 0.222 | 1418485251 | Calibration Amount Diff 25.937% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_C11 | 229566410 | 0.488 | C11 | C5H11NO2 | 0.164 | 1403486440 | Calibration Amount Diff 86.169% (absolute) > max calibration amount diff 20%
Valine pos | 7.85 | Target Compound | KD033022_033022_C12 | N/F | 0.244 | C12 | C5H11NO2 | N/F | 1364585819

## Valine U-13C, U-15N pos
Compound | RT | Type | Filename | Area | Theoretical Amt | Sample ID | Formula | Response Ratio | ISTD Response | Flag Details
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A1 | 396939654 | 1 | A01 | [13]C5H11[15]NO2 | N/A | N/A | 7.877:Peak area 396939654.174 is out of bounds (ISTD Minimum Recovery 534617011.906 and ISTD Max Recovery 1603851035.719)
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A2 | 577217810 | 1 | A02 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A3 | 701517657 | 1 | A03 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A4 | 899239321 | 1 | A04 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A5 | 1063777453 | 1 | A05 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A6 | 1147426052 | 1 | A06 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A7 | 1230637358 | 1 | A07 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A8 | 1316629753 | 1 | A08 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A9 | 1259777866 | 1 | A09 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A10 | 1329583142 | 1 | A10 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A11 | 1439388551 | 1 | A11 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_A12 | 1503531313 | 1 | A12 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B1 | 391963837 | 1 | B01 | [13]C5H11[15]NO2 | N/A | N/A | 7.883:Peak area 391963836.957 is out of bounds (ISTD Minimum Recovery 534617011.906 and ISTD Max Recovery 1603851035.719)
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B2 | 581934836 | 1 | B02 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B3 | 674524317 | 1 | B03 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B4 | 895528874 | 1 | B04 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B5 | 1097762054 | 1 | B05 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B6 | 1192940382 | 1 | B06 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B7 | 1271457277 | 1 | B07 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B8 | 1377248169 | 1 | B08 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B9 | 1436926483 | 1 | B09 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B10 | 1408045891 | 1 | B10 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B11 | 1440642055 | 1 | B11 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_B12 | 1392203141 | 1 | B12 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C1 | 373639210 | 1 | C01 | [13]C5H11[15]NO2 | N/A | N/A | 7.884:Peak area 373639210.123 is out of bounds (ISTD Minimum Recovery 534617011.906 and ISTD Max Recovery 1603851035.719)
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C2 | 514585747 | 1 | C02 | [13]C5H11[15]NO2 | N/A | N/A | 7.879:Peak area 514585746.802 is out of bounds (ISTD Minimum Recovery 534617011.906 and ISTD Max Recovery 1603851035.719)
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C3 | 610234571 | 1 | C03 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C4 | 842915123 | 1 | C04 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C5 | 978111553 | 1 | C05 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C6 | 1137806428 | 1 | C06 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C7 | 1187750040 | 1 | C07 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C8 | 1257024768 | 1 | C08 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C9 | 1376956661 | 1 | C09 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C10 | 1418485251 | 1 | C10 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C11 | 1403486440 | 1 | C11 | [13]C5H11[15]NO2 | N/A | N/A
Valine U-13C, U-15N pos | 7.85 | Internal Standard | KD033022_033022_C12 | 1364585819 | 1 | C12 | [13]C5H11[15]NO2 | N/A | N/A